# Single House Commissioning Plan

## Tunnel Work

### Connect pbar cables to pbar ends of BPM

## Service Building Decommissioning

### Remove old BPM modules and NIM crates

### Remove old BPM power supplies

### Remove TIM chassis

### Disconnect proton cables from analog box

### Remove analog box

### Remove External Device Bus chassis

## Service Building Hardware Installation

### Install VME crate

### Install cards in VME crate (done sooner?)

### Cabling

#### Connect proton and pbar heliax cables to transition point (done sooner?)

#### Connect beam signal cables from transition point to EchoTek cards

#### Connect Ethernet cable to processor

#### Connect BSYNC signal to timing card

#### Connect TCLK signal to processor and timing card

#### Connect 53MHz to timing card PLL

#### Connect External Device Bus cable between BLM chassis and BLM interface on BPM system

#### BPM Intra-card cabling (done sooner?)

##### Connect 80MHz clock to all EchoTek modules from timing module

##### Connect triggers to all EchoTek modules from timing module

##### Connect outputs of diagnostics boards to inputs of EchoTek modules

##### Connect diagnostic reference signal to all diagnostics boards.

##### Connect diagnostic board control and power

#### Connect VME Power

### Install necessary CAMAC modules?

## Power-up Sequence

### Turn on crate

### Allow crate to boot, examine FP status LEDs

### Processor test and setup

#### Verify Ethernet communication with processor (ping?)

#### Download configuration parameters (done on boot?)

#### Verify TCLK connection to processor board

#### Place timing card in diagnostic mode and verify change.

#### Examine and verify EchoTek module memory and setup.

### Timing card tests

#### Verify proper PLL output frequency (80 MHz?)

#### Verify triggers (closed orbit & TBT)

#### Verify communication with BLM system

#### Set up diagnostic signal for power up tests

#### Verify diagnostic signal for power up tests

### EchoTek card tests

#### Drive beam signal inputs with signal generator.

##### Verify raw signal, bypass downconverter

##### Verify closed orbit bandwidth measurement

##### Verify 10 Hz closed orbit bandwidth measurement

##### Verify TBT bandwidth measurement

#### Connect tunnel cables to beam signal inputs

##### Enable diagnostic outputs for channel 1 & 5 of each card.

##### Measure closed orbit bandwidth response of each signal path.

##### Archive the response measurements

##### Repeat the last three steps for channel pairs 2 & 6, 3 & 7, and 4 & 8.

## Beam First Turn Verification

### Scope verification

#### Compare timing of first turn trigger output from timing card with beam signal.

#### Verify that trigger is within 3 turns of beam injection.

#### Compare timing of sample trigger (AA marker) to beam signal.

#### Verify that trigger is within aperture delay of EchoTek sample.

#### Repeat for all proton channels.

### Trigger fine adjustment

#### Place EchoTek cards into raw data mode.

#### Trigger acquisition with first turn trigger.

#### Verify that first turn arrives at correct position in memory, fine tune trigger delay.

#### Place EchoTek cards into TBT mode and verify first turn goes to correct position in memory for all proton channels.

## Turn-by-turn Verification

### Place EchoTek cards into raw data mode.

### Trigger acquisition with turn marker (AA).

### Verify that beam signal occupies the proper location in memory, tune trigger delay.

### Place EchoTek modules into TBT mode.

## Closed Orbit Verification

### Place EchoTek cards into closed orbit mode.

### Verify reasonable position and intensity values.

## Proton Calibration

### Perform local bump around BPM and process a 7x7 grid of positions.

### Use closed orbit I&Q data from each of the inputs for each of the positions to calculate corrections for scaling, non-linearities, and directivity (pbar only).

### Save coefficients for corrections in memory for use in corrected position calculations.

## On-line Configuration Parameters

### Make changes to online configuration parameters to inform library functions and applications of change in house.

### Verify proper on-line communication and operation with diagnostics applications.