Fermilab/AD/TEV

Beams-doc1165-v2

# May 10, 2004

Specification for Passive Band Pass / Anti-aliasing Filter for the Tevatron BPM System

Vince Pavlicek

Tevatron BPM Project

May 10, 2004

# Scope

This note specifies an analog RF band pass filter to select the 53 MHz component of a complex analog signal presented by the Tevatron Beam Position Monitors. The filter serves an anti-aliasing function to permit digital under-sampling of the 53 MHz signal with a ~75 MHz A-D converter. Additionally, the impulse response of the filter must facilitate differentiation of impulses separated by 400 nanoseconds.

# Features

## Electrical Specifications

Generic Filter Type – Band pass

Filter Characteristic – Gaussian or Bessel preferred, Butterworth acceptable (see Impulse Response spec)

Source and Load Impedance – 50 ohms

Continuous Peak Input Signal – ± 30 volt impulses at 21 MHz repetition rate

Maximum Continuous Signal Power within 3 dB Bandwidth – 20 dBm

Center Frequency – 53.10 MHz

3 dB Bandwidth – approximately 6 to 8 MHz (see Impulse Response spec)

Flatness Across Pass Band – better than 1 dB over band ± 0.5 MHz of center frequency

Maximum Insertion Loss at Center Frequency – 3 dB

Minimum Return Loss – 15 dB over band ± 0.5 MHz of center frequency

Stop Band Attenuation – 50 dB or greater (relative to 53 MHz carrier) for all frequencies less than 30 MHz and greater than 95 MHz up to 250 MHz

Impulse Response – Envelope of impulse response must be down to 2% of peak envelope response at all times  400 ns after impulse

## Physical and Environmental Specifications

Package – circuit board mount package w/solder pins.  Maximum height 0.54 inches.  Package board area less than 1 sq inch is desired, with a maximum of 1.5 sq inch.

Environmental – hermetically sealed package

Identification – individual package ID with series/serial number.

Storage Temperature Range –  -20 to +45 degrees Celsius

Operating Temperature Range – +20 to +40 degrees Celsius

Shock – Withstand 1.5G without change in electrical performance greater than 0.1 dB insertion loss and 2 degrees transmission phase anywhere in 3 dB bandwidth

## Matching of Filter Pairs (option #1, bid separately)

Filters will be used in pairs.  Over the full 3 dB bandwidth paired filters are matched to within 0.1 dB insertion loss and to within 2 degrees transmission phase.  Provide documentation of test results to assist with rework and spare organization.

# Quantity

The total number of crates to ordered is .

# Acceptance Criteria

Delivered crates must meet manufacturer minimum specifications.

# References

# Change Log

| Version | Issue Date | Concurrence | Description of Change |
| --- | --- | --- | --- |
| 1.0 |  | SW | Original |

# Concurrence

The following persons concur with this document.

Steve Wolbers, Project Manager	(date)

Bob Webber, Deputy Project Manager	(date)

Jim Steimel, Technical Coordinator		(date)