# NuMI Toroid Scope and Functionality

Peter Prieto

7-May-04

This note describes the toroid elements used in NuMI transport line to measure the beam current.

## CT’s (Toroids)

NuMi beam line uses two Pearson current transformers (CT) with a scale factor of 1V/A to measure the integrated batch intensity. The scale factor of the CT is set by the burden resistor. In a standard Pearson CT, the stamped scale factor becomes 0.5 V/A when terminated into50 Ω. The CT has an RCA connector serving as an output, attached to this single-ended connector a male-to-male extender is used to attach a Pomona box containing a band-limiting filter with a corner frequency of 10 MHz made up of two 39 picoFarad capacitors and 5 turns of wire on a 3/8 in OD. 4C4 core making 1.2 microHenry inductance. The Pomona box turns the singe-ended signal into a differential signal and the output is connected through a twinax connector.

## Connections

The output of the Pomona box connects through a trumpeter TWC-78-2 cable using a twinax connector for both the upsteam CT (E:TOR101 now called E:NUPT) and the downstream CT (E:TORTGT).

E:NUPT will be instrumented at MI-60 North using NIM bin in rack MI60219.  Digital interface for the reading will be a CAMAC 284 module in Crate $73 N18.  This crate is also in rack MI60219.  The CAMAC 479 module in Crate $70 N17 will provide a timer for sampling. (Acnet parameter E:xxxxxx)  Crate $70 is in rack MI60227.

E:TORTGT will be instrumented at MI-65 underground Power Supply Support Room using NIM bin in rack THSR104.  Digital interface for the reading will be a CAMAC 284 module in Crate $E6 N14.  This crate is in rack THSR103.  The CAMAC 479 module in Crate $E6 N16 will provide a timer for sampling. (Acnet parameter E:xxxxx)

## Beam Pipe Mounting

The Pearson CT’s go around the outside of the beam pipe. There is a ceramic gap installed in the beam pipe where the CT overlaps the beam pipe. The CT is held in place by a stand described by drawing 8890.127-MC-422585. The stand is designed so the beam image current can flow across the ceramic gap by flowing along the inner walls of the CT stand.

## Integrator Module

The toroid integrator module receives a trigger signal (Acnet parameter E:xxxx)  from a MI clock system through a BNC in the rear panel to start the measurement. The CT signal connects through a 78 Ω twinax cable in the rear panel of the module and received as a differential signal. The burden resistor across the differential pair is a 127 Ω resistor with a 100 Ω resistor connected from each terminal to ground. The resistor combination presents a 78 Ω termination to the CT. The scale factor of 1V/A becomes 0.609 V/A. The signal gets integrated and digitized and is available as a 16-bit signal, on a D-connector on the rear of the module. Also the digitized signal goes to an FPGA where the integrated and digitized signal period is compared to timing set by on-board switches. The output of the FPGA is converted to an analog signal via a DAC, then amplified with a gain of 4 and is available on the rear panel BNC (50Ω driver). A front panel high-impedance monitor signal is also available.

## Toroid Scale Factor

The analog signal has a scale factor of 1.142 E12/Volt.  The digital signal is four times smaller or 0.2855e12/Volt.