Beam Loss Monitor Upgrade Users’ Guide

Al Baumbaugh
Charlie Briegel
Craig Drennan
Brian Fellenz
Kelly Knickerbocker
Jonathan Lewis
Marvin Olson
Stephen Pordes
Randy Thurman-Keup
Mike Utes
Jin-yuan Wu

June 21, 2010

BEAMS-DOC-1410-V9

Abstract

The new BLM readout system is designed to perform several tasks:  to provide a flexible and reliable abort system to protect Tevatron magnets; to provide loss monitor data during normal operations of the Tevatron, Main Injector and Booster; and to provide detailed diagnostic loss histories when an abort happens.   Beam losses are detected using the same ion chambers that have been used with the legacy system.

# System Overview

The basic principle of operation of the new BLM system is to integrate for a short period of time, typically 21 μs, and digitize to 16 bits.  There are two integrators per channel, running in a “Ping-Pong” mode, alternating between charge integration and digitization, so that no loss is missed.  While one channel is integrating, the other is digitized, its integrator is reset, and the data are processed.   The reset and processing time set a lower limit of 15 μs.  The digital data are used to construct several numbers that are compared against thresholds to generate abort signals.  These constructed data are sliding sums, which are a measure of the integrated loss over a variety of time scales from a single reading to a period of up to 64k cycles.  The abort signal is made in firmware by looking at these sums and thresholds as well as the number of channels requesting an abort.

The new BLM system uses a standard 6Ux160mm VME format crate.  Besides the VME crate computer in Slot 1 that communicates data to the main control system, the BLM system includes five types of custom cards:

Digitizer Cards (DC)

Timing Card (TC)

Control Card (CC)

High Voltage card (HV)

Abort Card (AC)

A custom J2 backplane is used for local system communication.  A Control Bus using the user-defined pins on the J2 VME connector handles all of the critical BLM controls.  This bus has 13 address lines and 8 data lines.  The CC is the only master on this bus, and the other cards are slaves. Also on the J2 connector is an Abort Bus where the AC is the master and the digitizer cards are the slaves (see Figure 1).

Figure

1: Block diagram of a BLM crate.

Figure

2: BLM Crate at A0 in the Tevatron.  The A0 abort card is present only in this crate and serves as the controller and concentrator for the ring-wide abort.

The sliding sum time scales and corresponding buffers and abort channels are referred to either by the sum (or abort number) or intended time scale.  These are as follows:

| Number | Name | Typical Time Scale | Circular Buffer Depth |
| --- | --- | --- | --- |
| 0 | Immediate | 20 μs | 64k |
| 1 | Fast | 1 ms | 16k |
| 2 | Slow | 50 ms | 4k |
| 3 | Very Slow | 1 s | 4k |

In this document, we include a summary description of each of the components followed by a description of the bus and communications protocol and detailed descriptions of the functions performed by each module including address maps.

## Digitizer Card

The Digitizer Card (DC) integrates and digitizes the current from four loss monitor chambers each beam revolution. To avoid dead time between measurements, signals for each input are switched between the two channels of a TI/Burr-Brown ACF2101 integrator chip (see [2] for tests of the integrator chip).  Results are digitized from the two channels on alternate cycles and fed to on-board FPGA devices.

The digitizer has 16 bit resolution with scaling such that one digitizer count represents 15.26 fC of charge in the integrator.  The sensitivity of the BLM ion chamber is approximately 70 nC of charge per Rad.

The logic maintains three running sums per channel with programmable durations of up to 65,536 base clocks (1.4 seconds for the Tevatron) and compares the current measurement and the running sums to abort thresholds (4 thresholds in all).  Each threshold can be set independently for each channel.  There can be up to 15 digitizer cards in a crate.  We envision sliding sums with periods of approximately 1 ms, 50 ms and 1s for normal operation.

The block diagram in Figure 3 illustrates the signal processing for each channel.  Note that the Sum registers will be read and the Threshold Registers written over the BLM Control Bus.  The SRAM memory which stores the integrator output values can be read over the VME bus (J1) by the crate computer.

Figure

3: Block diagram of the signal processing for one of the four channels on the Digitizer Card.

## Timing Card

The Timing Card (TC) receives accelerator system-wide timing information from three sources, the Tevatron Clock (TCLK), the Beam Sync Clock (BSCLK) and Machine State Data (MDAT).

The TC decodes BSCLK to generate the BLM system master clock which it distributes on the BLM Control Bus.  For the Tevatron this will be generated from the AA marker with a ~21 μs period.  For the Main Injector it will be half the AA marker frequency for a period of ~22 μs.   The master clock signal is known as Make_Meas (“Make Measurement”).

The TC maintains a 64k circular buffer of timing information for each cycle including a 32-bit Unix time (seconds since 1970) and a 24-bit microsecond counter which is reset at one second intervals; this buffer is in parallel with the circular buffer of loss measurements in the Digitizers.  The master clock defines the integration interval of the Digitizers and sets the threshold-comparison timing and abort-logic comparison timing.  The TC also generates signals at appropriate intervals to cause the Digitizers to latch the current values of the sliding sums and the Control Card to read these sums with the latched timing information.

The TC decodes TCLK and puts TCLKs requested by the CC into a queue that is readable by the CC through the Control Bus.  In addition, other events from TCLK are used to signal the BLM system to collect and store synchronous ring-wide data samples for beam studies.

The MDAT signal is decoded to determine the machine state.  State changes are made available to the CC through another Control Bus accessible queue that the Control Card reads to determine the appropriate abort thresholds and logic to load when the machine state changes.

## Control Card

To ensure that data communications and other tasks running on the VME crate computer do not impact the reliability of the BLM abort logic, the Control Card (CC) provides an independent dedicated processor that manages the setting of abort thresholds and other parameters used in the abort logic.  The Control Card CPU is a Zilog eZ80--a 24-bit address, 8-bit data, 45 MHz microcontroller.

The CC communicates with the other system cards over the dedicated custom J2 backplane keeping local communications separate from VME data transfers.  The CC also maintains circular buffers that store the histories of the three running sums for each digitizer channel with time stamps provided by the TC.  The histories will be at least 4096 time bins deep and are stored in a dual port memory, allowing the crate computer to read them at any time.  In addition, the most recent entry can be saved to one of three different buffers in response to a Flash, Profile, or Display TCLK event.  The CC also stores abort thresholds for each of the sums for each channel for up to 256 machine states. When a change in accelerator state is detected, the CC updates the thresholds in the Digitizer Cards as well as the abort masks and multiplicity requirements in the Abort Card.

## Abort Card

The four abort signals from each channel on each Digitizer Card are read by the Abort Card every integration interval.  The aborts of a particular type are counted and compared to a programmable multiplicity requirement for that abort type.  It is possible to mask channels off in the AC so they do not participate in the count.  If the multiplicity for that integration interval equals or exceeds the threshold, a beam abort signal is generated.  This logic is illustrated in Figure 4.  To accommodate the different operating conditions, the abort masks and multiplicity thresholds in the Abort Card can change depending on the Machine State.  We have also included a serial link on the Abort Card to allow a single point to receive information from all the BLM crates around the ring to be able to implement a ring-wide abort condition.

Figure

4: Abort Card multiplicity logic.

## Turn-by-Turn Buffers

In addition to the diagnostic buffers maintained by the Controller Card, which are circular buffers that are periodically overwritten, the BLM system has two linear buffers in the Digitizers and Timing Card that are triggered and are not automatically overwritten.  These buffers are Turn-By-Turn (TBT) and are each 8k deep.  The TBT Buffers are designed to match the capability of the new BPM system and allow the simultaneous sampling of beam position in the BPM and beam losses in the BLM.

The injection TBT buffer (ITBT) is designed to match the BPM injection TBT buffer and is triggered by ITBT_Trig at injection.  The studies TBT buffer is designed to match the BPM TBT buffer used for beam studies and is triggered by STBT_Trig.  Both of these triggers are generated on the TC by clock events on TCLK or BSCLK.  When either of these TBT buffers is triggered, the DCs and TC will set a bit in a status register indicating the TBT operations are in progress and that the TBT memory is not accessible from VME.  Once the TBT operation completes, or if an abort happens, the DCs and TC will reset the status bit, and VME will have access to the TBT memory.  TBT operations require the DCs and TC to not only write the large 64k circular buffer each sample period, but to also write the data into the TBT buffer each sample period.

Once it has been triggered, the TBT buffer will fill to its limit of 8k and stop.  It will not be overwritten until the next relevant clock event happens.  If another trigger happens prior to the completion of the current TBT operation, the TBT pointer will be reset to 0, and 8k of new TBT data will be written into the TBT buffer.

If a STBT_Trig happens during an ITBT operation, the ITBT operation is ended and the STBT pointer is set to 0 and a STBT operation is started.  Thus only a single TBT operation is allowed at any given time.  If this happens the host should be able to fully reconstruct the ITBT data using the ITBT and STBT buffers and their time stamps.

## Chassis

The Chassis is an integrated 6Ux160mm VME crate, power supply and fan fabricated by Weiner.  In addition to the J1 backplane that is being fabricated by Weiner, each crate includes a custom J2 backplane that handles the BLM control bus with all lines bussed on the A and C rows for slots 4-21.  Slots 1-3 will have no backplane connections on rows A and C.  Row B includes the standard extensions for A32D32 VME operation.  The power supply blocks the rear of the backplane, so transition modules cannot be used in a BLM crate.  The Wiener fan tray also includes an Ethernet – TCP/IP interface which enables slow control and monitoring via web browser, and provides a serial link to the crate processor via telnet.  See [3], [4], and [5] for various performance studies of the custom backplane.

## BLM Crate Normal Operations Sequence

Once the CC loads the settings into the TC, DCs and AC, the system is ready to run.  The BLM operations are initiated by a clock event such as “Prepare for Beam” which will cause the CC to tell the TC to issue a Digitizer Card Reset (DC_Reset) on the control bus.  The DC_Reset causes the DCs to zero all sliding sums and causes the DCs, TC, and AC to reset all circular buffer pointers.   This assures that all buffers are synchronized and ready to take data.

The primary clock for the BLM system, Make_Meas, is derived from the AA marker on the beam sync clock (typically 21 μs). Make_Meas is transmitted on the BLM control bus to all BLM cards.  Optionally the Make_Meas signal can be created by dividing the AA marker or by dividing down an internal clock.  The shortest allowable period for this signal is 15 microseconds due to the reset time needed by the integrators.

On the digitizer cards, the Make_Meas signal defines the sample period, causing the ACF2101 integrators to switch between channels for each input and triggering the ADCs to digitize the charge for the channel not being integrated.  After that, the sliding sums are updated and all abort comparisons are made.  At this time the new ADC readings are written to a 64k deep circular buffer which is used for diagnostic purposes as well as the source of the sliding sums.  The new ADC data may also be written to one of two turn-by-turn (TBT) dedicated studies buffers.  The abort states are latched on the next Make_Meas.  Thus the DC has the full sample period to do its conversions, make the sliding sums and do the abort compare with thresholds.  The timing card stores real-time clock data on each cycle in a 64k deep circular buffer that is synchronized with those of the digitizers.

On the AC, the Make_Meas signal causes the abort summing state machine to cycle through each BLM channel by putting the channel address ACS(5:0) on the abort bus and to read back from each channel the state of each of its abort requests ABORT(3:0).  For each abort type, each channel has an abort mask bit which determines if that channel is allowed to request an abort of that type.  A count is made for each of the four abort types of allowed AND requesting channels (i.e. those above threshold). If the number of channels requesting an abort for any of the four abort types equals or exceeds the abort multiplicity setting for that abort type, an abort request is transmitted from the card on a 50Ω TTL line driver.

The Make_Meas signal, therefore, causes the data to be taken and the abort logic to be updated every cycle. While a sliding sum might be the sum over 500 samples (10 ms) its abort threshold is compared every 21 µs.

During each 21μs cycle, the DCs make and update the three sliding sums of samples.  These sliding sums are compared every cycle to their abort limits.  However, for diagnostic purposes, these sums are stored periodically in circular buffers on the Control Card.  This process is controlled by the TC, which periodically generates 3 latch signals, one for each sliding sum.  The latch signals cause the DCs to latch the appropriate sum and the TC to latch the time stamp and to interrupt the CC so that it knows the data is latched and ready to be read and stored in the appropriate circular buffer.  The individual ADC readings are 16 bits; however, the sliding sums are 32 bit numbers.  Therefore, the dynamic range of, for example, the 1 second sliding sum is almost 32 bits.  These sliding sums are the total integrated loss over the sum interval, not just samples of losses spaced in time.

At any given time, the BLM has a variety of stored loss histories with different time resolutions:  the 64k deep raw measurement buffer provides 1.4 seconds of loss data with 21 µs resolution; the 16k Fast circular buffer provides 16 seconds of integrated loss data with 1 ms resolution, the 4k Slow circular buffer provides 200 seconds of integrated loss data with 50 ms resolution; and the 4k Very Slow buffer provides 4096 seconds, over an hour, of integrated loss data with 1 second resolution.   As one can see, in the event of an abort, there is a very detailed history of losses prior to the abort, which may be examined to aid in diagnosing the problem.

## Card Addresses

The system communicates with the outside world via ACNET through a standard VME host CPU.  This VME host sets the parameters in the BLM system trough a block of shared memory in the Control Card.  An overview of the BLM addressing scheme is given in Table 1.  In this document, unless otherwise specified, all VME and control-bus addresses are specified in hexadecimal.  VME addressing is based on 8-bit bytes, but the BLM system uses 16-bit data.  Therefore, the lowest address bit (A0) is always assumed to be 0 for VME.

Table

1: Overview of BLM addressing scheme.  Detailed switch/jumper settings are available in [1].

| Board Type and # | Switch or Jumper Setting (hex) | VME Address Type | VME Address range | Control Bus Address Range |
| --- | --- | --- | --- | --- |
| Control Card | 09 | A32 | 09000000 – 09FFFFFE | Master |
| Timing Card | 0B | A32 | 0B000000 – 0B0FFFFE | 1000 – 10FF |
| Timing Card Bridge |  | A32 | 0B100000 – 0B101FFE | Master |
| Abort Card | 0A | A32 | 0A000000 – 0A3FFFFE | 1100 – 11FF |
| A0 Abort Card | 0C | A32 | 0C000000 – 0C3FFFFE | N/A |
| HV Card | 08 | A32 | 08000000 – 08000FFE | N/A |
| Digitizer #0 | 0 | A24 | 000000 –0FFFFE | 0000 – 00FF |
| Digitizer #1 | 1 | A24 | 100000 –1FFFFE | 0100 – 01FF |
| Digitizer #2 | 2 | A24 | 200000 –2FFFFE | 0200 – 02FF |
| - - - - - - - | - - - - - | - - - - - | - - - - - - - - - - - - - - - - - | - - - - - - - - - |
| Digitizer #D | D | A24 | D00000 –DFFFFE | 0D00 – 0DFF |
| Digitizer #E | E | A24 | E00000 –EFFFFE | 0E00 – 0EFF |
| Digitizer Broadcast | N/A | N/A | N/A | 1300 – 13FF |

The BLM system has essentially 3 busses, a VME bus, the BLM Control bus on the A row of J2, and the Abort Bus on the C row of J2.  All operations on the Abort bus are done from the AC via its state machine.  All operations on the Control bus are done from the CC via eZ80 software.  In systems without a CC, VME can access the control bus only via a bridge on the TC.  Normally VME places data into the shared memory on the CC, and the CC puts that data into the hardware.  The pin assignments of the J2 backplane are shown in Table 2.  (Note that slots 1-3 on J2 are standard VME A32D32 with straight-through pins on rows A and C.)  All hardware parameters are sent by the CC over the dedicated BLM control bus.  This control bus has 13 address lines (CA12:0) and 8 data lines (CD7:0).  The digitizers use A24D16 addressing on VME while all other boards use A32D16.  All BLM data are integers where the lower byte corresponds to the less significant bits.

There can be up to 15 digitizers per crate.   Four switches on the digitizers set the high four bits of the VME address (A23:20).  They also set the high four bits of the control bus address and the abort channel select (Abort_CS).  Each card has 256 bytes of address space (CA(7:0)) which makes up its memory register space.  The Control Cards maintains a setup table for each of a possible 256 machine states, and will load this data into the hardware in response to an MDAT state change. These settings affect only the Digitizer Cards and the Abort concentrator card.  The settings for the Timing card do not change with machine state, and should only be changed during beam off periods.  An important feature of the BLM system is that all abort operations are handled by state machines, once setup by the CC, these operations proceed without intervention from either VME or eZ80.  The only things that the CPUs do is setup the parameters.

In order to smoothly update the settings in the DCs and AC, these cards must double buffer all registers. This double buffering is accomplished in two ways.  The AC transfers data to the actual usage register via a backplane signal, Update_Settings, which will occur after all settings have been written and will be generated on the TC, synchronously with Make_Meas, in response to a command from the CC.  When the DC is in original memory map mode, it functions just like the AC.  When the DC is in native memory map mode, it utilizes a paging system for storing thresholds and as such doesn’t use the Update_Settings line (with a few exceptions).  The double buffering is accomplished by the CC writing to a different page than is currently used and then switching the DC to the new page.  See Section 2.2 for details about the memory models in the DC.

## Backplane Signal Functions

The operation of the digitizer card is controlled by signals on the control bus.  These signals are:

Reset  [C21]:  This signal is the BLM Crate reset signal driven by the Controller Card (CC).  It resets the state machine on the DC and TC.  This pin is driven by eZ80 I/O pin PB1.

Reset_DC  [C20]  “Reset Digitizer Card.”:  This signal clears the sliding sums and resets the DC raw measurement data (RMD) pointer and TC time-stamp pointer to FFFF.  This signal typically is used to initialize the DC and TC and get ready for data taking.

Make_Meas   [C22]  “Make Measurement”:  This signal causes the DC to switch integrator channels, latches the previous abort states, digitizes the reading, resets the integrator, calculates the sliding sums, and makes the abort comparisons.  This signal is generated on the TC and typically runs continuously at about once every 20 microseconds.  The sliding sums are made by adding the current reading to the sum and subtracting the oldest reading from the sum.  The 3 sliding sums have 3 oldest data pointers, which point to the data in the RMD that is to be subtracted.  There is also new data pointer, which sets the location in the RMD where the newest reading is stored.    When Make_Meas is generated, the TC also stores the real time in the 64k deep time-stamp buffer.  Make_Meas triggers the AC to query the abort states of the digitizers and to store the results in the 64k deep abort buffer.

Fast_Latch  [C15]: This signal is produced on the TC to tell the DC state machine to latch the current Fast_Sum into a register so the controller can read it.  The TC latches the time stamp as well.  Each of these latch signals also sets a bit in a register and causes the TC to assert the IRQ2 line so that the CC knows to fetch data into its larger circular buffers.

Slow_Latch  [C16]:  This signal is produced on the TC to tell the DC state machine to latch the current Slow_Sum into a register so the controller can read it.  The TC latches the time stamp as well.

VSlow_Latch  [C17]:  This signal is produced on the TC to tell the DC state machine to latch the current Very_Slow_Sum into a register so the controller can read it.  The TC latches the time stamp as well.

Abort_CS(0:5)  [C10:14,27] “Abort Channel select”:  These signals come from the Abort Controller (AC) and are used to poll all the channels to request abort demand conditions.  The DC state machine compares Abort_CS(2:5) with the address switch values to determine if the requested channel is on the board. Abort_CS(0:1) signify channels 0-3 on the board. When selected, each channel’s four abort lines are driven onto the bus’s Abort(0:3) lines.  A channel requesting an abort drives the corresponding line low.

Abort(0:3)  [C5:8] “Abort lines 0,1,2,3”:  These are the abort requests driven by the DC to the AC in order for the AC to determine how many channels are requesting an abort for each of the four abort types.

Abort_Sync [C28] This signal is produced on the TC to trigger the ring-wide serial abort link.  The special abort card at A0 listens for this trigger which is generated every 2 ms and initiates the next data collection sequence on the link.

CA(0:12)  [A1:12] “Control Bus Address”:  These are used to address control and data within the state machine.  These are driven by the CC.

CD(0:7)  [A19:26] “Control Bus Data”: These are the data lines used to read and write data to and from the state machine.

MREQ* [A15] “Memory Request”:  This active-low signal indicates to the DC, AC, and TC that the CC is sending a valid address on the control bus for them to parse.  The eZ80 on the CC uses the same address space for control bus operations as well as memory  operations.  Thus addresses are valid on the bus only if MREQ* is low.

MemRD* [A16] “Memory Read”:  Indicates that the current Control Bus cycle is a read from the bus by the CC.  This is active low and should go down at the same time as MREQ*.

WE* [A17] “Memory Write Enable”:  This is the write strobe for data.  Data should be latched on the rising edge of WE*.  WE* is pulled low by the eZ80 on the CC (or by the control-bus bridge on the TC).

Update_Settings [C29]: The abort thresholds, masks and multiplicities are double buffered on the DC and AC with current values and set values.  After the CC is notified by the TC of an MDAT event, it updates the set values.  When all values have been written, it sets Update Settings to notify the DC and AC to move the set values to be the current values which are then used in abort decisions.

ITBT_Trig [C2] “Injection Turn-by-Turn Trigger”:  Causes the raw data to be written into the 8k ITBT buffer.  Once the buffer is full or the system stops, this data is readable via VME.  The trigger is caused by the TC receiving a programmable TCLK event.

STBT_Trig [C3] “Studies Turn-by-Turn Trigger”: Causes the next 8k raw measurement data to be written into the STBT Buffer. Once the buffer is full or the system stops, this data is readable via VME.  Note: If a STBT trigger happens while an ITBT is in progress, the ITBT is terminated and the STBT operation is initiated.  If either of these is in operation a bit is set in a status register and VME is not allowed to read the TBT data.  The TC sends STBT_Trig after receiving a programmable BSCLK event.

ChOK [C1] “Channel OK”:  Sent by the digitizer to the abort card with the abort data to indicate that the channel is working properly.  If ChOK is low, the AC ignores the data and interrupts the CC.

AIP [C19] “Abort in Progress”:  In normal operation AIP is sent by the TC with a programmable delay after receiving the Abort-in-progress TCLK event ($47).  AIP on the control bus stops all circular buffers and allows the 64k-deep raw data buffers on the DCs and TC to be read.  AIP is latched until reset by external control.  AIP can also be set by-hand in the TC.

IRQ0 [A28] “Interrupt Request 0”:  Timing Card notification to Control Card that a TCLK event has been loaded into the TCLK FIFO.

IRQ1 [A29] “Interrupt Request 1”:  Timing Card notification to Control Card that an MDAT state change has occurred and the new state is ready to be read.

IRQ2 [A30] “Interrupt Request 2”:  Timing Card notification to Control Card that one or more of the Fast, Slow, or Very Slow latches has fired and the data is ready to read and placed in the circular snapshot buffers on the CC

IRQ3 [A31] “Interrupt Request 3”:  Abort Card notification to Control Card that the AC needs servicing, either it has channel NOT OK bits, or is requesting ABORT.

CPU_Detect [C23]:  Pulled low by CC when plugged into the crate.  This disables the control-bus bridge on the TC.

Error [A32]:  Set if any of the cards senses an internal error.  This allows the CC to set a VME interrupt to notify the front-end CPU to send an alarm.  This pin is routed to pin 89 on the eZ80 chip (a.k.a. General Purpose I/O pin PB2).

Osc [C31] Oscillator:  10 MHz from timing card internal oscillator

BusReq [C25] and BusAck [C26]: For future expansion if needed.  These allow for a second master such as a DMA controller to request control of the bus from the eZ80 by pulling the BusReq low, the eZ80 responds by placing all of its control and address lines into a high impedance state and pulling the BusAck low to signal that the requesting device can now use the bus.

Table

2: J2 backplane connector pinout.

| Pin | Row A | Row B | Row C |
| --- | --- | --- | --- |
|  | Function | Write | Read | Function | Function | Write | Read |
| 1 | CA0 | C | DTA | +5V | ChOK | D | A |
| 2 | CA1 | C | DTA | Gnd | ITBT_Trig | T | D |
| 3 | CA2 | C | DTA | Reserved 1 | STBT_Trig | T | D |
| 4 | CA3 | C | DTA | VME-A24 | Abort4 (spare) |  |  |
| 5 | CA4 | C | DTA | VME-A25 | Abort0 | D | A |
| 6 | CA5 | C | DTA | VME-A26 | Abort1 | D | A |
| 7 | CA6 | C | DTA | VME-A27 | Abort2 | D | A |
| 8 | CA7 | C | DTA | VME-A28 | Abort3 | D | A |
| 9 | CA8 | C | DTA | VME-A29 | Gnd |  |  |
| 10 | CA9 | C | DTA | VME-A30 | Abort_CS0 | A | D |
| 11 | CA10 | C | DTA | VME-A31 | Abort_CS1 | A | D |
| 12 | CA11 | C | DTA | Gnd | Abort_CS2 | A | D |
| 13 | CA12 | C | DTA | +5V | Abort_CS3 | A | D |
| 14 | Gnd |  |  | VME-D16 | Abort_CS4 | A | D |
| 15 | MREQ* | C | DTA | VME-D17 | Fast_Latch | T | D |
| 16 | MEMRD* | C | DTA | VME-D18 | Slow_Latch | T | D |
| 17 | WR* | C | DTA | VME-D19 | V SIow_Latch | T | D |
| 18 | Gnd |  |  | VME-D20 | Gnd |  |  |
| 19 | CD0 | CDTA | CDTA | VME-D21 | AIP | T | DA |
| 20 | CD1 | CDTA | CDTA | VME-D22 | Reset DC | T | DTA |
| 21 | CD2 | CDTA | CDTA | VME-D23 | Reset | C | DTA |
| 22 | CD3 | CDTA | CDTA | Gnd | Make_Meas | T | DA |
| 23 | CD4 | CDTA | CDTA | VME-D24 | CPU_Detect | C | T |
| 24 | CD5 | CDTA | CDTA | VME-D25 | Gnd |  |  |
| 25 | CD6 | CDTA | CDTA | VME-D26 | BusReq |  | C |
| 26 | CD7 | CDTA | CDTA | VME-D27 | BusAck | C |  |
| 27 | Gnd |  |  | VME-D28 | Abort_CS5 | A | D |
| 28 | IRQ0 | T | C | VME-D29 | Abort_Sync | T | A |
| 29 | IRQ1 | T | C | VME-D30 | Update_Settings | T | DA |
| 30 | IRQ2 | T | C | VME-D31 | Gnd |  |  |
| 31 | IRQ3 | A | C | Gnd | Osc |  |  |
| 32 | Error | DTA | C | +5V | Spare2 |  |  |

## Standard Board ID Block

The BLM system uses a fairly standard scheme for identifying modules.  Each module has an ID block containing board type, firmware versions, and possibly authors (see Tables 3 and 4).  These ID blocks allow the crate to be mapped out to verify that the correct cards have been installed in each crate.  The ID check also forms a simple test for the basic functioning of each module.  In addition, each module has a serial number encoded in a DS2401 serial device (see Tables 5 and 6).

Table

3: ID block structure.

| Address | Size | Function |
| --- | --- | --- |
| 0000-0003 | 4 numbers | Board ID Number (ASCII) |
| 0004-002F | 44 characters | Board ID Text String (ASCII) |
| 0030-009F | 112 characters | Firmware version numbers, authors, etc |

Table

4: ID block locations for each board.  Note that the DC has 2 locations since the 2 FPGAs are programmed independently and each ID block contains the firmware version for that FPGA.  The CC has a separate ID block for the eZ80 software which is 48 bytes long.

| Board | Address | Board ID Number |
| --- | --- | --- |
| Control Card FPGA eZ80 | NN700000 NN810000 | 0001 |
| Timing Card | NN000000 | 0002 |
| Abort Card | NN000000 | 0003 |
| Digitizer Card Integ FPGA Sums FPGA | N00000 N02100 | 0004 |
| HV Card | NN000000 | 0005 |

Table

5: Word format for the serial number from the DS2401.  The family code is always 01 hex.

| Word 0 | Word 1 | Word 2 | Word 3 |
| --- | --- | --- | --- |
| XX | XX | XX | XX | XX | XX | XX | 01 |
| 8-bit CRC | 48-bit ID number | Family Code |
|  | Board serial number = ( ID mod 1000 ) |  |

Table

6: Location of Silicon IDs.

| Board | Address |
| --- | --- |
| Control Card | NN7000F8 |
| Timing Card | NN000200 |
| Abort Card | NN000200 |
| Digitizer Card | N00200 |
| HV Card | NN000200 |

## Temperature Sensors

The TC, AC, CC, and DCs are each equipped with a digital temperature sensor.  The CC, AC, and TC all use a National Semiconductor LM74 as the temperature sensing device.  The LM74 has 12-bit plus sign temperature resolution (0.0625°C per LSB) while operating over a temperature range of -55°C to +150°C. Temperature data is represented by a 13-bit, two’s complement word with an LSB (Least Significant Bit) equal to 0.0625°C.  The DC uses an Analog Devices TMP03 as the temperature sensing device.  The output of the TMP03 is encoded in the duty cycle of an output pulse which is measured by the FPGA firmware with a LSB of 0.0625°C.  The temperature words for the TC, AC, and DCs are periodically copied by the CC over the control bus and made available in the VME memory of the CC.  To convert to °C, one must multiply the signed value in Table 7 by 0.0625 (1/16 °C).

Table

7: Bit layout of temperature words.  The words are all available on the CC.

| Device | Address | Bits |
| --- | --- | --- |
|  |  | 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| TC | NN800060 | Signed Value | 1 | 1 | 1 |
| AC | NN800062 |  |  |  |  |
| CC | NN7000F2 |  |  |  |  |
| DC | NN800064… | Signed Value |

# Digitizer Card

The digitizer card is the primary data collector in the system.  It includes the following features:

4 loss-monitor channels, 4 dual integrators, 4 16-bit ADCs

4 16-bit DACs for analog (MADC) outputs

Control Bus Interface for setting parameters

512 Kbytes of RAM for storing Raw Measurement Data (RMD)

128 Kbytes of RAM for storing 2 banks of 8k TBT data

A VME interface used primarily for reading raw measurement data when the system is stopped and for reading TBT data when TBT is not in operation.

A 4-bit switch-selectable card number that sets the base address for the Control Bus, Abort Bus and VME.  For digitizer cards on the Control Bus, CA12=0 and CA8:11=Switch for card select.

A state machine that runs the digitizer section and makes the 3 sliding sums for each channel. It also maintains the RMD RAM circular buffer.

An Abort channel select and 4 abort line drivers, these 4 lines are for:

Immediate Loss Abort threshold (few microseconds)

Fast Loss Abort Threshold (few milliseconds)

Slow Loss Abort Threshold (few 10s of milliseconds)

Very Slow Loss Abort Threshold (few seconds)

The reading of the raw measurement data (RMD) is allowed only when the state machine is stopped, and this data is only available via the VME interface.  This is the only data available to VME on this card.  The VME Address lines 20:23 are compared with the 4-bit address switch to determine if the card is being addressed.  VME Address lines 1:2 are channel select, VME address 0 is byte select (only used if reading in byte mode) and VME address lines 3:18 are the 64k RMD pointer.  Note that the RMD is a 64k deep circular buffer used to store raw data; the time stamps for this data are stored on the TC.  In order to see raw measurement data while running, there are two turn-by-turn buffers that are described below in Section 1.5.  VME is also able to read the card ID Block and the board status at 30000.

The control bus address map of the digitizer is described in Tables 8 and 9.  These are the offsets from the control bus base address for data communication with the Control Card.

The response of the digitizer has been calibrated in prototypes.  The least count in the low integration range corresponds to a charge of ~0.02pC or a loss of 0.3µRad.  In the high range, the least count corresponds to a factor of 5 more charge.

## Address Maps

BLM digitizer cards respond to A24D16 and may take nearly all of the A24D16 address space (16 MB) since there may be up to 15 digitizer cards.  Each card uses the addresses N00000-NFFFFE where N is card number 0-E.  The base address is set by switches on the board which determine bits 20-23 [1].  The data are stored as follows as offset from the base address.

| Address | Description |
| --- | --- |
| 00000-00FFE | Standard ID block (Sec. 1.10) |
| 01000 | FPGA Reconfiguration Signal: COM_REG |
| 0100A | FPGA Reconfiguration Signal: DIAG1 |
| 10000-1FFFE | 64kB of Injection TBT Data (8k Turns, Increments) |
| 20000-2FFFE | 64kB of Studies TBT Data (8k Turns, Increments) |
| 30000 | Status Register |
| 80000-FFFFE | Raw Measurement Circular Buffers (64k turns) |

The turn-by-turn linear buffers and raw circular buffers are addressed as follows:

A2:A1 = Channel 0/1/2/3

A18:A3 = Index

The index for the raw buffer is set to 0000 on a reset and increments on each Make_Meas.  It resets when it reaches FFFF.  The current location is indicated by the RMD pointer which is located in a control bus register.

The following table shows the bit definitions for the status register at 30000.

| Bit |
| --- |
| 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
|  |  |  |  |  |  |  |  | Running |  |  |  |  |  | TBT | AIP |

## Control Bus Address Maps

The digitizer card can be accessed from the control bus (CB) with 8 data bits and 13 address lines.  The upper bits of the address lines are used to distinguish the modules in the crate.  Each digitizer card is normally given 256 bytes to be accessed by the control card.

The information accessible via CB is stored in the digitizer card using block memories and registers implemented in the lower (sums03) FPGA.  There are 4 types information accessible via the CB: thresholds, operation settings, sums and control registers.  The thresholds and operation settings are supported with a paging scheme.  The data being accessed by the CB and the data being used for module operation can be different.  The paging scheme allows users to avoid changing thresholds and operation settings while they are being used.  The pages being used and being accessed are changed by writing corresponding control registers.  The sums are latched as if they are double buffered.  However, the sums are actually implemented using a rotating paging scheme.  The users are allowed to take further advantage of this sum implementation in the sense that data from the previous 6 time periods are still accessible via other pages.

The particular addresses of the contents can be assigned in two different mappings: the original mapping and native mapping.  The registers are internally arranged in the digitizer cards using native mapping.  The registers can be remapped via an address translation ROM to the original mapping so that the original test software can still be used.  In the native mapping, registers that are normally read at the same time are grouped together to allow users to write faster and simpler software.  The switching between the mappings is purely a software operation; simply set or reset bit 7 of the FF register (at address 0xFF).  The power on default is the original mapping for backward compatibility with the Visual Basic test software.

The digitizer card also supports both individual and broadcast accesses.  Each digitizer card in the crate can be accessed individually at the assigned card address.  All the cards in the crate can be written simultaneously at the broadcast address.  The broadcast ability allows the cards to change operation state (such as threshold page to be used) together.

### Native Mapping

The native mapping of the digitizer card is shown in Table 8.  The digitizer presents the native register mapping when bit 7 of the FF register is set (=1xxxxxxx_b).

Table

8: Native memory map of the DC.

| Address | Access Type | Register | Mode | Channel |
| --- | --- | --- | --- | --- |
| 00 | R/W | Abort Thresholds | Immediate | 0 |
| 04 |  |  |  | 1 |
| 08 |  |  |  | 2 |
| 0C |  |  |  | 3 |
| 10 |  |  | Fast | 0 |
| 14 |  |  |  | 1 |
| 18 |  |  |  | 2 |
| 1C |  |  |  | 3 |
| 20 |  |  | Slow | 0 |
| 24 |  |  |  | 1 |
| 28 |  |  |  | 2 |
| 2C |  |  |  | 3 |
| 30 |  |  | Very Slow / Integrated | 0 |
| 34 |  |  |  | 1 |
| 38 |  |  |  | 2 |
| 3C |  |  |  | 3 |
| 40 | R/W | Sum Lengths | Immediate | N/A |
| 42 |  |  | Fast | N/A |
| 44 |  |  | Slow | N/A |
| 46 |  |  | Very Slow | N/A |
| 48 | R/W | Constants | 0 |
| 4A |  |  | 1 |
| 4C |  |  | 2 |
| 4E |  |  | 3 |
| 50 | R/W | Integration Squelch Level | 0 |
| 52 |  |  | 1 |
| 54 |  |  | 2 |
| 56 |  |  | 3 |
| 58 | R/W | Pedestal_LO | 0 |
| 5A |  |  | 1 |
| 5C |  |  | 2 |
| 5E |  |  | 3 |
| 60 | R/W | Pedestal_HI | 0 |
| 62 |  |  | 1 |
| 64 |  |  | 2 |
| 66 |  |  | 3 |
| 68 | R/W | Mode Select | 0 |
| 6A |  |  | 1 |
| 6C |  |  | 2 |
| 6E |  |  | 3 |
| 70 | R/W | DAC Out | 0 |
| 72 |  |  | 1 |
| 74 |  |  | 2 |
| 76 |  |  | 3 |
| 78 | R/W | Test DAC | N/A |
| 7A | R/W | MaxDY (=0x80) | N/A |
| 7C | R/W | CIC Length (=124-128) | N/A |
| 7E | R/W | Ped Length (=750-752) | N/A |
| 80 | R | Sums | De-rippled | 0 |
| 84 |  |  |  | 1 |
| 88 |  |  |  | 2 |
| 8C |  |  |  | 3 |
| 90 | R |  | Fast | 0 |
| 94 |  |  |  | 1 |
| 98 |  |  |  | 2 |
| 9C |  |  |  | 3 |
| A0 | R |  | Slow | 0 |
| A4 |  |  |  | 1 |
| A8 |  |  |  | 2 |
| AC |  |  |  | 3 |
| B0 | R |  | Very Slow | 0 |
| B4 |  |  |  | 1 |
| B8 |  |  |  | 2 |
| BC |  |  |  | 3 |
| C0 | R | Integration / CIC_WF | 0 |
| C8 |  |  | 1 |
| D0 |  |  | 2 |
| D8 |  |  | 3 |
| E0 | R/W | THR U page | N/A |
| E1 | R/W | THR C page | N/A |
| E2 | R/W | Constant U page | N/A |
| E3 | R/W | Constant C page | N/A |
| E4 | R/W | Sum U page | N/A |
| E5 | R/W | Sum F page | N/A |
| E6 | R/W | Sum S page | N/A |
| E7 | R/W | Sum V page | N/A |
| E8 | N/A | Unused | N/A |
| F6 | R | Circular Buffer PT | N/A |
| F8 | R/W | FPGA Ctrl Reg. | N/A |
| FA | N/A | Unused | N/A |
| FC | R | Temperature | N/A |
| FE | N/A | Unused | N/A |
| FF | R/W | FF reg | N/A |

The thresholds from addresses 0x00 to 0x3F are 32-bit numbers each occupying 4 bytes and stored in Little Endian format, i.e. the LSB of the threshold occupies the lowest address.  There are 64 pages of thresholds.  The pages being used and being accessed via CB are represented by the lower 6 bits of the “THR U page” (at 0xE0) and the “THR C page” (at 0xE1) registers, respectively.

The addresses from 0x40 to 0x7F contain the settings for the module operation.  There are 8 pages for this class of information.  The pages being used and being accessed via CB are represented by the lower 3 bits of the “Constant U page” (at 0xE2) and the “Constant C page” (at 0xE3) registers, respectively.

The addresses from 0x80 to 0xDF are assigned for sums.  The fast, slow and very slow sliding sums are 32-bit integers occupying 4 bytes each in Little Endian format.  The immediate measurements are implemented as sliding sums with a sum length = 1.  On power on, the sum length (at 0x40 and 0x41) for immediate is automatically preset to 1, i.e., (0x40) = 0x01 and (0x41) = 0x00.  The users need not set the immediate sum length for regular operation.  The immediate sums are stored internally but not output.  Instead, the de-rippled sums are output at addresses 0x80 to 0x8F, 4 bytes for each channel.  The de-rippled sums are for display only and are not used to compare with thresholds.  The integrations are 64-bit integers and are readout from 0xC0 to 0xDF in normal operation.  Two other numbers useful for debugging, the baseline waveform, WF, and the CIC sums can be readout form the integration locations if bit 4 of the FF register is set to 1.  The CIC sum is a sliding sum of a sliding sum.  It filters out higher frequency noise better than sliding sum.  The baseline waveform is stored in internal memory and used in de-ripple process.  It is a sample of the noise present when there are no losses.

There are 8 pages for the sums.  The pages being used and being accessed via CB are specified by the lower 3 bits of the “Sum U page” (at 0xE4), the “Sum F page” (at 0xE5), the “Sum S page” (at 0xE6) and the “Sum V page” (at 0xE7) registers, respectively.  All the page registers are reset on power up and can be read and written by the Controller Card.  When the latch sum signals for fast, slow and/or very slow become active, the sum F, S and/or V pages are updated to the current value of the Sum U page and then the Sum U page increases by 1.  So the corresponding pages for different sum types point to the last latched values unless more than 8 latch commands have been issued before the sums are readout.  Bit 7 of the sum F, S and V page registers are used to prevent the registers from being updated by the latch sum signals.  The user may set the bit 7 of the page registers before reading the sums so that the contents are not corrupted by the new latch sum signals.

The contents of de-rippled sums and the integrations are pointed by the Sum F page and the Sum V page registers, respectively.

### Original Mapping

The original mapping is shown in Table 9.  The digitizer presents the original mapping when bit 7 of the FF register is reset (=0xxxxxxx_b).  The original mapping is also the power on default condition since the original test software doesn’t know about the two mappings.

Table

9: Original memory map of the DC.  This mode is used during bridge program use.

| Address | Access Type | Register | Mode | Channel |
| --- | --- | --- | --- | --- |
| 00 | R/W | Abort Thresholds | Fast | 0 |
| 04 |  |  | Slow | 0 |
| 08 |  |  | Very Slow / Integrated | 0 |
| 0C |  |  | Immediate | 0 |
| 10 | R | Sums | Fast | 0 |
| 14 |  |  | Slow | 0 |
| 18 |  |  | Very Slow | 0 |
| 1C | R/W | Mode Select | 0 |
| 1E | R/W | DAC Out | 0 |
| 20 | R | Sum | De-rippled | 0 |
| 22 | R/W | Pedestal_HI | 0 |
| 24 | R/W | Constant | 0 |
| 26 | R/W | Pedestal_LO | 0 |
| 28 | R | Integration / CIC_WF | 0 |
| 30 | R/W | Threshold | Fast | 1 |
| 34 |  |  | Slow | 1 |
| 38 |  |  | Very Slow / Integrated | 1 |
| 3C |  |  | Immediate | 1 |
| 40 | R | Sums | Fast | 1 |
| 44 |  |  | Slow | 1 |
| 48 |  |  | Very Slow | 1 |
| 4C | R/W | Mode Select | 1 |
| 4E | R/W | DAC Out | 1 |
| 50 | R | Sum | De-rippled | 1 |
| 52 | R/W | Pedestal_HI | 1 |
| 54 | R/W | Constant | 1 |
| 56 | R/W | Pedestal_LO | 1 |
| 58 | R | Integration / CIC_WF | 1 |
| 60 | R/W | Threshold | Fast | 2 |
| 64 |  |  | Slow | 2 |
| 68 |  |  | Very Slow / Integrated | 2 |
| 6C |  |  | Immediate | 2 |
| 70 | R | Sums | Fast | 2 |
| 74 |  |  | Slow | 2 |
| 78 |  |  | Very Slow | 2 |
| 7C | R/W | Mode Select | 2 |
| 7E | R/W | DAC Out | 2 |
| 80 | R | Sum | De-rippled | 2 |
| 82 | R/W | Pedestal_HI | 2 |
| 84 | R/W | Constant | 2 |
| 86 | R/W | Pedestal_LO | 2 |
| 88 | R | Integration / CIC_WF | 2 |
| 90 | R/W | Threshold | Fast | 3 |
| 94 |  |  | Slow | 3 |
| 98 |  |  | Very Slow / Integrated | 3 |
| 9C |  |  | Immediate | 3 |
| A0 | R | Sums | Fast | 3 |
| A4 |  |  | Slow | 3 |
| A8 |  |  | Very Slow | 3 |
| AC | R/W | Mode Select | 3 |
| AE | R/W | DAC Out | 3 |
| B0 | R | Sum | De-rippled | 3 |
| B2 | R/W | Pedestal_HI | 3 |
| B4 | R/W | Constant | 3 |
| B6 | R/W | Pedestal_LO | 3 |
| B8 | R | Integration / CIC_WF | 3 |
| C0 | R | Sums | Immediate | 0 |
| C4 |  |  |  | 1 |
| C8 |  |  |  | 2 |
| CC |  |  |  | 3 |
| D0 | R/W | Integration Squelch Level | 0 |
| D2 |  |  | 1 |
| D4 |  |  | 2 |
| D6 |  |  | 3 |
| DA | R/W | MaxDY (=62) | N/A |
| DC | R/W | CIC Length (=124-128) | N/A |
| E0 | R/W | THR U page | N/A |
| E1 | R/W | THR C page | N/A |
| E2 | R/W | Constant U page | N/A |
| E3 | R/W | Constant C page | N/A |
| E4 | R/W | Sum U page | N/A |
| E5 | R/W | Sum F page | N/A |
| E6 | R/W | Sum S page | N/A |
| E7 | R/W | Sum V page | N/A |
| EA | R/W | Ped Length (=750-752) | N/A |
| EC | R/W | Sum Length | Immediate | N/A |
| F0 |  |  | Fast | N/A |
| F2 |  |  | Slow | N/A |
| F4 |  |  | Very Slow | N/A |
| F6 | R | Circular Buffer PT | N/A |
| F8 | R/W | FPGA Ctrl Reg. | N/A |
| FA | R/W | Test DAC | N/A |
| FC | R | Temperature | N/A |
| FE | N/A | Unused | N/A |
| FF | R/W | FF reg |  | N/A |

### Broadcast Access

The CB address space has 13 bits.  The digitizer card individual access address is 0x0NXX, where N = 0 to E is the board switch setting and XX = 00 to FF is the byte address.

The digitizer card broadcast access uses the address space outside the above.  The address is 0x13XX which is unused by any card.  In the broadcast access address space, all digitizer cards are selected as being individually accessed.

Writing to a location in 0x13XX causes the addressed byte in all cards being written with the same value.

Reading in the 0x13XX space is not recommended since the bus conflict may be resulted in.  Reading-after-writing is allowed since the anticipated values from all cards are identical but may not be very useful.

The broadcast access is essentially used to change the operating condition of all cards simultaneously.  For example, one may change the threshold page being used when a machine state change is needed.

## Control Bus Registers

### FF Register

The FF register is always at address 0xFF in both native and original mapping.

| Bits |
| --- |
| 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| Mapping | Ch OK |  | Out De-Bug |  | Do Sums | Clear Sums |  |

Bits 1 and 2 are used to initialize sum operations without a timing card.  Turning bit 1 high and then low causes the digitizer card to initialize the summing operation.  Turning bit 2 high and then low causes the digitizer card to accumulate sliding sums and integrations.

The summing operation via toggling the FF register bits 1 and 2 is usually for debugging purposes and the input data are always the constants stored in addresses 0x48 to 0x4F.  The bit 0 in the FPGA Control Register should be set in order to use constant as the summing and integration input.

Bit 4 specifies that the baseline waveform and the CIC sums should be made available in the space normally used for integrated values.

Bit 6 Enable/Disable the ChOK line on the Abort Bus backplane.  The channel OK line is not currently connected to anything meaningful, so this bit should be 0.

Bit 7 of the FF register is used to control the mapping for the Control Bus.  In this document, native mapping is used when referring to Digitizer Card Control Bus addresses, which requires bit 7 being set.  To use the original mapping, bit 7 is cleared.

### FPGA Control Register

The FPGA Control Register is a 16-bit register at address 0xF8 and 0xF9 in both native and original mapping.

| Bits |
| --- |
| 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| NSKIP16 | Decimation Offset | Use constant |

Bit 0 set means constants loaded at 0x48 to 0x4F (in native mapping) are used as input values for computing sliding sums and integrations.  In normal operation using the real ADC input, bit 0 is clear.

Bits 1 to 7 specify an offset to the default decimation constant used in the noise filter. The default value of the decimation constant is 22336.  Bit 7 specifies the sign of the offset (1=negative), and bits 1..6 specify a value which gets multiplied by 32 to form the actual offset.  So the final decimation constant is 22336 + sign(bit[7]) * bit[6..1] * 32.  See reference [6] for details about the noise filter.

Bits 8 to 15 represent a number NSKIP16 that is loaded by the user.  The digitizer card ignores the first 16*NSKIP16 measurements before starting summing and integration processes.

### Mode Select Registers

The Mode Select Registers are 16-bit words at 0x68 to 0x6F (in native mapping).  Each channel is controlled by its own Mode Select Register.  The bits of the registers are assigned as shown in the following table:

| Bits |
| --- |
| 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
|  |  |  |  |  |  |  | Squelch | /INP_ON | MADC Function Select [2:0] | Use Integration Mode | Test Mode On | Enable Long Time Constant | Enable High Range |

Bit 0 specifies whether to use low or high range which causes the integration capacitor for the integrator chips to be set to either 100 pF or 500 pF; a difference of a factor of 5 in the least count and range for the data.

Bit 1 enables the long time constant for the integration circuit.  The normal series resistance of the input to the integrator is about 2 kΩ.  When this control bit is set, an additional 15 kΩ resistor is included in series, increasing the effective time constant of the integrator.

Bit 2 enables test mode, whereby the on-board DAC is used to send a fixed current to the input of the integrator.  The value of the DAC is set from the Test DAC register.

When bit 3 is turned on, the 32-bit very slow abort thresholds will be compared with bits 16 to 48 of the integration values.  When this bit is clear, the integrations will still be accumulated, but they are not compared with the thresholds.

Bits 4 to 6 indicate what should be put on the analog outputs of the DC.  The analog output for each channel can provide up to 8 different signals derived from the raw data.  These are specified in the firmware.  The output choices are:

| 3-bit Value | Description |
| --- | --- |
| 0 | Value from MADC Manual Setting |
| 1 | Raw measurement value |
| 2 | Integrated current value (if enabled) |
| 3-7 | Unused |

Bit 7 opens or closes the integrator input switch.  A value of zero closes the input switch (normal operation), and a value of 1 opens it.

Bit 8 “Squelch” turns on the squelch operation for integration mode.  Inputs lower than the sum of the pedestal plus squelch level are considered noise and are not accumulated into the integration.  (See Section 2.11 for a description of the squelch level)

## Remote FPGA Programming

The two FPGA devices, Integ03 and Sums03 on the digitizer card are each configured from an EEPROM device AT17LV002A at power up.  The contents of the EEPROM can be both read and rewritten from VME.  Once the EEPROM is successfully rewritten, the FPGA will be configured with new firmware at the next power up.

Write access to the EEPROM signals is via a command register COM_REG at N01000 and the read access is from a diagnostic register DIAG1 at N0100A.  Tables 10 and 11 list the relevant bits of the reprogramming registers.  The remaining bits are used for other purposes and should be set to 0 when writing to COM_REG.

Table

10: Bit definitions for the COM_REG reprogramming register showing which EEPROM pins are connected to which bits.

| COM_REG (N01000) |
| --- |
| Bit 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7-0 |
| Unused | DATA0_V1 | DCLK_V1 | SER_EN1 | Unused | DATA0_V | DCLK_V | SER_EN | Reserved |
|  | To Sums03 FPGA |  | To Integ03 FPGA |  |

Table

11: Bit definitions for the DIAG1 reprogramming register.

| DIAG1 (N0100A) |
| --- |
| Bit 15-6 | 5 | 4 | 3-0 |
| Reserved | DATA0_in1 | DATA0_in | Reserved |
|  | From Sums03 FPGA | From Integ03 FPGA |  |

The signals to the EEPROM pins are set by writing to the COM_REG register.  To set a signal high or low, write 1 or 0 to the given bit in the COM_REG.  A signal stays at the level after being written until next time an opposite level is written.  The level of DATA can be read from DIAG1 register bit 4 and 5 as shown above.

Not that there is an inverter each for the SER_EN and SER_EN1 signals.  They are active high from program view while the inverted versions of them presented at the device pins are active low.  After power-up, these register bits are 0, which disables access to the EEPROM devices.  To read or write an EEPROM, set SER_EN (or SER_EN1) to 1.  The polarities of the DATA0 and the DCLK signals are non-inverted.

Check the card for the following to ensure correct reconfiguration:

- Jumpers J53, J54 and J66 must be installed for Integ03 FPGA reconfiguration.

- Jumpers J51, J52 and J64 must be installed for Sums03 FPGA reconfiguration.

## Noise Filtering

Because of the EMF pickup encountered in the tunnels, the fast sums in particular can display a rather large noise of up to 200 mR/s.  To combat this from a display perspective, an additional sum type was created that consists of a sliding sum of a sliding sum where the sum length is picked to be approximately 360 Hz.  The double sliding sum, a.k.a. Cascaded Integration Comb filter (CIC), provides better rejection of the 360 Hz multiples when the sampling frequency is not exactly aligned with 360 Hz.  In particular, during acceleration, the sampling frequency, which is the turn frequency, is changing by small amounts.  In addition, instead of just an average pedestal subtraction, the actual shape of the noise (taken when there are no losses) is subtracted from the CIC sum.  This last step removes 60 Hz and multiples.  This filtering process is described in detail in [6].

## Integration Operation

The digitizer card computes sliding sums and integrations triggered by the Make_Meas signal.  In typical operation, the signal is generated by the TC about every 21 μs.  The integrations are 64-bit long integers readable from the control bus at 8 bytes each.

The integration process begins after the sums are reset.  The first set of measurement points can be ignored and the number of points to be ignored is defined by users.  Then the pedestals of the 4 channels are accumulated.  The pedestals are stored once the accumulation is done.  In the integration mode, very slow sliding sum for each channel is used as input data to the squelch checking and the integration accumulation.

After the pedestals become available, the very slow sliding sums are used to compare with the pedestals plus the squelch levels to see if valid data is present or not.  If the very slow sum of a channel is higher than the pedestal plus the squelch level, the pedestal is subtracted from the very slow sum and the result is added to the integration.  A sufficiently large initial value is given to the integrations which prevents the integrations from drifting negative when the pedestals fluctuate too high.

In integration mode, bits 16-48 of the 64-bit integration value are compared with the 32-bit very slow thresholds every 21 μs. The details are described below.

## Sum Lengths

The lengths of the sliding sums are 16-bit integers.  In addition to the fast, slow and very slow sliding sums, the immediate measurement for each channel is also implemented as a sliding sum with length = 1.  All these sums are implemented identically and their lengths can be arbitrarily defined by the users.

The pedestal sum length at 0x7E is defined similarly.  The length is used to calculate the pedestals at the beginning of the run.  In integration mode, there must be a fixed ratio of 16 between the very slow sum length, which is the input to the integration, and the pedestal sum length.  Typically, the very slow sum length is chosen to be about 47 and the pedestal length to be about 752.  Note that the users are allowed to choose very slow sum length freely.  For example, the very slow sum length can be 55 and the pedestal sum length 880.

## Pedestals

All digitizer measurements are positive integers.  In order to avoid bias from negative noise signals, a fixed current is added to the input of each integrator circuit.  This presents a pedestal for each measurement.  As discussed in Section 2.4, a pedestal is subtracted in making the integrated sum.  However, for the raw measurement data or sliding sum data, pedestals must be subtracted external to the digitizer.  In order to minimize the load on the control card, these subtractions are done by the front-end CPU.  Pedestals are taken at the beginning of a measurement cycle, after Reset_DC is asserted.  As mentioned earlier, the first 16*NSKIP16 measurements are ignored before the pedestals are calculated.  The calculated pedestals are placed into the RAM block interfacing with the control bus and can be read by the CC.  The pedestals can be changed by the users also, if necessary.  But to avoid being overwritten, they should only be changed after the accumulation of the pedestals at the beginning of each run.

Abort thresholds for each channel are stored in the ACNet database and correspond to actual losses.  The front-end corrects the thresholds for the pedestals before loading them to the Control Card.  While care was taken to minimize analog noise, some residual 60Hz and 720Hz noise remains.  Therefore, the length of the averaging period used to determine the pedestal should be an integral multiple of 1/60 second.  For the Tevatron, this is 795 cycles; while for the main injector, it is 748 cycles at 8 GeV with Make_Meas running at half the AA marker frequency.

Because the on-board pedestal must be an integer, it is possible to see a slow drift in the integrated measurement with no losses.  In the low (high) range this is 14 mRad/s (70 mRad/s).

## Thresholds

Pre-loaded sets of thresholds are organized as pages and there are 64 pages total in a digitizer card.  The users are allowed to select a particular page to be used at any time.  A clean way to change pages is via a control bus broadcast.

## Saturation Protection

In order to be sensitive to extremely large losses, the Digitizer protects the integrator from becoming saturated.  During the integration phase of the cycle, the output voltage is checked by a comparator.  If it reaches approximately 95% of full scale for the ADC, the input current is shut off, and the remaining charge is stored on the cable.  Thus for large losses, the leading edge time is preserved; however, the raw measurement will be near full scale until all charge from the loss has been collected.  Those measurements will be included as normal in the sliding sums.  In this way, by summing over measurements, it is possible to record the full energy deposition.

## Calculations

This section presents the exact mathematical prescriptions for the various quantities calculated by the DC.  It also provides conversions from physical values to digitizer values.

Table

12: Table of BLM quantities.  Highlighted quantities are calculations done by the DC.  Other quantities are calculated by the front end or other entitiy.

| Quantity | Registers | Description |
| --- | --- | --- |
| NS | F9 | ( number of Make_Meas to skip before summing ) / 16 |
| LPED | 7E-7F | Pedestal sum length (must be 16 * LVS for integration mode to work) |
| T | --- | Type of sum (F=Fast, S=Slow, VS=Very Slow) |
| LT | 40-47 | Sum length for type T |
| LC | 7C-7D | CIC length (used in the derippled sum) |
| Xi | --- | Raw data measurement (this quantity is only available in the original mapping) |
| SiT | 90-BF | Sliding sum |
| SiD | 80-8F | Derippled sum |
| SjC | C0-DF† | CIC sliding sum which forms part of the derippled sum |
| P | 58-67 | Pedestal |
| Wj-j0 | C0-DF† | Baseline waveform for subtraction in derippled output |
| q | --- | Squelch level in units of # of sigma above noise RMS |
| σ | --- | Raw data measurement RMS |
| Q | 50-57 | Squelch level in units of 16 * the very slow sum |
| Yi | C0-DF | Integrated value on Digitizer Card (64 bits) |
| Zi | --- | Integrated value in Control Card circular buffer |
| AiT | --- | Average pedestal subtracted measurement using the specified sum in units of mR/s |
| AiC | --- | Pedestal subtracted, derippled measurement in units of mR/s |
| Ii | --- | Real integrated value in units of mR |
| PI | --- | Integration pedestal (only really needed if input switches are opened) |
| IThresh | 30-3F | Abort threshold for integrations in ADC counts |
| ATThresh | 00-2F | Abort threshold for sum type T in ADC counts |
| ILim | --- | Abort threshold for integrations in mR |
| ATLim | --- | Abort threshold for sum type T in mR/s |

† These registers normally contain the integrated values.  The baseline waveform and the CIC sum is available if bit 4 of the FF register is set.

The pedestal starts after a delay of (16*NS) Make_Meas cycles to give the analog portion of the board time to settle

The derippled baseline waveform is collected from data when there are no losses

The sliding sum is the sum of the previous LT raw data measurements

The DeRippled sum, SjD, is a sum of sums

minus the baseline waveform plus the average pedestal

The squelch level that must be downloaded to the digitizer is ~16 times the very slow sum noise level times the number of sigma desired (it is 16*17 rather than 16*16 because of the contribution of the pedestal)

To determine the squelch level σ from a fast time plot of the fast sum, look at the plot and determine the level, R, where you’d like to place the squelch. Then use the following formula.

The integral starts accumulating after the delay and the pedestal measurement and has the slow sum added every Make_Meas cycle.  The integral also has an offset of 128 Meg added to it to prevent negative numbers on the digitizer.  The U( ) function is the unit step function ( U(x) = 0 if x< 0 and = 1 if x>0 )

The integration value is stored in the digitizer as a 64-bit number.  The control card reads out bits 16-47 and puts that in the very slow sum buffer.

In the front end, the sums must be turned into averages and have pedestals subtracted and conversion factors applied.  For the Fast, Slow, and Very Slow,

whereas for the DeRippled,

To obtain a true integrated value, the very slow sum buffer value must have the 16 bit shift undone along with the factor of 16 and the over counting by LVS.  Then the value can be converted to mR.  The pedestal, PI, is either 227/216, or a measured value.  A measured value is needed if the input switch is opened during pedestals.

The abort thresholds are simply the inverse of the AT and I.

### Bit Alignment

As can be see above, divisions are not actually taken when the variables are to be interpreted as averages such as pedestals and very slow sliding sums.  The values are scaled according the corresponding sum lengths.

In the FPGA implementation the variables are aligned based on the following picture:

The pedestals and sliding sums are accumulated aligned with the raw measurement data.  The squelch level is added to the pedestal with bit 0 aligned.  Therefore, it is possible set squelch level with fractional part.

When the very slow sliding sums are operated with pedestals and squelch levels, the sliding sums are shifted up for 4 bits equivalent to multiplying by 16.

The integration is aligned with the pedestals.  Therefore, dividing the integration value by LPED scales it back to the raw measurement data unit.

The thresholds are aligned to the bit 16 to 47 of the integration.

# Timing Card

The timing card (TC) is the primary controller for data collection in the BLM system.  It includes the following features:

Control Bus Interface for setting parameters

TCLK decoder
Used to receive encoded TCLK events.  One of several actions can be taken upon receipt of a selected TCLK.  TCLK events which are written to the FIFO cause IRQ0 to be asserted. A 1 MHz clock is derived from this clock.  In locations without a TCLK, an optional 10 MHz oscillator is provided.  The list of chosen TCLK events is downloaded through the control bus and stored in RAM.

BSCLK decoder
Used to receive the encoded Beam-Sync Clock events such as $AA or TBT trigger events.   As with the TCLK decoder, this one also features several actions that can be taken for each BSCLK event.  The main use for this decoder is in receiving the AA marker which normally feeds the Make_Meas signal logic.

MDAT decoder
Used to receive machine state information from the MDAT system.  Different machine states require different abort thresholds.  Receiving an MDAT event causes IRQ1 to be asserted.  The MDAT frames of interest are set in the firmware and are $12 for the Tevatron and $56 for the Main Injector.  An MDAT event is defined as a change in the value of the relevant frame.  Only the lowest 7 bits of the frame are considered.  The upper bit is set to 1 if the frame was a MI frame and 0 if it was a TEV frame.

Real-time clock
The 32-bit Unix time in seconds is incremented by the TCLK 1 Hz event ($8F).  A 24-bit microseconds counter increments from each 1 Hz event.  These time values are used to mark the digitizer raw and sum data.

512 Kbytes of RAM for storing 64k of Raw Measurement Data Time Stamps.  These are stored as 8 bits of state, 32 bits Unix time in seconds, and 24 bits of microseconds.

128 Kbytes of RAM for storing 2 sets of 8k worth of TBT time stamps.

A VME interface used for reading the Raw Data Time Stamps when the system is stopped and for reading TBT Time Stamps when TBT is not in operation.

Make_Meas backplane output
Make_Meas can be generated from either of two sources:

The1 MHz clock with an 8-bit divisor

The AA marker from BSCLK or the AA marker / 2 for MI.

The Make_Meas signal also latches the 64-bit time-stamp data and writes it into the time stamp RAM array and increments the RAM Pointer.  The Reset_DC signal resets this pointer.

Divisors of the Make_Meas signal used to make the Fast_Latch, Slow_Latch, and Very_Slow_Latch signals. The divisors are 16 bit numbers.  These latch signals set a bit in a register and cause IRQ2 to be asserted so that the CC knows to read the latched data from the digitizers.  These three latch signals also latch the current time stamp into three registers so that the time stamp can also be read.

Writeable registers to force the creation of Reset_DC, AIP, Clear AIP etc

Front Panel lemo inputs which also make Reset_DC, AIP, and Clear AIP

Abort-in-Progress (AIP) signal
AIP being asserted causes the state machine to stop making all control pulses: Make_Meas, xxxx_Latch, etc., effectively freezing all buffers in the system.  AIP is a latched signal cleared by Clear_AIP and by Reset.

ITBT Trigger
This line is driven onto the backplane in response to a TCLK Injection clock event, it causes the 8k Injection TBT buffer to be filled on the digitizer cards as well as the injection TBT time stamps to be filled on the TC.

STBT Trigger
This line is driven onto the backplane in response to a BSCLK Studies clock event, it causes the 8k Studies TBT buffer to be filled on the digitizer cards as well as the Studies TBT time stamps to be filled on the TC.

The control bus registers for the Timing Card are described in Table 13.  The Timing Card is controlled by signals on the control bus and by TCLK, BSCLK, and MDAT events.

The reading of the Raw Measurement Time Stamp circular buffer is allowed only when the state machine is stopped, i.e. AIP is true.  The TBT time stamp buffers can be read whenever a turn study is not in progress.  The VME address map is described in Section 3.1 below.

The TC also includes a VME to Control Bus Bridge to allow for setting up the DC and TC hardware in situations where a CC module is not present. VME Address line 20 set to 1 indicates that the Bridge is being accessed.  VME Address lines 1:13 are copied to Control Bus address lines 0:12 and control bus data lines 0:7 are connected to VME data lines 0:7.  The bridge must also use appropriate VME signals to create the following Control Bus signals: MEMRD*, MREQ*, and WE*.  The TC will pull Bus line CPU_Detect high.  If this line is sensed low, the bridge will be disabled since there is a CC present.

To set the seconds counter, one writes the time setting to the appropriate registers and then sets the register to update the time setting on the next $8F TCLK event.  The TC resets that register when the time has been updated.  When using a host CPU to set the time on the Control Card, there is a similar feature whereby one sets the time to use in the update and sets a register that the CC polls periodically to see if there is an updated setting to pass to the TC.

## Address Maps

BLM Timing Cards respond to A32D16 and occupy 2Mb.  Each card uses the addresses NN000000-NN1FFFFE where NN is the base address.  The base address is set by switches on the board which determine bits 24-31.  The data are stored as follows as offset from the base address.

| Address | Description |
| --- | --- |
| 000000-000FFE | Standard ID Block (see Sec. 1.10) |
| 00A000 | JTAG programming register |
| 00A002 | JTAG enable register |
| 010000-01FFFE | 64KB of TBT Data (8k Turns) Injection, linear buffer |
| 020000-02FFFE | 64KB of TBT Data (8k Turns) Studies, linear buffer |
| 030000 | Status Register |
| 080000-0FFFFE | 512KB of TBT Data (64k Turns), Circular buffer |
| 100000-103FFE | VME to Control Bus Bridge |

The following table shows the bit definitions for the status register at 30000.

| Bit |
| --- |
| 15 | 14 | 13 | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
| 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | Running |  |  |  |  |  | TBT | AIP |

The time stamp buffers are addressed as follows

A2:A1 = Time Stamp Words (0-3)

A18:A3 = Turns number

The structure of the time stamp data is:

| Word | Most Significant Byte | Least Significant Byte |
| --- | --- | --- |
| 0 | Abort State | Microseconds, byte 2 |
| 1 | Microseconds, byte 1 | Microseconds, byte 0 |
| 2 | Unix Time, byte 3 | Unix Time, byte 2 |
| 3 | Unix Time, byte 1 | Unix Time, byte 0 |

For the 64k turns of raw measurement time stamps, the index increments from 0000 to FFFF and then wraps.  This data can be read only when the system is stopped (i.e. AIP is asserted).  The timestamps for turn-by-turn data may not be read during a turn study.

When using the control bus bridge VME A13:1 map to CA12:0 and A19:14 are ignored.  The 8-bit Control Bus data is on VME D7:0.

Table

13: Timing Card Control Bus registers.

| Function | Access | Address (Hex) | Bytes | Bits |
| --- | --- | --- | --- | --- |
| Setup |
| CSR | R/W | 00 | 2 | 16 |
| Abort State | R/W | 02 | 1 | 8 |
| TBT Status (bit 4) | R | 03 | 1 | 8 |
| Seconds Time Clock setting | W | 04 | 4 | 32 |
| Set Clock on next ($8F) | W | 08 | 1 | 0 |
| Update Settings on next Make_Meas | W | 09 | 1 | 0 |
| Control |
| Make_Meas Divisor | R/W | 10 | 1 | 8 |
| Fast Sum Length | R/W | 12 | 2 | 16 |
| Slow Sum Length | R/W | 14 | 2 | 16 |
| Very Slow Sum Length | R/W | 16 | 2 | 16 |
| Raw Data Time Stamp Pointer | R | 20 | 2 | 16 |
| TCLK Event FIFO | R/W | 30 | 1 | 8 |
| TCLK Event FIFO Status | R/W | 31 | 1 | 8 |
| BSCLK Event FIFO | R/W | 32 | 1 | 8 |
| BSCLK Event FIFO Status | R/W | 33 | 1 | 8 |
| MDAT Event FIFO | R | 34 | 1 | 8 |
| MDAT Event FIFO Status | R/W | 35 | 1 | 8 |
| TCLK RAM Pointer | R/W | 36 | 1 | 8 |
| TCLK RAM Data | R/W | 37 | 1 | 4 |
| BSCLK RAM Pointer | R/W | 38 | 1 | 8 |
| BSCLK RAM Data | R/W | 39 | 1 | 3 |
| Clocks/Timing |
| Fast Latch Time Stamp | R | 40 | 8 | 64 |
| Slow Latch Time Stamp | R | 48 | 8 | 64 |
| Very Slow Latch Time Stamp | R | 50 | 8 | 64 |
| General |
| Force Clear | W | F0 | 1 | 0 |
| Force Start (Reset DC) | W | Fl | 1 | 0 |
| Force Stop | W | F2 | 1 | 0 |
| Force ClkReset ($8F) | W | F3 | 1 | 0 |
| Force AIP | W | F4 | 1 | 0 |
| Force Sum Latches | W | F5 | 1 | 0 |
| Force ITBT Trigger | W | F6 | 1 | 0 |
| Force STBT Trigger | W | F7 | 1 | 0 |
| IRQ2 Status Reg Fast | R/W=Clr | F8 | 1 | 0 |
| IRQ2 Status Reg Slow | R/W=Clr | F9 | 1 | 0 |
| IRQ2 Status Reg Very Slow | R/W=Clr | FA | 1 | 0 |
| Diagnostics |
| Board Temperature | R | FE | 2 | 16 |

## Control Bus Registers

### CSR

| Bits |
| --- |
| 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |
|  |  |  |  |  | Make_Meas source |

Bits 0…2 specify the source of Make_Meas.

| 0 | Use MI AA marker ( 2 |
| --- | --- |
| 3 | Use 1 MHz internal oscillator |
| 4 | Use TEV AA marker |

### Make_Meas Divisor

This specifies the divisor to use if the 1 MHz internal oscillator is selected in the CSR register.

### Event FIFO Status

| Bit | TCLK | BSCLK | MDAT |
| --- | --- | --- | --- |
| 0 | FIFO Empty | FIFO Empty | FIFO Empty |
| 1 | FIFO Full | FIFO Full | FIFO Full |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  | MDAT decode error |
| 5 |  |  |  |
| 6 |  |  |  |
| 7 |  |  |  |

### Timestamps

| Byte | Description |
| --- | --- |
| 0 | Microseconds Low Byte |
| 1 | Microseconds Middle Byte |
| 2 | Microseconds High Byte |
| 3 | State |
| 4 | Seconds Low Byte |
| 5 | Seconds Middle Byte |
| 6 | Seconds More Middle Byte |
| 7 | Seconds High Byte |

## Remote FPGA Programming

The FPGAs are normally programmed by connecting a cable called a “ByteBlaster” between the parallel port of a computer and the JTAG connector on the front panel of the card, and downloading the proper programming file using Altera’s Quartus program into the card’s configuration device.  The configuration device holds the FPGA programming bitstream in non-volatile memory, and programs the FPGAs whenever the board is powered-up.  An alternative method of loading the configuration device exists using four bits of a VME register to emulate the four bits of the JTAG lines.  Table 14 shows the bit assignments.

Table

14: VME register NN00A000 for loading the configuration device.

| Bit | JTAG Function |
| --- | --- |
| 0 | TCK |
| 1 | TMS |
| 6 | TDI |
| 7 | TDO |

This programming register is at NN00A000.  As a precaution, in order for this programming operation to occur, a special code must be written to register NN00A002.  Currently the only requirement in this special code is that a 1 be written to bit 0.  The idea is obviously to ensure that no random or erroneous VME operations corrupt the FPGA configurations.  The software routines for carrying out the remote programming are part of the front end code collection.

## Clock Decoders

### Introduction

The timing cards contain decoders for TCLK, BSCLK, and MDAT.  Both TCLK and BSCLK are programmable with respect to which clock events are acted upon.  It is the responsibility of the Control Card to load the TC with the proper list of clock events to use.  The MDAT decoder is hardwired in firmware to listen to two MDAT frames, one for the TEV and one for the MI.

### Tevatron Clock (TCLK) Decoder

The following is a list of Control Bus Addresses and functions involved in setting up the TCLK Clock Event Decoder RAM.  The RAM is 256 x 4 bits where each 4 bit value is associated with clock event (entries 0-255) and each bit specifies an action to be taken when that particular clock event happens.

Table

15: Table of TCLK decoder associated registers.

| CB Address | Description | Read | Write |
| --- | --- | --- | --- |
| nn30 | Event FIFO 32 events deep (each entry is 8 bits) | Gets the next properly decoded TCLK event. | Clears the FIFO. |
| nn31 | Event FIFO Status | Gets the FIFO status. Disables writing to decoder RAM and enables normal event decoding | Disables decoding and enables writing to decoder RAM. |
| nn36 | Decoder RAM Address Pointer | Gets the current address pointer. The RAM is 256 x 4 bits. | Specifies the RAM address pointer. |
| nn37 | Decoder RAM Data | Gets the 4 bits of data at the current Address Pointer location. | Writes the lower 4 bits to the current address pointer location. |

The 4 bits of the RAM data are defined as follows:

| Bit | Description |
| --- | --- |
| 0 | A one in this bit location causes a 1Hz interrupt to occur. (Only used by the 8F Clock Event.) |
| 1 | A one in this bit location causes an Injection TBT Trigger to occur. |
| 2 | A one in this bit location causes a hardware Abort in Progress (AIP) to be generated. |
| 3 | A one in this bit location causes the decoded TCLK event to be written into the FIFO. |

#### Loading the TCLK Clock Event Decoder RAM

The procedure to load the clock event decoder RAM for the TCLK decoder is as follows:

Execute a Control Bus write to the TCLK FIFO status address (nn31). This enables the Decoder RAM Address Pointer and allows data to be written into the RAM. It also disables normal TCLK Event Decoding.

Write the value of the address you want to change within the Clock Event Decoder RAM to the Decoder RAM Address Pointer (nn36).

Write the value of the data you want within the RAM to the Decoder RAM Data (nn37).

Repeat steps 2 and 3 until the Decoder RAM is filled.

Execute a Control Bus read from the FIFO status address (nn31). This will disable writing to the Decoder RAM, and enable normal TCLK Event Decoding.

#### Example setup for a couple of TCLK Clock Events

This example sets up the Tevatron Prepare for Beam Clock Event ($71) and the Tevatron 1Hz Clock Event ($8F).

- Write any value to the TCLK event FIFO status address (nn31).

Write $71 to the TCLK Event Decoder RAM Address Pointer (nn36).

Write $08 to the TCLK Clock Event Decoder RAM Data (nn37).  This will cause the $71 clock event to be loaded into the Event FIFO when decoded.

Write $8F to the TCLK Event Decoder RAM Address Pointer (nn36).

Write $01 to the TCLK Clock Event Decoder RAM Data (nn37). This will cause a hardware 1 Hz interrupt when the $8F clock event is decoded.

Read from the TCLK Event FIFO status address (nn31).

### Beam Sync Clock (BSCLK) Decoder

The following is a list of Control Bus Addresses and functions involved in setting up the BSCLK Clock Event Decoder RAM.  The RAM is 256 x 3 bits where each 3 bit value is associated with a clock event (entries 0-255) and each bit specifies an action to be taken when that particular clock event happens.

Table

16: Table of BSCLK decoder registers.

| CB Address | Description | Read | Write |
| --- | --- | --- | --- |
| nn32 | Event FIFO 32 events deep (each entry is 8 bits) | Gets the next properly decoded BSCLK event. | Clears the FIFO. |
| nn33 | Event FIFO Status | Gets the FIFO status. Disables writing to decoder RAM and enables normal event decoding | Disables decoding and enables writing to decoder RAM. |
| nn38 | Decoder RAM Address Pointer | Gets the current address pointer. The RAM is 256 x 3 bits. | Specifies the RAM address pointer. |
| nn39 | Decoder RAM Data | Gets the 3 bits of data at the current Address Pointer location. | Writes the lower 3 bits to the current address pointer location. |

The 3 bits of RAM data are defined as follows:

| Bit | Description |
| --- | --- |
| 0 | A one in this bit location causes an internal AA event to occur. (Used by the $AA Clock Event.) An AA event feeds into the Make_Meas generating logic. |
| 1 | A one in this bit location causes a Studies TBT Trigger to occur. |
| 2 | A one in this bit location causes the decoded BSCLK event to be written into the FIFO. |

#### Loading the BSCLK Clock Event Decoder RAM

The procedure to load the clock event decoder RAM for the BSCLK decoder is as follows:

- Execute a Control Bus write to the BSCLK event FIFO status address (nn33). This enables the BSCLK Event Decoder RAM address pointer and allows data to be written into the BSCLK Event Decoder RAM. It also disables normal BSCLK Event Decoding.

- Write the value of the address you want to change within the Clock Event Decoder RAM to the BSCLK Event Decoder RAM Address Pointer (nn38).

- Write the value of the data you want within the RAM to the BSCLK Clock Event Decoder RAM Data (nn39).

- Repeat steps 2 and 3 until the BSCLK Clock Event Decoder RAM is filled.

- Execute a Control Bus read from the BSCLK Event FIFO status address (nn33). This will disable writing to the BSCLK Event Decoder RAM, and enable normal BSCLK Event Decoding.

#### Example setup for a couple of BSCLK Clock Events

This example sets up the Tevatron Data Acquisition Event ($DA) and the Tevatron Beam Rotation Marker ($AA).

- Write any value to the BSCLK event FIFO status address (nn33).

- Write $XX to the BSCLK Event Decoder RAM Address Pointer (nn38).

- Write $04 to the BSCLK Clock Event Decoder RAM Data (nn39).

- Write $AA to the BSCLK Event Decoder RAM Address Pointer (nn38).

- Write $01 to the BSCLK Clock Event Decoder RAM Data (nn39).

- Read from the BSCLK Event FIFO status address (nn33).

### MDAT Decoder

The MDAT decoder decodes the MDAT stream listening for frames $12 and $56.  Frame $12($56) is the TEV(MI) state frame and contains the 16-bit machine state for the Tevatron(MI).  Currently, only states 1-63 are supported by the BLM system, so the MDAT decoder copies bits 0…6 of the state value into the MDAT queue and puts a 0(1) in bit 7 if the machine is TEV(MI).

Table

17: MDAT queue registers.

| CB Address | Description | Read | Write |
| --- | --- | --- | --- |
| nn34 | Event FIFO 32 events deep (each entry is 8 bits) | Gets the next properly decoded MDAT frame. | Clears the FIFO. |
| nn35 | Event FIFO Status | Gets the FIFO status. | N/A |

# High Voltage Card

The High Voltage card is a double-wide 6U VME module that can carry one to four high voltage modules that are independently controlled through the VME bus. It has an 8-bit switch selectable card number which sets the card address corresponding to VME address (A31:A24). A Quad 12-bit DAC provides the program voltage for each of the high voltage modules for voltage output control. The combination multiplexer and 16-bit ADC reads the high voltage monitors for each module.

The card incorporates an FPGA (Altera Cyclone) to interface with the VME bus and provide local timing and control. The FPGA receives all the VME bus control, address, and data lines for reading/writing of data on the card. The FPGA can be programmed from the front panel through the Active Serial Program connector. The card has a serial number and ID memory to identify individual cards (see Section 1.10).

The program voltage circuitry consists of a 12-bit four-channel DAC device and an op-amp gain circuit. The DAC is controlled from the VME bus through an FPGA to select and send data to the selected channel. The DAC has an output range of 0 to 10V which feeds the control input of the HV module and results in a HV output of 0 to 2250V.  So a setting value of 0 ( FFF corresponds to an output of 0 ( 2250V.

To read all the high voltage monitors, the card uses a 16-to-1 multiplexer and 16-bit bipolar input digitizer. The circuit operates in a circular mode, such that the digitizer is continually digitizing all the monitor signals from each channel and storing the data into registers. The registers then can be read at anytime through the VME bus. The ADC has an input range of ±10V with an output of 0 to FFFF (0V input ( 8000 output). The timing and control is done through the FPGA.  Voltage readings are scaled such that 1V ADC input corresponds to 1kV and current readings are scaled such that 1V ADC input corresponds to 100μA.

## Address Maps

BLM HV cards respond to A32D16 and occupy <1 kb.  The base address of the board is set by a dip switch which corresponds to A31:A24.  The address offsets for the registers on the board are given in Table 18.  The only data that can be written are the HV settings and the HV enable words.  Note that only the lowest 12 bits are significant in the HV setting.

Table

18: VME address map of HV board.

| Offset | Access | Size | Description |
| --- | --- | --- | --- |
|  |  | Bytes | Bits |  |
| 000000 | R | 160 | --- | Board ID |
| 000200 | R | 8 | --- | Serial Number |
| Channel 1 |
| 000250 | R/W | 2 | 12 | Channel 1 HV Setting |
| 000252 | R | 2 | 16 | Channel 1 Output Current |
| 000254 | R | 2 | 16 | Channel 1 Output Voltage |
| 000256 | R | 2 | 16 | Channel 1 Return Voltage |
| Channel 2 |
| 000260 | R/W | 2 | 12 | Channel 2 HV Setting |
| 000262 | R | 2 | 16 | Channel 2 Output Current |
| 000264 | R | 2 | 16 | Channel 2 Output Voltage |
| 000266 | R | 2 | 16 | Channel 2 Return Voltage |
| Channel 3 |
| 000270 | R/W | 2 | 12 | Channel 3 HV Setting |
| 000272 | R | 2 | 16 | Channel 3 Output Current |
| 000274 | R | 2 | 16 | Channel 3 Output Voltage |
| 000276 | R | 2 | 16 | Channel 3 Return Voltage |
| Channel 4 |
| 000280 | R/W | 2 | 12 | Channel 4 HV Setting |
| 000282 | R | 2 | 16 | Channel 4 Output Current |
| 000284 | R | 2 | 16 | Channel 4 Output Voltage |
| 000286 | R | 2 | 16 | Channel 4 Return Voltage |
| Other stuff |
| 000290 | R/W | 1 | 1 | 0/1 disables/enables channel 1 supply |
| 000292 | R/W | 1 | 1 | 0/1 disables/enables channel 2 supply |
| 000294 | R/W | 1 | 1 | 0/1 disables/enables channel 3 supply |
| 000296 | R/W | 1 | 1 | 0/1 disables/enables channel 4 supply |

## Positive High Voltage Module

The High Voltage module is a self-contain module with two SHV connectors, one for high voltage output, the other for a high voltage return. A third connector is used as an I/O for input power, program voltage and high voltage monitoring. High voltage is produced by a dc to high voltage converter that is controlled by a programmed voltage input. The high voltage output is regulated and has low ripple output. The module also has three high voltage monitors for voltage output, voltage input return and current output monitoring. Shown below are the specifications for the module.

DC to HV Converter: 0 to 3000V @ 500µA max. Emco model: G30

Voltage Output: 0 to +2.25kV

Maximum current output: 500µA

Output Ripple: < 200mVp-p (full load) freq. @ 150KHz

Output Regulation: < .1% @ full load

Program input voltage: 0 to +10V

Voltage & current inputs: +12V @ 220mA, -12V @ 5mA (full load)

Monitor outputs:

HV current monitor: +1V/100uA

HV output monitor: +1V/1000V

HV return monitor: +1V/1000V

Status monitor (on/off): ON = +3.3V, OFF = 0V

HV input and output connectors: SHV, Kings No. 1704-1

I/O connector: AMP 534237-8

Module size: L 5” x W 2” x H 1.125”

# Abort Card

The Abort Card (AC) is the primary controller for abort collection and reporting.  The BLM abort status will be sent to the existing abort chassis in each service building.  In addition, each card will be part of a serial link which sends Tevatron abort information from house to house to a dedicated collection board located at A0.  A diagram of the abort cards data flow is given in Figure 5.  Functions and features of the abort card are as follows:

Reads the Abort information from each Digitizer Card once per measurement cycle.

Applies channel mask and multiplicity requirements before requesting an abort.

Stores the abort information in a VME-readable circular buffer for 65536 turns.

Periodically sends this abort information to the next house, forming a serial data stream of abort information.  Each successive house appends its data to the data stream.

VME registers exist for reading the circular buffer, testing the memory, sending diagnostic signals, reading board ID information.

Control Bus for setting the mask and multiplicity values, and for reading the Channel OK bits that were latched while reading abort information from the DCs.

Abort bus for reading the abort bits from the Digitizer Cards.

Masks and Multiplicity numbers are double-buffered..

Front Panel Lemo outputs: Immediate Abort, Fast Abort, Slow Abort, Very Slow Abort.

Front Panel BNCs:  Abort Out, Serial Link In, and Serial Link Out..

VME Test of Serial Link

VME interrupt capability to be specified.

Front panel displays and LEDs show the board status.

Remote FPGA programming capability using the VME bus

The Abort Card is master of the Abort Bus, which is used to gather abort data for the crate.  Upon receiving a Make_Meas signal,  the Abort Card sends addresses out on the AbortCS[5..0] and for each address sent it receives one channel’s worth of abort data on the Abort[3..0] lines.  These four bits correspond to the four “abort types” Immediate abort, Fast Abort, Slow Abort, Very Slow Abort.  If a bit is a “0” then the channel has a digitized value greater than its threshold value (preset in the Digitizer Card via the Control Bus).  AbortCS[5..3] represent the DC board address, and AbortCS[1..0] represent the four individual channels per board.  Upon receiving all the abort information, the Abort Card counts the number of channel hits per abort type and compares the value to a preset multiplicity value.  If any value is greater than the multiplicity threshold the AC sends an abort request to the Abort System via a BNC connector on the front panel.  Each of the four abort types behaves independently in this manner, and four Lemo connectors exist on the front panel in case individual pulse information is desired.  The abort outputs are active-low.  Thus the TTL high output corresponds to a beam permit.

Each digitizer channel in the crate can also be “masked” to prevent its inclusion in an abort request.  There are four masks per channel corresponding to the four abort types.

Abort 0	Immediate Abort

Abort 1 	Fast Abort

Abort 2 	Slow Abort

Abort 3 	Very Slow Abort.

A mask setting of “1” allows that channel to cause an abort.

The abort bus also includes a Channel OK line.  This indicates that the digitizer state machine is functioning properly.  In case of an error, the abort card keeps a mask of problem channels and can interrupt the control card to generate an error associated with the problem.  If a CHOK is false, the channel is not used in the abort multiplicity calculation.

The register map for control bus data is given in Table 23.

Figure

5: Overview of Abort Card functions.

## Address Maps

BLM Abort Cards respond to A32D16 and occupy 4Mb. The address offset and data for the Abort Card are as follows:

| Address | Description |
| --- | --- |
| 000000-000FFE | 8kB of Settings, ID etc. |
| 00A000 | JTAG FPGA programming register |
| 00A002 | JTAG enable register |
| 020000 | Interrupt Register |
| 030000 | Abort Status Register |
| 040000 | Memory Pointer LSBs |
| 040002 | Memory Pointer MSBs |
| 050000 | Memory Test Register |
| 050010 | Abort Control Register |
| 060000 | Serial Diagnostic Register |
| 060010 | Serial Link Test Register |
| 070000 | Position in Serial data stream (reserved) |
| 200000-3FFFFE | Abort History Circular Buffer (64k turns) |

The bits of the various settings and status registers are decoded according to Table 19.

Table

19: Definition of bits for VME Registers.

| Bit | Interrupt Register (reserved) | Abort Status Register | Memory Test Register | Serial Link Test Register | Abort Control Register |
| --- | --- | --- | --- | --- | --- |
| 0 | R/W | AIP | Memory Test Mode (allows VME write cycles to Circular Buffer) | Error Count Bit0 | Reset for VME and CB Registers |
| 1 | R/W | Circular Buffer “Wrapped” bit | Zero | Error Cnt Bit1 | Zero |
| 2 | R/W | Immed. Abort | Zero | Error Cnt Bit2 | Zero |
| 3 | R/W | Fast Abort | Zero | Error Cnt Bit3 | Zero |
| 4 | R/W | Slow Abort | Zero | Error Cnt Bit4 | Serial Link Reset |
| 5 | R/W | V. Slow Abort | Zero | Error Cnt Bit5 | Zero |
| 6 | R/W | Zero | Zero | Error Cnt Bit6 | Zero |
| 7 | R/W | Running | Zero | Error Cnt Bit7 | Zero |
| 8 | N/A | Reset Abort Status | R | Error Cnt Bit8 | N/A |
| 9 | N/A | N/A | N/A | Error Cnt Bit9 | N/A |
| 10 | N/A | N/A | N/A | Error Cnt Bit10 | N/A |
| 11 | N/A | N/A | N/A | Error Cnt Bit11 | N/A |
| 12 | N/A | N/A | N/A | Error Cnt Bit12 | N/A |
| 12 | N/A | N/A | N/A | Error Cnt Bit13 | N/A |
| 14 | N/A | N/A | N/A | Reset Err Count | N/A |
| 15 | N/A | N/A | N/A | Link Test Mode | N/A |

The Abort History is a circular buffer parallel to those in the Digitizer and Timing cards.  Each data record frame is 32 bytes, and corresponds to one turn.  The first 28 bytes contain the unmasked abort status for 56 of the 64 possible DC channels in a crate.  A 16-bit status word for four channels is structured as in Table 20.

Table

20: Definition of bits for each 16-bit word in the circular buffer.

| Bit | Channel | Status |
| --- | --- | --- |
| 0 | 0 | Immediate |
| 1 | 0 | Fast |
| 2 | 0 | Slow |
| 3 | 0 | Very Slow |
| 4 | 1 | Immediate |
| 5 | 1 | Fast |
| 6 | 1 | Slow |
| 7 | 1 | Very Slow |
| 8 | 2 | Immediate |
| 9 | 2 | Fast |
| 10 | 2 | Slow |
| 11 | 2 | Very Slow |
| 12 | 3 | Immediate |
| 13 | 3 | Fast |
| 14 | 3 | Slow |
| 15 | 3 | Very Slow |

The Abort History data record has the following structure; channels are numbered commensurate with backplane lines AbortCS[5..0].

Table

21: VME address offset corresponding to DC channels.

| Offset | Function |
| --- | --- |
| 0 | Status, Channels 0-3 |
| 2 | Status, Channels 4-7 |
| 4 | Status, Channels 8-11 |
| ……. | ………………. |
| 1A | Status, Channels 52-55 |
| 1C | Summary word 1 |
| 1E | Summary word 2 |

The two summary words of Table 21 contain the multiplicities of the aborts after the mask and the abort bits that are output to the beam-permit system:

Table

22: Last two words in each circular buffer frame (one frame per turn).

| Word | Bits | Function |
| --- | --- | --- |
| 1 | 0:3 | Abort bits (VS, S, F, I) |
| 1 | 4:9 | Fast Abort Multiplicity |
| 1 | 10:15 | Immediate Abort Multiplicity |
| 2 | 0:3 | 4-bit counter incrementing on MM |
| 2 | 4:9 | Very Slow Abort Multiplicity |
| 2 | 10:15 | Slow Abort Multiplicity |

The Abort Status Register looks at the status the four abort bits as seen on the front panel, except they are latched in this status register to record if any bits had gone high and subsequently cleared.  This was intended as a tool for the Operators to keep track of aborts around the ring.  The reset Abort Status bit resets only the four abort bit readbacks in this register.  Actually any write to this register will perform the reset.   These bits do not affect the front panel outputs.  The bits are also cleared by Sysreset, front panel reset, and bit D0 of the Abort Control Register. Also in this register is “running” which is the inverse of AIP, and the Circular buffer “wrapped” bit.  The Wrapped bit is cleared by a front-panel reset or the backplane signal Reset DC.

Table

23: Control Bus register map for Abort Card.

| Function | Type | Address | Bytes | Bits |
| --- | --- | --- | --- | --- |
| Masks |
| Abort Mask | R/W | 00 | 1 | 6 |
| Abort Channel Count | R/W | 01 | 1 | 6 |
| Abort Enable | R/W | 0A | 1 | 1 |
| Immediate Abort Mask | R/W | 10 | 8 | 64 |
| Fast Abort Masks | R/W | 18 | 8 | 64 |
| Slow Abort Masks | R/W | 20 | 8 | 64 |
| Very Slow Abort Masks | R/W | 28 | 8 | 64 |
| Multiplicities |
| Immediate Multiplicity | R/W | 40 | 1 | 6 |
| Fast Multiplicity | R/W | 41 | 1 | 6 |
| Slow Multiplicity | R/W | 42 | 1 | 6 |
| Very Slow Multiplicity | R/W | 43 | 1 | 6 |
| General |
| Force Abort H | R/W | 45 | 1 | 6 |
| Force Abort L | R/W | A4 | 1 | 6 |
| Force Abort Permit | R/W | A5 | 1 | 6 |
| Diagnostic |
| ADC Reading Temperature | R | 4E | 2 | 16 |
| Channel OK | R | 50 | 8 | 64 |
| Memory Pointer (Bits 11..04) | R | 02 | 1 | 8 |
| Memory Pointer (Bits 19..12) | R | 03 | 1 | 8 |
| Interrupt Register | R | A6 | 1 | 7 |
| Abort Status Register | R | A7 | 1 | 5 |
| Abort Snapshot Register | R/W | 60 | 32 | 256 |
| Abort OR Register | R | 80 | 32 | 256 |
| Snapshot Pointer (Bits 11..04) | R | 04 | 1 | 8 |
| Snapshot Pointer (Bits 19..12) | R | 05 | 1 | 8 |

Table

24: Control Bus register bits.

| Bit | Interrupt Register | Abort Enable | Abort Mask | Abort Status |
| --- | --- | --- | --- | --- |
| 0 | Assert IRQ3 when an abort occurs before multiplicity is applied | Abort Enable | Immediate | Immediate |
| 1 | Assert IRQ3 when an abort occurs | Zero | Fast | Fast |
| 2 | Assert IRQ3 when any Channel “not OK” | Zero | Slow | Slow |
| 3 | Zero | Zero | Very Slow | Very Slow |
| 4 | Raw Abort Interrupt | Two consecutive aborts |  |  |
| 5 | Abort Interrupt | Reset channel OK list | All | Logical OR |
| 6 | Ch not OK Interrupt | Reset Abort OR Register |  |  |
| 7 | Clear IRQ3 | Zero |  |  |

The Abort Snapshot register captures the last frame of the Abort History circular buffer just prior to an abort.   It will keep this frame until anything is written to Control Bus register 60.

The Abort OR register keeps a running tab of which abort bits went over threshold since the last reset.  This reset is bit D6 of the Abort Enable register or power-up or Reset CB.  The Abort OR register does not register the “sum words” that exist in the circular buffer.

## Remote FPGA Programming

The FPGAs are normally programmed by connecting a cable called a “ByteBlaster” between the parallel port of a computer and the JTAG connector on the front panel of the card, and downloading the proper programming file using Altera’s Quartus program into the card’s configuration device.  The configuration device holds the FPGA programming bitstream in non-volatile memory, and programs the FPGAs whenever the board is powered-up.  An alternative method of loading the configuration device exists using four bits of a VME register to emulate the four bits of the JTAG lines.  Table 25 shows the bit assignments.

Table

25: VME register NN00A000 for loading the configuration device.

| Bit | JTAG Function |
| --- | --- |
| 0 | TCK |
| 1 | TMS |
| 6 | TDI |
| 7 | TDO |

This programming register is at NN00A000.  As a precaution, in order for this programming operation to occur, a special code must be written to register NN00A002.  Currently the only requirement in this special code is that a 1 be written to bit 0.  The idea is obviously to ensure that no random or erroneous VME operations corrupt the FPGA configurations.  The software routines for carrying out the remote programming are part of the front end code collection.

## Reset Functions

A breakdown of the reset functions of the board follows:

Front Panel Reset Button:  VME State Machine, VME Interrupt Register, Control Bus sequencing, CB Interrupt Register, Abort Outputs, Abort enable bit, Snapshot register, wrapped bit, circular buffer pointer, all Serial state machines and FIFOs, Serial error detector, remote programming special code.

VME SysReset:  VME State Machine, VME Interrupt Register, Control Bus sequencing, CB Interrupt Register, Abort Outputs, Abort enable bit, Snapshot register, all Serial state machines and FIFOs, Serial error detector, remote programming special code.

Abort Control Register Bit D0:  VME Interrupt Register, AbortStatusRegister, MemTest Register, JTAG programming registers, Control Bus sequencing, CB Interrupt Register, Abort Outputs, remote programming special code.

Abort Control Register Bit D4:  All Serial state machines and FIFOs, Serial error detectors.

Reset DC (backplane):  Wrapped bit, circular buffer pointer, Snapshot register.

Reset CB (backplane):   Abort enable bit, Snapshot register.

# Abort Card at A0

A special Abort Card at A0 is used for initiating the Serial link data stream and processing all the houses’ data from the serial link once they make it completely around the ring.  All three FPGAs on this abort card need special firmware.  Although the data is serialized for transmission on the buried coax between houses, the Serial Link FPGA on each Abort Card treats the data a byte at a time.

Upon receipt of the signal “Abort Sync” from the Timing Card, A sixteen byte “preamble” is sent out of the A0 card to the first house, A1.  This house checks for the correct preamble and puts its 32 bytes of data immediately at the end of the preamble and sends out the data to A2.  Each subsequent house tacks its 32 bytes of data immediately onto the end of the data from the upstream house until the whole ring’s worth of data arrives back at the A0 Abort Card.  The A0 card largely ignores the four sum bytes from each house, separates out the bits based on abort type, adds up the bits and compares these four sums to serial-link multiplicity values also based on abort type.  If any are over threshold, and also “enabled”, an abort signal goes out the front panel.  The Abort_Sync occurs every 2 msec.

## Address Maps

The following table contains the VME registers for the A0 Abort Card.  There is no Control Bus access.

| Address | Description |
| --- | --- |
| 060000 | Serial Link Data Buffer FIFO |
| 060020 | Upstream House Link Errors |
| 070000 | Link Errors |
| A00000,A00002 | Multiplicity Registers |
| A00004 | Abort Enable Register |

## Abort Multiplicity Registers

There are four abort multiplicity settings, one for each abort type.

| Address | Bits |
| --- | --- |
|  | 15…8 | 7…0 |
| A00000 | Fast multiplicity | Immediate multiplicity |
| A00002 | Very Slow multiplicity | Slow multiplicity |

## Abort Enable Register

The Abort Enable register (NNA00004) allows individual inclusion of each abort type from around the ring to be included in the ringwide abort signal.

| Bit | Abort Enable |
| --- | --- |
| 0 | Immediate abort enable |
| 1 | Fast abort enable |
| 2 | Slow abort enable |
| 3 | Very Slow abort enable |
| 4 | Zero |
| 5 | Zero |
| 6 | Preamble Error Read and Reset Error |
| 7 | Abort Read and Reset Abort |

## A0 Serial Link Data Buffer

The A0 Serial Link Data Buffer (NN060000) is a FIFO memory that stores the complete data set that makes it to A0 prior to the latest abort.  The data is comprised of all 32 bytes of circular buffer data from each house, including sum words, for a total of 880 bytes (16 byte preamble + 27 houses x 32 bytes).   Data order is the preamble from the A0 Abort Card, A1 data, A2 data, and so on.  To obtain this data, read NN060000 a minimum of 880 times.  As long as Abort_Syncs are occurring, this buffer will accumulate data, writing over the previous ring’s worth of data with each Abort_Sync.  When an abort occurs the buffer will continue filling until it receives all the data from the last house in the chain.  Further writing is then prohibited.  The buffer is re-armed when bit 7 of the Abort Enable register is written high.

## A0 Serial Link Error Handling

The Serial Link Receiver on each Abort Card checks for valid 8B/10B data arriving at the input.  If an invalid code is detected an error bit is generated.  This error bit is placed on bit D3 of the last of the four sum bytes in the serial data (it does not affect the counter in the circular buffer described above).  This bit is the exception; the rest of the bits in the four sum bytes from each card are ignored by the A0 Abort Card.

If this bit is set from any of the ACs in the ring, it is assumed that the data for the present sample prior to this point is wrong and the serial link multiplicity counters are reset, preventing an abort.  Error free transmission after this point in the turn may still abort if there are enough bits set.  Upon receiving the next Abort Sync pulse from the TC, the algorithm starts anew.   A register at NN070000 is available for reading the number of errors detected around the ring.  A fourteen bit number counts the errors between resets to the register.  The number resides in bits D13-D0, and the reset is bit D14.  To date, no errors have been detected on the link.

A similar register which monitors only the errors from the upstream house resides at 14060020.  It is also has a fourteen bit counter and is reset by bit D14.

# Control Card

The Control Card (CC) is the primary controller for the Tevatron BLM system.  It includes the following features:

eZ80 microprocessor running the main control program

1 MB flash memory for storing code

2 MB RAM for running the control program

Control Bus Interface Master for setting parameters and reading back sliding sums, etc.

8 Megabyte Block of shared memory with the VME bus.  This is used for getting setup parameters and for storing the circular buffers and other frame memories.

Two RS-232 ports for diagnostics and standalone operation

A TCP/IP 10/100 baseT port for standalone operation (not implemented)

The Control Card functions by running a control program [7] on the eZ80 that

initializes the other cards,

responds to TCLK events

prepare for beam and end of beam events which start and stop the DAQ

Flash, profile and display frame events

reads the sums from the digitizers,

responds to MDAT state changes by changing abort thresholds and masks,

responds to requests from the front end crate controller

changes to abort thresholds and masks

changes to other settings

The program is stored in the 1 MB flash memory and can be reprogrammed remotely.  At boot time, the program is copied to the 2 MB RAM where it then runs.  The program is kept in the Controls Department’s CVS repository on nova.fnal.gov under project blmez80.

## Address Maps

The control card address space is a 24-bit address space, enough to address 16MB of memory.  The eZ80 is allowed to access the entire space while the VMEbus is allowed to access a subset of it.  The memory space of the control card is shown in the following table.

| eZ80 address | VMEbus address | Contents | Size |
| --- | --- | --- | --- |
| 000000-0FFFFF |  | Flash | 1MB |
| 200000-3FFFFF |  | eZ80 Program RAM | 2MB |
| 400000-401FFF |  | Control Bus | 8KB |
| 500000-507FFF |  | Internet Chip W3100A (not implemented) | 32KB |
| 610000 | NN610000 | FPGA reconfiguration signal: DCLK |  |
| 620000 | NN620000 | FPGA reconfiguration signal: DATA |  |
| 630000 | NN630000 | FPGA reconfiguration signal: SER_ENn |  |
| 700000-7000EF | NN700000- | ID block, utility RAM | 240B |
| 7000F2- 7000F3 | NN7000F2- | Temperature | 2B |
| 7000F8-7000FF | NN7000F8- | Silicon ID: family code, ID, check sum | 8B |
| 700100-7001EF | NN700100- | Utility RAM | 240B |
| 800000-FFFFFF | NN800000- | Shared RAM | 8MB |

The VMEbus uses A32D16 mode.  The higher 8 bits of the address, “NN”, are set by a DIP switch on the control card [1].  Note that the byte alignment is reversed from VMEbus standard.  The byte written to the even address by the eZ80 is the lower byte in the 16-bit VMEbus data.  The VMEbus can only access memory space from NN610000 to NNFFFFFE.  The control card firmware does not guard on invalid accesses.  Attempts of accessing memory addresses not listed in the table will have effects not documented here.  Users are advised to access only the listed space.  The lower addresses are only accessible by the eZ80.

### The 8MB Shared Memory Space

On the control card, 8MB of RAM are designed to be accessible from both VME and eZ80 sides.  The access is time shared using sequential circuits.  In practical program writing, the users are allowed to access the memory “simultaneously” from both VME and eZ80 sides.  However, it is a common practice to use a handshaking protocol so that a block of data is readout by one process only after the other process has finished writing valid data into the block.  The shared memory begins at NN800000.  Addresses described below are added to this address.  Table 26 lists the main blocks of memory in the 8MB of shared memory on the Control Card.  For the details of every address, see [7].

Table

26: Map of the shared memory on the Control Card.  This table enumerates blocks of addresses and not individual address details.  For that information see [7].

| Address | Size | Description |
| --- | --- | --- |
| 000000 | 2 | eZ80 Program Status Word |
| 000002 | 30 | Various front end triggers and settings |
| 000020 | 32 | Indices and pointers to the various data storage buffers |
| 000050 | 34 | Temperatures |
| 000080 | 8 | Channel Not OK List from Abort Card |
| 000100 | 512 | Various settings (not all addresses are used) |
| 000300 | 256 | 32-bit pedestals stored in standard data record |
| 000400 | 120 | 16-bit squelch values |
| 000500 | 74 | Abort Snapshot Info |
| 000600 | 256 | 32-bit Integration pedestals |
| 010000 | 48 | Program name, version, and build date |
| 010030 | ~64K | Debug information |
| 020000 | 128K | Buffer for new flash memory contents |
| 080000 | 128K | 256 Flash Frames |
| 0A0000 | 128K | 256 Profile Frames |
| 0C0000 | 512 | Display Frame |
| 0E0000 | 256 | MDAT state to Abort state map |
| 0E0100 | 512 | TEV F Sector - MI TCLK Info |
| 0F0000 | 64K | TEV F Sector Info |
| 100000 | 256 K | Abort Settings for Abort States 0-255. Each state is 1K of data. The front end writes new settings to this block. |
| 140000 | 256 K | Copy of settings that are in use. The eZ80 program copies the data from the previous block to this block. These are the settings that are copied to the AC and DCs. |
| 180000 | 512 K | DeRippled Sum Buffer 2K frames deep |
| 200000 | 4 Meg | Fast Sum Buffer 16K frames deep |
| 600000 | 1 Meg | Slow Sum Buffer 4K frames deep |
| 700000 | 1 Meg | Very Slow Sum Buffer 4K frames deep |

### The ID Block and Utility RAM

The address ranges NN700000-NN7000EF and NN700100-NN7001EF is implemented with a small memory block.  (In fact, it is the same memory block that supports the temperature and silicon ID)  This memory contains the ID block.

Users are allowed to read and write into this small memory block; therefore it is also called utility RAM.  However, care must be taken while writing from eZ80 side.  The utility RAM is organized as 16-bit words that are can be read or written in a single operation from VME.  From eZ80, the even and odd bytes can be readout separately.  However, while writing from the eZ80, either to the even or odd byte, both bytes in the 16-bit word are written with the same number.  For example, if a byte XY is written into 700000 or 700001 the 16-bit word becomes XYXY.

The utility RAM can be used as handshaking information exchange resource.  It is important that a data item being exchanged is written in a single operation.  The inconvenience while writing from eZ80 side described above was primarily for simplicity of the implementation.  However, it also prevents the users from mistakenly sending 2-byte information that needs two operations from eZ80 to VME for handshaking.

## Remote Programming

### FPGA

The FPGA on the control card is configured from an EEPROM device AT17LV002A at each power up.  The contents of the EEPROM can be read and rewritten from both eZ80 and VME sides.  Once the EEPROM is successfully rewritten, the FPGA will be configured with new firmware at next power up.

There are three signals used to control reading and writing information into the EEPROM:

- The DCLK signal at NN610000

- The DATA signal at NN620000

- The SER_ENn signal at NN630000

These signals are single bit registers.  To set a signal high or low, write 1 or 0 to the given address from the eZ80.  From VME, write FFFF to set the signal high and 0 to set it low.  A signal stays at the level after being written until it is overwritten with the opposite level.  The level of DATA can be read.  Bit 0 of the byte at 620000 or word at NN620000 represents the level of signal DATA.

Check the card for the following to ensure correct reconfiguration:

- BP112 must be present.

- Jumper J5, J6 and J7 must be installed.

### eZ80

The eZ80 program must be running to reprogram it remotely since VME does not have access to the Flash memory.  The general approach is to write the new code to a section of dual-port memory and then trigger the eZ80 program to copy it to flash memory.  See [7] for details about reprogramming the eZ80.

## Data Structures

BLM data is reported in two different ways: Snapshots and Profiles.  A profile is taken in response to a TCLK event.  Profile data are the most recent data from the specified sum (Fast or Slow) from each BLM channel.  The “Flash” and “Display” frames are the same as a Profile, but are taken in response to a different TCLK event.  It is possible to store up to 256 Flash frames and 256 Profile frames in the CC for later readout and analysis.  There is only one Display Frame.  The current frame indexes are stored in registers.  The Snapshot buffers contain the history of losses for all channels for data from each of the sliding sums.  These data are stored in the circular buffers of the Control Card and can be read by the front-end CPU from the CC’s dual-port memory and then reported to ACNET.  Snapshot data is stored in circular buffers in the same format as the Profile data.  In addition to the sliding sum data, Profile/Flash/Display frames include the integrated loss since the last reset for the Main Injector and the very slow sum for the Tevatron.  Thus these data records are 512 bytes long, while the Snapshot (circular buffer) records are only 256 bytes long.  Table 27 enumerates the contents of a 256 byte data frame.

Table

27: Contents of a 256 byte data frame.  Profile/Flash/Display frames contain 2 of these frames.  The Timestamps are stored in Little-Endian format.

| Byte Offset | Length | Description |
| --- | --- | --- |
| 00 | 1 | Machine state (0-255) |
| 01 | 1 | Measurement Divisor |
| 02 | 2 | Sum Divisor |
| 04 | 1 | Instantaneous Abort Status from AC |
| 05 | 1 | Channel Count |
| 06 | 1 | Data flag to indicate normal (0), last of cycle (1), new cycle (2), or waiting for stable data (3) frame |
| 07 | 1 | Unused |
| 08 | 4 | Time stamp: Microseconds since last TCLK 1 Hz Event |
| 0C | 4 | Time stamp: seconds since T0 (i.e. 1 January 1970) |
| 10 + n*4 | 240 | Sliding Sum data for channel n |

## Abort Settings

The abort settings for the system are downloaded into the CC memory by the front end.  In addition to the thresholds and the masks, there is also an array which maps MDAT state to abort state.  The thresholds, masks, and the map are all entered by the user through a console application.  There is currently support for states 1…63.  This When a new MDAT state appears, the map is used to determine which abort state to switch to and the CC

# Revision History

| 1.00.0 | 10/14/04 | AB | First Posting in Doc DB. |
| --- | --- | --- | --- |
| 1.00.1 | 10/19/04 | AB | Improved Figures for JDL. |
| 1.00.2 | 10/26/04 | AB | Modified address map to have maximum of 60 Channels. |
| 1.00.3 | 11/2/04 | AB | Modified Address map card IDs etc. and updated J2 Connector pin list to include standard VME in B row. |
| 1.00.4 | 11/2/04 | AB | Fixed errors in Modified Address map for Card Ids etc. |
| 1.00.5 | 4/14/05 | JL | Major formatting overhaul. Clean up address maps. Add introduction. |
| 1.00.6 | 6/6/05 | JL | Modified CC map to include integral measurements. Added info on HV and Abort Cards. Minor cleanups. |
| 1.01.0 | 6/10/05 | JL | CC map fixed. Ready for prime time. (FLW) |
| 1.01.1 | 7/25/05 | JL | Cleanups. Update AC description |
| 2.00.0 | 4/5/07 | RMTK | Semi-major overhaul to reflect all the changes that occurred during production and integration. |
| 2.00.1 | 6/21/10 | RMTK | Tweaks and fancification. |

# References

BEAMS-DOC-2360, “VME Base Address Switch Settings for the BLM System”.

BEAMS-DOC-1126, “Tests of BLM Integrator”.

BEAMS-DOC-1736, “BLM VME J2 Backplane Crosstalk Evaluation”.

BEAMS-DOC-1994, “BLM VME J2 Backplane Impedance”.

BEAMS-DOC-2000, “BLM System Control Bus Testing”.

BEAMS-DOC-3077, “Digitizer Card Operating Guides”.  See http://www-ppd.fnal.gov/EEDOffice-W/Projects/ckm/comadc/DCopGuide.doc for the latest version.

BEAMS-DOC-2433, “BLM Upgrade Control Card Program Design”.

# Footnotes

The dual porting is implemented in the on-board FPGA as a time-sliced sharing between VME and the eZ80.