Fermilab/BD/MI

Beams-doc-2083

July 24, 2006

Electronic Design Document

Steve Wolbers, Vince Pavlicek, Bill Haynes, Greg Deuerling,

Luciano Piccoli, Steve Foulkes,

Fermilab - Computing Division - CEPA

Bob Webber, Fermilab - Accelerator Division - Accelerator Controls Department

Abstract

This document contains the specification for the Timing Generator Fanout (TGF-II) for the Beam Position Monitor (BPM) upgrade for the Main Injector at Fermilab.  All required operating modes, system timing, diagnostics and interactions with the BPM software are described in this document

Overview	3

Beam Position System	4

BPM Sensor	4

BPM Data Acquisition System	5

BPM Requirements	8

Data Acquisition	8

Functional Overview	8

Echotek Data Converter Module	8

Timing Generator Fanout (TGF-II) Module	10

TGF-II Firmware	14

TGF-II Flow Diagram of Main Injector	14

TGF-II Simplified Block Diagram	16

TGF-II Simplified Block Diagram	17

TGF-II  Register Map  (from Offset = 2000h)	18

TGF-II Schematics, Layouts and drawings	34

Sub Rack	34

Sub Rack Processor Module	34

Sub Rack Cables	34

Installation	35

Cable Specifications	36

Sub rack Specifications	36

Echotek DSR Module	36

# Overview

This document describes the Timing Generator Fanout (TGF-II) needed for the precision timing and synchronization of the data acquisition readout of the Main Injector Beam Position Monitor (BPM) upgrade project.  The data acquisition (DA) hardware digitizes the analog position signal from the BPM sensors, amplify them and digitally filter the signals and make them available to a VME front-end computer.    Figure 1 shows the functional block diagram and the different elements involved with the BPM upgrade project.

The hardware of Figure 1 consists of:

- Amplification and signal conditioning of the BPM sensor signals in the Combiner Box and the Transition card, designed at FNAL.

- The precision timing functions are implemented in the VME Timing Generator Fanout (TGF-II) module, designed at FNAL.

- The control and diagnostics of the Transition Cards are performed by the Transition Controller interfaced to the TGF-II.

- The Digital Signal Receiver (DSR) is a Commercial-Off-The-Shelf (COTS) module from Echotek Corporation, the ECDR-GC814/8-FV2.

- The VME sub rack controller is a COTS Single Board Computer from Motorola Corporation, the MVME5500

# Beam Position System

## BPM Sensor

The Main Injector has four BPM detectors per betatron wavelength in both the horizontal and the vertical planes. There are a total of 208 BPMs for the 3320 m ring. Out of these, 203 are MI style ring BPMs, and are located in the downstream end of every MI quadrupole. The other 5 are wide aperture BPMs located at Q101, Q402, Q522, Q608 and Q620 locations.

The BPM detectors in MI consist of four transmission line strips, or strip-lines, located on the perimeter of the beam pipe. The strip-lines are 250 mm long and 12.7 mm wide. They are positioned at a distance of 46.7 mm horizontally and 44.5 mm vertically, center to center as shown below.

BPM Sensor inside quadrupole beam pipe

The pick-up has a characteristic impedance of 50(, determined by the gap between the strip and the beam pipe. The current BPM RF module input impedance is matched to 50( within a 5 MHz bandwidth centered at 53MHz. The outputs are combined in pairs externally to form either a horizontal or a vertical detector. Each strip-line is shorted at one end and connected to a ceramic feed-through at the other end, which makes these BPMs non-directional.  At present any BPM measures either the horizontal or the vertical position per cycle at each quadrupole. It can be remotely switched to the orthogonal mode.

The wide aperture BPMs, with 6” long plates and a 4.625” aperture, are located adjacent to the Lambertson magnets and are mounted external to the quadrupole magnets.

## BPM Data Acquisition System

The BPM signals pass through the Combiner Box for band pass filtering (53 MHz & 2.5 MHz) and to the 8-Channel Transition Card for and additional filtering and amplitude control.  These conditioned signals are fed to an 8-channel digitizer from Echotek. Timing for the digitizer and the CPU are performed by the Timing Generator Fanout (TGF-II). Timing to the digitizer consists of a Clock and a Sync (or gate) input based on six Main Injector timing signals.  1) The Main Injector RF Clock (RFClk) operating from a range of 52.8Mhz to 53.1Mhz, 2) Main Injector BSync (MIBS), operating at 7.5Mhz and synchronous the MI beam, 3) Recycler BSync (RRBS), operating at 7.5Mhz and synchronous to the Recycler Beam, 4) Booster Extraction Sync (BES), 5) TClk operating at 10Mhz and asynchronous to the beam containing accelerator-complex-wide commands,  and  6) MDat operating at 10Mhz also asynchronous to the beam and containing accelerator-complex-wide commands.

The sensor front-end electronics depicted in Figure 1 are contained in two separate sub racks.  The 8-channel Transition cards is installed in a VME style sub rack but is strictly for mechanical purposes and convenience.  Up to 10 Echotek digitizers can be installed in a standard VME sub rack and controlled by a commercial CPU and timed synchronously to the Main Injector beam by the TGF-II.  The BPM sensor front-end electronics gain and filter selections are controlled by the Front End Controller (a VME transition card inserted in the rear I/O area in the same slot as the TGF-II card).  Figure 2 illustrates the organization of the cards in the BPM VME sub rack.  One sub rack can hold the necessary electronics needed for a service building (sometimes called a house).  There are 7 service buildings or houses around the Main Injector ring and each BPM house can have up to 10 BPM digitizing boards.

The BPM digitizing modules are Echotek module ECDR-GC814/8-FV2, and each module digitizes signals from 8 channels.  An individual BPM sensor generates information on 4 channels – the A and B plates for the proton end of the sensor and the A and B plates for the antiproton end.  Therefore one Echotek module will be capable of accepting signals from two BPM sensors.  The digitized output that results from each channel of the digitizing module may be sent on for further processing or be used to generate a calculated position when processed with data from the matching plate channel.  The details are in the Software Specifications document.  If the output is the digitized and filtered signal data, each channel is actually represented by two components: a real (Q) and imaginary (I) part.  The digital filtering on the Echotek module can also be disabled allowing a raw, oscilloscope type view of the signal.

The operation of the Echotek module is controlled by two inputs that are common for all eight channels.  They are the CLK and SYNC signals.  The digitizing clock paces the Echotek module and comes from the BPM Timing Generator Fanout (TGF-II) Module.  A phase-lock-loop on the Timing module creates a clock output that has a fractional relationship to the Tevatron RF clock frequency.  TeV RF is approximately 52.8 MHz.  The Echotek clock is 10/7ths or approximately 75.43 MHz.  This means that the 52.8 MHz signal is technically under-sampled but the signals-of-interest are modulated onto that RF signal and filtering and signal processing techniques allow the module to extract those signals-of-interest.  The digitization process can be gated, singly triggered or triggered for a specified number of samples by the SYNC signal.  SYNCs are generated in the TGF-II and normally initiated with respect to the Main Injector BSync (MIBS) or Booster Extraction Sync (BES) event. There are two additional decoders (MDAT & TClk) used to interrupt the VME CPU for notification of a Machine state change and alarms relative to a beam reset.  The MIBS decoder or BES to the TGF-II can initiate digitizing sequences and can add fixed delays with a resolution of one-half RFClk period, in order to remove cable delays or to time-in similar actions between the widely separated service buildings around the Main Injector ring.  The TGF-II card is capable of additional delay resolutions by changing the PLL count values of the FPGA firmware.

Figure 2.  Sub Rack Module, Example Configuration

# BPM Requirements

The requirements for this system are based on the system requirements documented in Main Injector BPM Upgrade Requirements Document (Beams Doc #1949).  A summary of the functional requirements are:

- User defined turn-by-turn: A measurement of the orbit on every turn (588 53Mhz RF buckets) of every BPM for 2048 turns, performed in wide bandwidth mode.

- Injection and extraction turn-by-turn: A wide bandwidth measurement of the orbit on every turn of every BPM for each batch of the beam injected into the machine. It also provides the capability to measure the extraction turn of at least one portion of the beam extracted from the machine.

- Flash Frame:  A single orbit measurement, performed in wide bandwidth mode for each BPM. The flash frame consists of the first turn that has beam for injections, or the last turn that had beam for extractions.

- Average Orbit: A narrow bandwidth measurement that consists of the orbit averaged over the first 512 turns of the flash turn buffer.

- Closed Orbit: A narrow bandwidth measurement collected from all BPMs at a rate of 500Hz.

- Display Frame: A narrow bandwidth measurement triggered by the profile frame TClk (0x7B). Data for this measurement is pulled from the closed orbit buffer.

- Profile Frame: A narrow bandwidth measurement triggered by the profile frame TClk (0x7A). Data is pulled from the closed orbit buffer.

# Data Acquisition

## Functional Overview

The communication between the front-end processor and the Echotek modules happens through the VME backplane.

Data coming from the BPM are stored into a set of buffers in the sub rack controller.  The depth and number of logical buffers that reside on the sub rack controller is defined in the Front-end Software Design Document (Beams Doc #1949).  Some of the buffers, most notably the turn by turn data, is first stored in memory on the Echotek boards and then transferred to the sub rack controller on demand, as the processor/backplane does not have the bandwidth to acquire it in real time.  The actual physical implementation of the sub rack controller buffers will be described in the front-end software design document.

## Echotek Data Converter Module

The Echotek module ECDR-GC814/8-FV2 is the baseline data converter.  This module belongs to a general type of modules called Digital Signal Receiver (or Radio) (DSR).  These modules have high speed digitizers that convert the analog, usually radio frequency, inputs into digital information very early in the signal processing.  This reduces the amount of analog circuitry in the system.  The subsequent digital circuitry is used to down convert or base band the RF information, removing the high frequency carrier (52.8 MHz in this case) and revealing the information modulated onto that carrier, 48 KHz revolution information, 20 KHz betatron signals etc.  The base banded signals are filtered in n stages of digital FIR filters and the output data is stored in a buffer memory on the module.  In this Echotek module the 85 MHZ, 14 bit Analog-to-Digital converter is followed by a Greychip Digital Down Converter (DDC) chip and then there is a re-programmable FPGA to route the data to onboard memory accessible to the VME bus.  The sub rack processors can read out the data memory over the sub rack VME backplane.

Operation of the DAQ software is described in Beams doc #1949, Main Injector Beam Position Monitor Upgrade Software Specifications for Data Acquisition.  The Echotek boards are configured by initialization files that are loaded during the startup process.

The data sheets for module are in the Appendix as Echotek DSR Module.  The module front panel has eight input connectors, a digitizer clock and a synchronization input that can be used as a trigger or a gate.  The module signals are described in more detail in the data sheets.

The firmware in the module can be updated by running the driver code software (?) and directing it to the new firmware object file.  The system should be not running any other software at this time to avoid interference with the loading process.  One Engineering Change Order (ECO) was applied to this module after consultation with Echotek Corp.  The SYNC input termination resistor was lowered from 1000 ohms to 50 ohms to match the signal source and to fix a cross talk problem found early in testing.

## Timing Generator Fanout (TGF-II) Module

The Timing Generator Fanout (TGF-II) is a double width, 6U VME card that generates timing signals that control and initiate acquisition of the BPM signals. The timing signals are based on inputs from the Main Injector clock system.  These signal inputs include the Main Injector RFClk, BSync, TClk, MDat, Recycler BSync and Booster Extraction Sync (BES).  The TGF-II is based on the design of the Tevatron TGF and is backwards compatible with a three exceptions. 1) There are ten Clock and Sync outputs instead of eight.  2) The turn counter firmware is 1176 for the Main Injector and 2226 half buckets for Tevatron.  3) the Clock output of the TGF-II is the RFClk times 10/7 for the main Injector rather than 7/5 for Tevatron.

The RFClk is the synchronization time base for the major functionality of the TGF.  The RF frequency shifts as the energy of the beam changes and thus frequency values shown are usually approximations.  The TGF internally doubles the RFClk to over sample and decode all other timing inputs (MIBS, RRBS, BES, MDat and TCLK) signals.

The BSync base frequency is one seventh of the RF Clock or approximately 7.5 MHz. events are encoded as an eight bit serial byte onto this carrier and can be extracted from it.  An important event is the revolution timing mark (0xAA).  This event allows synchronization with the pattern of particles circling around the Main Injector ring.

The TClk and MDat signals are encoded similar to BSync but with a 10MHz carrier frequency and from it the TGF-II can extract up 255 events from each of the input signals. The TGF-II can be programmed to respond to any of up to sixteen selectable TClk events to interrupt the sub rack controller and any of those TClk events can generate up to 32 timer alarm interrupts with ½ bucket resolution.

For each Echotek DSR module the TGF-II creates a digitization time base clock that paces the data conversions and filtering on the board.  It also supplies a trigger signal for control of the data acquisition.  For the Main Injector BPM system, the TGF-II multiplies the RF clock frequency by exactly 10/7 to produce the digitization clock.  The range of the Main Injector RFClk is approximately 52.8 MHz to 53.1 MHz.  This produces range of the digitization rate of 75.43Mhz to 75.86Mhz.  The TGF-II module can also generate other multiples of the RF Clock for the digitization clock.

Digitization operations will be triggered by and/or delayed from the BSync revolution timing mark or Booster Extraction Sync (BES) so that the acquired data are related to the particle bunches and measurements in separate service buildings are synchronized to each other as necessary.

#### Front Panel I/O Description

See Error! Reference source not found.

The Timing Generator Fanout Module inputs are:

- RFCLK – Main Injector RF Clock (52.8Mhz -53.1MHz).  The RFClk must have signal level greater than 200mv and capable of driving an ac-coupled 50-ohm termination.

- BSync0 – Main Injector Beam Sync is an approximately ~7.5 MHz carrier with a modified Manchester encoding of Tevatron beam specific events. MIBS must be capable of driving 50-ohms with a signal level greater than +1.2v.

- BSync1 – Recycler Beam Sync is an approximately ~7.5 MHz carrier with a modified Manchester encoding of Tevatron beam specific events.  RRBS must be capable of driving 50-ohms with a signal level greater than +1.2v.

- TCLK – Main Injector TClk, a 10 MHz carrier modulated by a fixed protocol of bit patterns that can carry commands and data events to electronics across the accelerator complex. TClk must be capable of driving 50-ohms with a signal level greater than +1.2v.

- MDAT – A 10 MHz carrier modulated like TCLK to carry Machine state data and events.  TClk must be capable of driving 50-ohms with a signal level greater than +1.2v.

- BES – Booster Extraction Sync, a trigger pulse from the Booster that indicates a beam batch is being extracted to the Main Injector. BES must be capable of driving 50-ohms with a signal level greater than +1.2v.

The Timing Generator Fanout Module outputs are:

- A2D Clock signals [0-9] – the digitization clock to the Echotek modules. These regenerated clock outputs are based on a signal that is phase locked to the RFClk with a frequency of RFClk multiplied by 10/7.  Output levels are >700mv into 50-ohms.

- A2D Sync signals [0-9] - the synchronization output for a sequence of digitization.  The Sync output signals have a pulse width of ~50ns and are capable of driving a 50-ohm termination.

- An External Trigger Output (Trig0)–Allows an external device to be triggered from an internal firmware trigger condition such as a decoded BSync event.  Currently set up to trigger on a Bsync Start event.

- An External Trigger Output (Trig1)–Allows an external device to be triggered from an internal firmware trigger condition such as a decoded BSync event.  Currently set up to trigger on a BSync Turn event.

- VME Interrupts for signaling the sub rack controller.

- Backplane connections that allow the TGF-II to control the Tevatron Filter Boards.  These lines will also allow a level of auto-detection on a per sub rack basis so that the TGF-II will ‘know’ how many Filter boards there are to control.

The A2D Sync signal variations are:

- BSync trigger direct AA (turn marker) detect.

- BSync + programmable delay

- BSync + repeated programmable delay

- BES – Similar to above BSync except it is a direct trigger and bypasses the Manchester decoding section.

Diagnostic Signal (for Tevatron Filter card)

The TGF-II can create a diagnostic signal called Diagnostic Clock above that is distributed to the Analog Filter Boards Tevatron and can be used to test different parts of the hardware.  It is not a calibration signal but is expected to be stable enough to look for long term degradation in individual channels relative to the other channels.  The diagnostic signal can be set to a continuous 53Mhz wave or a short burst of 53Mhz pulses via VME writes to TGF-II offset 0x100.  Some example waveforms are shown in Beams Doc #1586.  From the TGF, the diagnostic signal is distributed to the Filter Boards on a User-Specified bus on the P2/J2 portion of the VME backplane.  The signal is approximately 53 MHz which is near the backplane frequency capability so termination of the buss is provided on the TGF-II module and by a dedicated termination module as far away from the TGF-II as possible, usually in slot 20.  This termination module was made by replacing the blank PCB portion of a commercial air-dam slot filler module with a PCB containing the necessary resistors.  All of the unused slots in the installed sub racks have air-dam modules which are used to evenly distribute the cooling airflow in the sub racks.

#### TGF VME Base Address Selection

The VME interface section of the TGF-II is capable of 16-bit address and data transfers only.  The VME base address is selectable via a four position dip switch labeled ID3 – ID0 as shown below.  For the BPM project, only ID1 will be turned off and all the others will be on.  This gives a base address of 0x2000 where ID3 – ID0 controls the upper 4-bits of the VME address. See Appendix “A” for a complete register map of the TGF-II module.

### TGF-II Firmware

The overall TGF-II firmware will not be changeable remotely.  The board will need to be directly connected to the programmer workstation or laptop to change the firmware.  However the functional programmability is controlled by the registers indicated in the “TGF-II Register Map” section of this document.  These resistors are accessible over the VME bus by the sub rack controller.  All registers are 16 bits wide although the effective width is noted.  The following simplified flow diagram demonstrates the basic operation of the TGF-II firmware.

### TGF-II Flow Diagram of Main Injector

TGF-II Simplified Block Diagram

### TGF-II Register Map (from Offset = 2000h)

| Table 1 | REGISTER | MAP |  |
| --- | --- | --- | --- |
| I/O Address | Register Name | Effective width |  |
| 0x00 | Acquisition Bucket Delay 0 | 12-bits | R/W |
| 0x02 | Acquisition Bucket Delay 1 | 12-bits | R/W |
| 0x04 | Acquisition Bucket Delay 2 | 12-bits | R/W |
| 0x06 | Acquisition Bucket Delay 3 | 12-bits | R/W |
| 0x08 | Acquisition Bucket Delay 4 | 12-bits | R/W |
| 0x0A | Acquisition Bucket Delay 5 | 12-bits | R/W |
| 0x0C | Acquisition Bucket Delay 6 | 12-bits | R/W |
| 0x0E | Acquisition Bucket Delay 7 | 12-bits | R/W |
| 0x10 | Gate Count 0 | 16-bits | R/W |
| 0x12 | Gate Count 1 | 16-bits | R/W |
| 0x14 | Gate Count 2 | 16-bits | R/W |
| 0x16 | Gate Count 3 | 16-bits | R/W |
| 0x18 | Gate Count 4 | 16-bits | R/W |
| 0x1A | Gate Count 5 | 16-bits | R/W |
| 0x1C | Gate Count 6 | 16-bits | R/W |
| 0x1E | Gate Count 7 | 16-bits | R/W |
| 0x20 | MDat Type Matched (Last) | 8-bits | R |
| 0x22 | MDat Value Matched (Last) | 16-bits | R |
| 0x24 | MDat Type Event (interrupt) | 8-bits | R/W |
| 0x26 | Turn Scaler for Background | 16-bits | R/W |
| 0x28 | Diagnostics Counter | 16-bits | Read Only |
| 0x2A | Diagnostic Counter w/ wait |  | Wait on read |
| 0x2C |  |  |  |
| 0x2E | Pre-Trigger Delay All | 16-bits | R/W |
| 0x30 | Control | See Table 5 | R/W |
| 0x32 | Status_Register0 | See Table 6 | Read Only |
| 0x34 | BSync Turn Event | 8-bits | R/W |
| 0x36 | BSync Start Event | See Table 10 | R/W |
| 0x38 | BSync Clear Timestamp Event | See Table 10 | R/W |
| 0x3A | First Turn Timestamp (Wrd0) | Quad Word | Read Only |
| 0x3C | First Turn Timestamp (Wrd1) | Quad Word | Read Only |
| 0x3E | First Turn Timestamp (Wrd2) | Quad Word | Read Only |
| 0x40 | First Turn Timestamp (Wrd3) | Quad Word | Read Only |
| 0x42 | FirstTurnsCount (Low) | Low+High=32 | Read Only |
| 0x44 | FirstTurnsCount (High) | Low+High=32 | Read Only |
| 0x46 | LastTurnsCount (Low) | Low+High=32 | Read Only |
| 0x48 | LastTurnsCount (High) | Low+High=32 | Read Only |
| 0x4A | MissingTurnsCount (Low) | Low+High=32 | Read Only |
| 0x4C | MissingTurnsCount (High) | Low+High=32 | Read Only |
| 0x4E | # of ADCSync0Count Low | Low+High=32 | Read Only |
| 0x50 | # of ADCSync0Count High | Low+High=32 | Read Only |
| 0x52 | Turns Counter Low | Low+High=32 | Read Only |
| 0x54 | Turns Counter High | Low+High=32 | Read Only |
| 0x56 | BSync TimeStamp Event | See Table 10 | R/W |
| 0x58 | BSync TimeStamp Wrd0 | Quad Word | Read Only |
| 0x5A | BSync TimeStamp Wrd1 | Quad Word | Read Only |
| 0x5C | BSync TimeStamp Wrd2 | Quad Word | Read Only |
| 0x5E | BSync TimeStamp Wrd3 | Quad Word | Read Only |
| 0x60 | Irq Vector Reg 1 | See Figure 6 | R/W |
| 0x62 | Irq Vector Reg 2 | See Figure 6 | R/W |
| 0x64 | Last Interrupted TClk Event | 8-bits | Read Only |
| 0x66 | TClk IRQ #Words Pending | 8-bits |  |
| 0x68 | Switch | 16-bits | Read Only |
| 0x6A | Start Event Counter | Note 1 | Read Only |
| 0x70 | IRQ1 Source Mux | See Table 3. | R/W |
| 0x72 | IRQ2 Source Mux | See Table 3. | R/W |
| 0x74 | TClk TimeStamp Event | 8-bits | R/W |
| 0x76 | TClk TimeStamp Wrd0 | Quad Word | Read Only |
| 0x78 | TClk TimeStamp Wrd1 | Quad Word | Read Only |
| 0x7A | TClk TimeStamp Wrd2 | Quad Word | Read Only |
| 0x7C | TClk TimeStamp Wrd3 | Quad Word | Read Only |
| 0x7E | Status_Register1 | 16-bits | Read Only |
| 0x80 | TClk Event0 | See Table 9. | R/W |
| 0x82 | TClk Event1 | See Table 9. | R/W |
| 0x84 | TClk Event2 | See Table 9. | R/W |
| 0x86 | TClk Event3 | See Table 9. | R/W |
| 0x88 | TClk Event4 | See Table 9. | R/W |
| 0x8A | TClk Event5 | See Table 9. | R/W |
| 0x8C | TClk Event6 | See Table 9. | R/W |
| 0x8E | TClk Event7 | See Table 9. | R/W |
| 0x90 | TClk Event8 | See Table 9. | R/W |
| 0x92 | TClk Event9 | See Table 9. | R/W |
| 0x94 | TClk Event10 | See Table 9. | R/W |
| 0x96 | TClk Event11 | See Table 9. | R/W |
| 0x98 | TClk Event12 | See Table 9. | R/W |
| 0x9A | TClk Event13 | See Table 9. | R/W |
| 0x9C | TClk Event14 | See Table 9. | R/W |
| 0x9E | TClk Event15 | See Table 9. | R/W |
| 0xA0 | BSync_P_Event(00_0F) Found | 16-bits | R/W |
| 0xA2 | BSync_P_Event(10_1F) Found | 16-bits | R/W |
| 0xA4 | BSync_P_Event(20_2F) Found | 16-bits | R/W |
| 0xA6 | BSync_P_Event(30_3F) Found | 16-bits | R/W |
| 0xA8 | BSync_P_Event(40_4F) Found | 16-bits | R/W |
| 0xAA | BSync_P_Event(50_5F) Found | 16-bits | R/W |
| 0xAC | BSync_P_Event(60_6F) Found | 16-bits | R/W |
| 0xAE | BSync_P_Event(70_7F) Found | 16-bits | R/W |
| 0xB0 | BSync_P_Event(80_8F) Found | 16-bits | R/W |
| 0xB2 | BSync_P_Event(90_9F) Found | 16-bits | R/W |
| 0xB4 | BSync_P_Event(A0_AF) Found | 16-bits | R/W |
| 0xB6 | BSync_P_Event(B0_BF) Found | 16-bits | R/W |
| 0xB8 | BSync_P_Event(C0_CF) Found | 16-bits | R/W |
| 0xBA | BSync_P_Event(D0_DF) Found | 16-bits | R/W |
| 0xBC | BSync_P_Event(E0_EF) Found | 16-bits | R/W |
| 0xBE | BSync_P_Event(F0_FF) Found | 16-bits | R/W |
|  |  | 16-bits |  |
| 0xC0 | TClkEvent(00_0F) Found | 16-bits | R/W |
| 0xC2 | TClkEvent (10_1F) Found | 16-bits | R/W |
| 0xC4 | TClkEvent (20_2F) Found | 16-bits | R/W |
| 0xC6 | TClkEvent (30_3F) Found | 16-bits | R/W |
| 0xC8 | TClkEvent (40_4F) Found | 16-bits | R/W |
| 0xCA | TClkEvent (50_5F) Found | 16-bits | R/W |
| 0xCC | TClkEvent (60_6F) Found | 16-bits | R/W |
| 0xCE | TClkEvent (70_7F) Found | 16-bits | R/W |
| 0xD0 | TClkEvent (80_8F) Found | 16-bits | R/W |
| 0xD2 | TClkEvent (90_9F) Found | 16-bits | R/W |
| 0xD4 | TClkEvent (A0_AF) Found | 16-bits | R/W |
| 0xD6 | TClkEvent (B0_BF) Found | 16-bits | R/W |
| 0xD8 | TClkEvent (C0_CF) Found | 16-bits | R/W |
| 0xDA | TClkEvent (D0_DF) Found | 16-bits | R/W |
| 0xDC | TClkEvent (E0_EF) Found | 16-bits | R/W |
| 0xDE | TClkEvent (F0_FF) Found | 16-bits | R/W |
| 0xE0 | MDatEvent(00_0F) Found | 16-bits | R/W |
| 0xE2 | MDatEvent (10_1F) Found | 16-bits | R/W |
| 0xE4 | MDatEvent (20_2F) Found | 16-bits | R/W |
| 0xE6 | MDatEvent (30_3F) Found | 16-bits | R/W |
| 0xE8 | MDatEvent (40_4F) Found | 16-bits | R/W |
| 0xEA | MDatEvent (50_5F) Found | 16-bits | R/W |
| 0xEC | MDatEvent (60_6F) Found | 16-bits | R/W |
| 0xEE | MDatEvent (70_7F) Found | 16-bits | R/W |
| 0xF0 | MDatEvent (80_8F) Found | 16-bits | R/W |
| 0xF2 | MDatEvent (90_9F) Found | 16-bits | R/W |
| 0xF4 | MDatEvent (A0_AF) Found | 16-bits | R/W |
| 0xF6 | MDatEvent (B0_BF) Found | 16-bits | R/W |
| 0xF8 | MDatEvent (C0_CF) Found | 16-bits | R/W |
| 0xFA | MDatEvent (D0_DF) Found | 16-bits | R/W |
| 0xFC | MDatEvent (E0_EF) Found | 16-bits | R/W |
| 0xFE | MDatEvent (F0_FF) Found | 16-bits | R/W |
| 0X100 | Filter Card Control Register | See Table 7 | R/W |
| 0X102 | Relay Data, Card 0 | See Table 8 | R/W |
| 0X104 | Relay Data, Card 1 | See Table 8 | R/W |
| 0X106 | Relay Data, Card 2 | See Table 8 | R/W |
| 0X108 | Relay Data, Card 3 | See Table 8 | R/W |
| 0X10A | Relay Data, Card 4 | See Table 8 | R/W |
| 0X10C | Relay Data, Card 5 | See Table 8 | R/W |
| 0X10E | Relay Data, Card 6 | See Table 8 | R/W |
| 0X110 | Relay Data, Card 7 | See Table 8 | R/W |
| 0x112 | FEC_Control_0 | 16-bits | R/W |
| 0x114 | FEC_Control_1 | 16-bits | R/W |
| 0x116 | FEC_Dignostic Signal Control | See Table 11. | R/W |
| 0x120 | BSync_S_Event(00_0F) Found | 16-bits | R/W |
| 0x122 | BSync_S_Event(10_1F) Found | 16-bits | R/W |
| 0x124 | BSync_S_Event(20_2F) Found | 16-bits | R/W |
| 0x126 | BSync_S_Event(30_3F) Found | 16-bits | R/W |
| 0x128 | BSync_S_Event(40_4F) Found | 16-bits | R/W |
| 0x12A | BSync_S_Event(50_5F) Found | 16-bits | R/W |
| 0x12C | BSync_S_Event(60_6F) Found | 16-bits | R/W |
| 0x12E | BSync_S_Event(70_7F) Found | 16-bits | R/W |
| 0x130 | BSync_S_Event(80_8F) Found | 16-bits | R/W |
| 0x132 | BSync_S_Event(90_9F) Found | 16-bits | R/W |
| 0x134 | BSync_S_Event(A0_AF) Found | 16-bits | R/W |
| 0x136 | BSync_S_Event(B0_BF) Found | 16-bits | R/W |
| 0x138 | BSync_S_Event(C0_CF) Found | 16-bits | R/W |
| 0x13A | BSync_S_Event(D0_DF) Found | 16-bits | R/W |
| 0x13C | BSync_S_Event(E0_EF) Found | 16-bits | R/W |
| 0x13E | BSync_S_Event(F0_FF) Found | 16-bits | R/W |
| 0x200 | Acquisition Bucket Delay 0 | 12-bits | R/W |
| 0x202 | Acquisition Bucket Delay 1 | 12-bits | R/W |
| 0x204 | Acquisition Bucket Delay 2 | 12-bits | R/W |
| 0x206 | Acquisition Bucket Delay 3 | 12-bits | R/W |
| 0x208 | Acquisition Bucket Delay 4 | 12-bits | R/W |
| 0x20A | Acquisition Bucket Delay 5 | 12-bits | R/W |
| 0x20C | Acquisition Bucket Delay 6 | 12-bits | R/W |
| 0x20E | Acquisition Bucket Delay 7 | 12-bits | R/W |
| 0x210 | Acquisition Bucket Delay 8 | 12-bits | R/W |
| 0x212 | Acquisition Bucket Delay 9 | 12-bits | R/W |
| 0x220 | Gate Count 0 | 16-bits | R/W |
| 0x222 | Gate Count 1 | 16-bits | R/W |
| 0x224 | Gate Count 2 | 16-bits | R/W |
| 0x226 | Gate Count 3 | 16-bits | R/W |
| 0x228 | Gate Count 4 | 16-bits | R/W |
| 0x22A | Gate Count 5 | 16-bits | R/W |
| 0x22C | Gate Count 6 | 16-bits | R/W |
| 0x22E | Gate Count 7 | 16-bits | R/W |
| 0x230 | Gate Count 8 | 16-bits | R/W |
| 0x232 | Gate Count 9 | 16-bits | R/W |
| 0x240 | Irq Vector Reg 1 | See Figure 6 | R/W |
| 0x242 | Irq Vector Reg 2 | See Figure 6 | R/W |
| 0x244 | Irq Vector Reg 3 | See Figure 6 | R/W |
| 0x246 | Irq Vector Reg 4 | See Figure 6 | R/W |
| 0x248 | Irq Vector Reg 5 | See Figure 6 | R/W |
| 0x24A | Irq Vector Reg 6 | See Figure 6 | R/W |
| 0x24C | Irq Vector Reg 7 | See Figure 6 | R/W |
| 0x24E | Irq Vector Reg 8 | See Figure 6 | R/W |
| 0x250 | IRQ1 Source Mux | See Table 3. | R/W |
| 0x252 | IRQ2 Source Mux | See Table 3. | R/W |
| 0x254 | IRQ3 Source Mux | See Table 3. | R/W |
| 0x256 | IRQ4 Source Mux | See Table 3. | R/W |
| 0x258 | IRQ5 Source Mux | See Table 3. | R/W |
| 0x25A | IRQ6 Source Mux | See Table 3. | R/W |
| 0x25C | IRQ7 Source Mux | See Table 3. | R/W |
| 0x25E | IRQ8 Source Mux | See Table 3. | R/W |
| 0x260 | TClk IRQ Timer Alarm1 | Low 16-bits | R/W |
| 0x262 | TClk IRQ Timer Alarm1 | High 16-bits | R/W |
| 0x264 | TClk IRQ Timer Alarm2 | Low 16-bits | R/W |
| 0x266 | TClk IRQ Timer Alarm2 | High 16-bits | R/W |
| 0x268 | TClk IRQ Timer Alarm3 | Low 16-bits | R/W |
| 0x26A | TClk IRQ Timer Alarm3 | High 16-bits | R/W |
| 0x26C | TClk IRQ Timer Alarm4 | Low 16-bits | R/W |
| 0x26E | TClk IRQ Timer Alarm4 | High 16-bits | R/W |
| 0x270 | TClk IRQ Timer Alarm5 | Low 16-bits | R/W |
| 0x272 | TClk IRQ Timer Alarm5 | High 16-bits | R/W |
| 0x274 | TClk IRQ Timer Alarm6 | Low 16-bits | R/W |
| 0x276 | TClk IRQ Timer Alarm6 | High 16-bits | R/W |
| 0x278 | TClk IRQ Timer Alarm7 | Low 16-bits | R/W |
| 0x27A | TClk IRQ Timer Alarm7 | High 16-bits | R/W |
| 0x27C | TClk IRQ Timer Alarm8 | Low 16-bits | R/W |
| 0x27E | TClk IRQ Timer Alarm8 | High 16-bits | R/W |
| 0x280 | TClk IRQ Timer Alarm9 | Low 16-bits | R/W |
| 0x282 | TClk IRQ Timer Alarm9 | High 16-bits | R/W |
| 0x284 | TClk IRQ Timer Alarm10 | Low 16-bits | R/W |
| 0x286 | TClk IRQ Timer Alarm10 | High 16-bits | R/W |
| 0x288 | TClk IRQ Timer Alarm11 | Low 16-bits | R/W |
| 0x28A | TClk IRQ Timer Alarm11 | High 16-bits | R/W |
| 0x28C | TClk IRQ Timer Alarm12 | Low 16-bits | R/W |
| 0x28E | TClk IRQ Timer Alarm12 | High 16-bits | R/W |
| 0x290 | TClk IRQ Timer Alarm13 | Low 16-bits | R/W |
| 0x292 | TClk IRQ Timer Alarm13 | High 16-bits | R/W |
| 0x294 | TClk IRQ Timer Alarm14 | Low 16-bits | R/W |
| 0x296 | TClk IRQ Timer Alarm14 | High 16-bits | R/W |
| 0x298 | TClk IRQ Timer Alarm15 | Low 16-bits | R/W |
| 0x29A | TClk IRQ Timer Alarm15 | High 16-bits | R/W |
| 0x29C | TClk IRQ Timer Alarm16 | Low 16-bits | R/W |
| 0x29E | TClk IRQ Timer Alarm16 | High 16-bits | R/W |
| 0x2A0 | TClk IRQ Timer Alarm17 | Low 16-bits | R/W |
| 0x2A2 | TClk IRQ Timer Alarm17 | High 16-bits | R/W |
| 0x2A4 | TClk IRQ Timer Alarm18 | Low 16-bits | R/W |
| 0x2A6 | TClk IRQ Timer Alarm18 | High 16-bits | R/W |
| 0x2A8 | TClk IRQ Timer Alarm19 | Low 16-bits | R/W |
| 0x2AA | TClk IRQ Timer Alarm19 | High 16-bits | R/W |
| 0x2AC | TClk IRQ Timer Alarm20 | Low 16-bits | R/W |
| 0x2AE | TClk IRQ Timer Alarm20 | High 16-bits | R/W |
| 0x2B0 | TClk IRQ Timer Alarm21 | Low 16-bits | R/W |
| 0x2B2 | TClk IRQ Timer Alarm21 | High 16-bits | R/W |
| 0x2B4 | TClk IRQ Timer Alarm22 | Low 16-bits | R/W |
| 0x2B6 | TClk IRQ Timer Alarm22 | High 16-bits | R/W |
| 0x2B8 | TClk IRQ Timer Alarm23 | Low 16-bits | R/W |
| 0x2BA | TClk IRQ Timer Alarm23 | High 16-bits | R/W |
| 0x2BC | TClk IRQ Timer Alarm24 | Low 16-bits | R/W |
| 0x2BE | TClk IRQ Timer Alarm24 | High 16-bits | R/W |
| 0x2C0 | TClk IRQ Timer Alarm25 | Low 16-bits | R/W |
| 0x2C2 | TClk IRQ Timer Alarm25 | High 16-bits | R/W |
| 0x2C4 | TClk IRQ Timer Alarm26 | Low 16-bits | R/W |
| 0x2C6 | TClk IRQ Timer Alarm26 | High 16-bits | R/W |
| 0x2C8 | TClk IRQ Timer Alarm27 | Low 16-bits | R/W |
| 0x2CA | TClk IRQ Timer Alarm27 | High 16-bits | R/W |
| 0x2CC | TClk IRQ Timer Alarm28 | Low 16-bits | R/W |
| 0x2CE | TClk IRQ Timer Alarm28 | High 16-bits | R/W |
| 0x2D0 | TClk IRQ Timer Alarm29 | Low 16-bits | R/W |
| 0x2D2 | TClk IRQ Timer Alarm29 | High 16-bits | R/W |
| 0x2D4 | TClk IRQ Timer Alarm30 | Low 16-bits | R/W |
| 0x2D6 | TClk IRQ Timer Alarm30 | High 16-bits | R/W |
| 0x2D8 | TClk IRQ Timer Alarm31 | Low 16-bits | R/W |
| 0x2DA | TClk IRQ Timer Alarm31 | High 16-bits | R/W |
| 0x2DC | TClk IRQ Timer Alarm32 | Low 16-bits | R/W |
| 0x2DE | TClk IRQ Timer Alarm32 | High 16-bits | R/W |
| 0x2E0 | Alarm Counter when IRQ Issued | Low 16-bits | R Only |
| 0x2E2 | Alarm Counter when IRQ Issued | High 16-bits | R Only |
| 0x2E4 | TClk Timer Alarm Counter | Low 16-bits | R Only |
| 0x2E6 | TClk Timer Alarm Counter | High 16-bits | R Only |
|  | ID ROM |  |  |
| 0X400 | x0054, “T” | 8-bit | R only |
| 0X402 | x0047, “G” | 8-bit | R only |
| 0X404 | x0046, “F” | 8-bit | R only |
| 0X406 | x0000, “ “ | 8-bit | R only |
| 0X408 | x0056, “V” | 8-bit | R only |
| 0X40A | x0045, “E” | 8-bit | R only |
| 0X40C | x0052, “R” | 8-bit | R only |
| 0X40E | x0053, “S” | 8-bit | R only |
| 0X410 | x0049, “I” | 8-bit | R only |
| 0X412 | x004F, “O” | 8-bit | R only |
| 0X414 | x004E, “N” | 8-bit | R only |
| 0X416 | x0000, “ “ | 8-bit | R only |
| 0X418 | x0033, “3” | 8-bit | R only |
| 0X41A | x002E, “.” | 8-bit | R only |
| 0X41C | x0031, “1” | 8-bit | R only |
| 0X41E | x002E, “.” | 8-bit | R only |
| 0X420 | x0032, “2” | 8-bit | R only |
| 0X422 | x0020, “ “ | 8-bit | R only |
| 0X500-0X6FE | Last MDat Values Stored | 16-bit | R/W |

Note 1 The Start Event Counter located at offset x”6A” counts the number of BSync start events (offset=x”36”).  This counter is cleared on BSync Clear Timestamp Event (x”38”.

| Table 2 | IRQ SOURCE MUX | I/O Address 0x70...0x72 |  |
| --- | --- | --- | --- |
| DATA | SOURCE |  |  |
| 0x00 | Null |  |  |
| 0x01 | Start Event |  |  |
| 0x02 | Revolution Event |  |  |
| 0x03 | MDAT Type Match |  |  |
| 0x04 | Turn Scaler Trigger Out |  |  |
| 0x05 | TimerAlarm |  |  |
| 0x06 | Delay Timer Trigger |  |  |
| 0x07 | Delay Timer Terminal Count |  |  |
| 0x08 | Turn Counter Trigger |  |  |
| 0x09 | Periodic timer (BGFlash rate) |  |  |
| 0x0A | TClk Event Match | 16-bits | R/W |

Figure 3

| Table 3 | CONTROL AND MODE REGISTER |  | I/O Address 0x30 |
| --- | --- | --- | --- |
| DATA |  |  |  |
| Low Byte |  |  |  |
| 0xXX01 | Control[0] | Enable Pre-Trigger Counter | Zero = Disabled, One = Enabled |
| 0xXX02 | Control[1] | Enable Delay Timer | Zero = Disabled, One = Enabled |
| 0xXX04 | Control[2] | Reset Alarms | Zero = No action, One = Reset |
| 0xXX08 | Control[3] | Not Used |  |
| 0xXX10 | Control[4] | Trigger Source | Zero = PreTrigger from turn event, One = External, Two = PreTrigger from start event |
| 0xXX20 | Control[5] | Trigger Source |  |
| 0xXX40 | Control[6] | “Start” Source | Zero = Start Event, One = Null, Two = Turn Scaler (periodic) |
| 0xXX80 | Control[7] | “Start” Source |  |
| High Byte |  |  |  |
| 0x01XX | Control[8] | Enable MDat |  |
| 0x02XX | Control[9] | Enable TClk | Zero = Disabled, One = Enabled |
| 0x04XX | Control[10] | Enable Beam Sync1 (MI_BS) | Zero = Disabled, One = Enabled |
| 0x08XX | Control[11] | Enable Beam Sync2 (RR_BS) | Zero = Disabled, One = Enabled |
| 0x10XX | Control [12] | Single/Repetitive ADC Sync Out | Zero = Repetitive, One = Single |
| 0x20XX | Control[13] | Enable Beam Extraction Sync (BES) | Zero = Disabled, One = Enabled |
| 0x40XX | Control[14] | Reset TClk IRQ Pending | Zero = No action, One = Reset |
| 0x80XX | Control[15] | Reset TSG | Zero = No action, One = Reset |

Control Register Comments:

Control (2)

When equal to ‘1’ the TClkIrq Alarm Timer state machine will be initialized and wait for the next TClk alarm event.

Control (5:4)

When equal to “00”, the pre-trigger counter is selected as the trigger source for the gate count and bucket delay. The pre-trigger counter decrements on each turn event and when it equals zero, the bucket delay is started, when the bucket delay is decremented to zero, the ADCsync signal is issued and  then the gate count is decremented on the next turn followed by another ADCSync after the bucket delay, etc until the gate count equals zero.

When equal to “01”, an external trigger input bypasses the pre-trigger counter and immediately starts the bucket delay counter, then issues the ADCsync signal, then decrements the gate count every 1113 buckets for the Tevatron (588 buckets for Main Injector).

When equal to “10” ,  a start event also bypasses the pre-trigger counter and immediately starts the bucket delay, then issues the ADCsync signal, then decrements the gate count every 1113 buckets for the Tevatron (588 buckets for Main Injector).

Control (7:6)

When equal to “00”, the Start Source is the BSync start event register (0x36) is equal the decoded BSync serial input.

When equal to “01”, No Start Source.

When equal to“10”, Start is generated when the turn scaler counter (0x26) is equal to zero. The Turn scaler counter is decremented by the decoded BSync Turn event.

Control (12)

When equal to a zero, an ADC_Sync Out pulse is generated for each event trigger (BSync event or external trigger) repetitively.  When equal to a one, an ADC_Sync Out pulse is generated for the first event trigger after writing to the control register. Default is zero for repetitive mode.

| Table 4 | STATUS_0 REGISTER BIT DEFINITIONS | I/O Address 0x32 |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | Turn Marker |  |
| Bit-1 | Start Marker |  |
| Bit-2 | TClk Parity Error |  |
| Bit-3 | MDat Parity Error |  |
| Bit-4 | BSync Parity Error |  |
| Bit-5 |  |  |
| Bit-6 |  |  |
| Bit-7 | Missing Turn Error |  |
| Bit-8 | Turn Scaler Enabled |  |
| Bit-9 | 53MHz present |  |
| Bit-10 | BSync Present |  |
| Bit_11 | TClk Present |  |
| Bit_12 | MDat Present |  |
| Bit_13 | BSync Enabled |  |
| Bit_14 | TClk Enabled |  |
| Bit_15 | MDat Enabled |  |
| Table 6 | STATUS_1 REGISTER BIT DEFINITIONS | I/O Address 0x7E |
| DATA | DESCRIPTION |  |
| Bit-0 | BSync_P_ Enabled |  |
| Bit-1 | BSync_S_ Enabled |  |
| Bit-2 | BES Enabled |  |
| Bit-3 | MDat Enabled |  |
| Bit-4 | TClk Enabled |  |
| Bit-5 | BSync_P_ Present |  |
| Bit-6 | BSync_S_ Present |  |
| Bit-7 | BES_ Present |  |
| Bit-8 | MDat_ Present |  |
| Bit-9 | TClk_ Present |  |
| Bit-10 | 53MHz present |  |
| Bit_11 | Turn Marker |  |
| Bit_12 | Start Marker |  |
| Bit_13 | Turn Scaler Enabled |  |
| Bit_14 | Missing Turn Error |  |
| Bit_15 |  |  |

| Table 7 | FilterCard Control Register BIT DEFINITIONS | I/O Address 0x100 |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | FilterCard Start/Done Card 0 | Note 1 |
| Bit-1 | FilterCard Start/Done Card 1 | Note 1 |
| Bit-2 | FilterCard Start/Done Card 2 | Note 1 |
| Bit-3 | FilterCard Start/Done Card 3 | Note 1 |
| Bit-4 | FilterCard Start/Done Card 4 | Note 1 |
| Bit-5 | FilterCard Start/Done Card 5 | Note 1 |
| Bit-6 | FilterCard Start/Done Card 6 | Note 1 |
| Bit-7 | FilterCard Start/Done Card 7 | Note 1 |
| Bit-8 | FilterCard R/W | Note 2 |
| Bit-9 | FilterCard Enable Diagnostic Signal |  |
| Bit-10 | Diagnostic Signal Type | Note 3 |
| Bit-11 | RelayState0 |  |
| Bit-12 | RelayState1 |  |
| Bit-13 | RelayState2 |  |
| Bit-14 |  |  |
| Bit-15 | FilterCard Global Reset |  |

Note 1:  After loading relay information into the Filter Card Data Registers (0x102 – 0x110) (table 2) and setting a one in the corresponding FilterCard Start/Done l register bit will immediately start sending the serialized data to the appropriate Filter/Relay Card.  When the transfer of data is complete the start bit will be cleared automatically by the FPGA Firmware.

Note 2:  The FilterCard R/W bit (Bit-8) is currently write capable only and read capability will be implemented in the future.  Default is zero for writes to Filter Card.

Note 3:  The Diagnostic Signal Type can be set to a continuous 53MHz clock when set to a logic zero or a 53 MHz clock pulse train of seven pulses triggered from the BSync turn event when set to a logic one.  Default is a continuous 53MHz clock.

| Table 8 | FilterCard Data Register |  |  |
| --- | --- | --- | --- |
|  | Table 8 |  |  |
| Ch # | MSB Data LSB | Source | Destination |
| 1 | xx xx xx xx xx xx xx 00 | BPM | Digitizer |
| 1 | xx xx xx xx xx xx xx 01 | Diagnostic | BPM |
| 1 | xx xx xx xx xx xx xx 10 | Diagnostic | BPM & Digitizer |
| 1 | xx xx xx xx xx xx xx 11 | Diagnostic | Digitizer |
| 2 | xx xx xx xx xx xx 00 xx | BPM | Digitizer |
| 2 | xx xx xx xx xx xx 01 xx | Diagnostic | BPM |
| 2 | xx xx xx xx xx xx 10 xx | Diagnostic | BPM & Digitizer |
| 2 | xx xx xx xx xx xx 11 xx | Diagnostic | Digitizer |
| 3 | xx xx xx xx xx 00 xx xx | BPM | Digitizer |
| 3 | xx xx xx xx xx 01 xx xx | Diagnostic | BPM |
| 3 | xx xx xx xx xx 10 xx xx | Diagnostic | BPM & Digitizer |
| 3 | xx xx xx xx xx 11 xx xx | Diagnostic | Digitizer |
| 4 | xx xx xx xx 00 xx xx xx | BPM | Digitizer |
| 4 | xx xx xx xx 01 xx xx xx | Diagnostic | BPM |
| 4 | xx xx xx xx 10 xx xx xx | Diagnostic | BPM & Digitizer |
| 4 | xx xx xx xx 11 xx xx xx | Diagnostic | Digitizer |
| 5 | xx xx xx 00 xx xx xx xx | BPM | Digitizer |
| 5 | xx xx xx 01 xx xx xx xx | Diagnostic | BPM |
| 5 | xx xx xx 10 xx xx xx xx | Diagnostic | BPM & Digitizer |
| 5 | xx xx xx 11 xx xx xx xx | Diagnostic | Digitizer |
| 6 | xx xx 00 xx xx xx xx xx | BPM | Digitizer |
| 6 | xx xx 01 xx xx xx xx xx | Diagnostic | BPM |
| 6 | xx xx 10 xx xx xx xx xx | Diagnostic | BPM & Digitizer |
| 6 | xx xx 11 xx xx xx xx xx | Diagnostic | Digitizer |
| 7 | xx 00 xx xx xx xx xx xx | BPM | Digitizer |
| 7 | xx 01 xx xx xx xx xx xx | Diagnostic | BPM |
| 7 | xx 10 xx xx xx xx xx xx | Diagnostic | BPM & Digitizer |
| 7 | xx 11 xx xx xx xx xx xx | Diagnostic | Digitizer |
| 8 | 00 xx xx xx xx xx xx xx | BPM | Digitizer |
| 8 | 01 xx xx xx xx xx xx xx | Diagnostic | BPM |
| 8 | 10 xx xx xx xx xx xx xx | Diagnostic | BPM & Digitizer |
| 8 | 11 xx xx xx xx xx xx xx | Diagnostic | Digitizer |

| Table 9 | TClk Event Register BIT DEFINITIONS | I/O Address 0x80-0x9E |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | TClk Event bit 0 |  |
| Bit-1 | TClk Event bit 1 |  |
| Bit-2 | TClk Event bit 2 |  |
| Bit-3 | TClk Event bit 3 |  |
| Bit-4 | TClk Event bit 4 |  |
| Bit-5 | TClk Event bit 5 |  |
| Bit-6 | TClk Event bit 6 |  |
| Bit-7 | TClk Event bit 7 |  |
| Bit-8 | One = UseAlarm Seq , Zero = No Alarm Seq | Note 1 |
| Bit-9 |  |  |
| Bit-10 |  |  |
| Bit-11 |  |  |
| Bit-12 |  |  |
| Bit-13 |  |  |
| Bit-14 |  |  |
| Bit-15 |  |  |

Notes: 1) When set to a logic one and the TClk event occurs, start Timer Alarm and generate a VME interrupt to load the Timer Alarm sequence. When Timer Alarm1 times-out a VME interrupt will be generated, followed by Timer Alarm2, Timer Alarm3, etc until the next Timer Alarm value is zero or the Timer Alarm16 has timed-out. Values loaded into the Timer Alarm registers are elapsed time between alarms not total time.

| Table 10 | Event Register BIT DEFINITIONS | I/O Address 0x34, 0x36, 0x56 |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | Event bit 0 |  |
| Bit-1 | Event bit 1 |  |
| Bit-2 | Event bit 2 |  |
| Bit-3 | Event bit 3 |  |
| Bit-4 | Event bit 4 |  |
| Bit-5 | Event bit 5 |  |
| Bit-6 | Event bit 6 |  |
| Bit-7 | Event bit 7 |  |
| Bit-8 | Sync_Sel0 | Note 1 |
| Bit-9 | Sync_Sel1 | Note 1 |
| Bit-10 |  |  |
| Bit-11 |  |  |
| Bit-12 |  |  |
| Bit-13 |  |  |
| Bit-14 |  |  |
| Bit-15 |  |  |

Notes: 1) When SyncSel = “00” then BSync_P_Event is selected, when SyncSel = “01” then BSync_S_Event is selected and when SyncSel = “10” then BES event is selected.

| Table 11 | Event Register BIT DEFINITIONS | I/O Address 0x112 |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | FEC_Data0 |  |
| Bit-1 | FEC_Data1 |  |
| Bit-2 | FEC_Data2 |  |
| Bit-3 | FEC_Data3 |  |
| Bit-4 | FEC_Data4 |  |
| Bit-5 | FEC_Data5 |  |
| Bit-6 | FEC_Data6 |  |
| Bit-7 | FEC_Data7 |  |
| Bit-8 | FEC_Data8 |  |
| Bit-9 | FEC_Data9 |  |
| Bit-10 | FEC_Data10 |  |
| Bit-11 | FEC_Data11 |  |
| Bit-12 | FEC_Data12 |  |
| Bit-13 | FEC_Data13 |  |
| Bit-14 | FEC_Data14 |  |
| Bit-15 | FEC_Data15 |  |

| Table 12 | Event Register BIT DEFINITIONS | I/O Address 0x114 |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | FEC_Address(0) |  |
| Bit-1 | FEC_Address(1) |  |
| Bit-2 | Spare(0) |  |
| Bit-3 | Spare(1) |  |
| Bit-4 | Spare(2) |  |
| Bit-5 | Spare(3) |  |
| Bit-6 | Spare(4) |  |
| Bit-7 | Spare(5) |  |
| Bit-8 | Spare(6) |  |
| Bit-9 | Spare(7) |  |
| Bit-10 | Spare(8) |  |
| Bit-11 | Spare(9) |  |
| Bit-12 | Read/Write, Read = ‘1’, Write = ‘0’ |  |
| Bit-13 | Not Used |  |
| Bit-14 | Not Used |  |
| Bit-15 | Not Used |  |

| Table 13 | Event Register BIT DEFINITIONS | I/O Address 0x116 |
| --- | --- | --- |
| DATA | DESCRIPTION |  |
| Bit-0 | Diagnostic Signal ½ Bucket Delay bit 0 |  |
| Bit-1 | Diagnostic Signal ½ Bucket Delay bit 1 |  |
| Bit-2 | Diagnostic Signal ½ Bucket Delay bit 2 |  |
| Bit-3 | Diagnostic Signal ½ Bucket Delay bit 3 |  |
| Bit-4 | Diagnostic Signal ½ Bucket Delay bit 4 |  |
| Bit-5 | Diagnostic Signal ½ Bucket Delay bit 5 |  |
| Bit-6 | Diagnostic Signal ½ Bucket Delay bit 6 |  |
| Bit-7 | Diagnostic Signal ½ Bucket Delay bit 7 |  |
| Bit-8 | Diagnostic Signal ½ Bucket Delay bit 8 |  |
| Bit-9 | Diagnostic Signal ½ Bucket Delay bit 9 |  |
| Bit-10 | Diagnostic Signal ½ Bucket Delay bit 10 |  |
| Bit-11 | Diagnostic Signal ½ Bucket Delay bit 11 |  |
| Bit-12 | Enable FEC Diagnostic Signal |  |
| Bit-13 | FEC type Diagnostic Signal= PulseTrain/Continuous |  |
| Bit-14 | Freq of Diagnostic Signal, 53Mhz = ‘0’, 2.5Mhz=’1’ |  |
| Bit-15 |  |  |

Notes: 1) Diagnostic signal delay is used to generate the diagnostic signal delayed from the decoded BSync start event.

2) The FEC diagnostic signal type is selected to a continuous (53Mhz/2.5Mhz) signal when bit-13 = ‘0’ and a (53Mhz/2.5Mhz) pulse train of 84-pulses when bit-13 = ‘1’.

3) The frequency of the diagnostic signal is 53Mhz when bit-14 =’1’ and 2.5Mhz when ‘0’.

### TGF-II Schematics, Layouts and drawings

Schematic: Beams Document #

Layout: Beams Document #

Bill of Materials: Beams Document #

## Sub Rack

The sub rack is a standard 6U VME64 by Dawn VME Products Corp, see Sub rack Specifications.  The 800 watt power supply is built in and provides VME64 standard voltages and the crate has complete fan cooling.  There is a crate monitoring micro controller that can report voltages, temperatures and fan speeds.  It has minimum control over the power supply such that it can turn the outputs on and off and it can reset the VME bus.  It has a serial port and connects to the Accelerator Ethernet network through an Optilogic module that is also used on systems in the Tevatron.

## Sub Rack Processor Module

The sub rack controller CPU board is responsible for establishing the communication link between user applications and the Echotek modules.  All control and data transferred to/from the BPM modules pass through the sub rack controller.

The processor chosen for the Main Injector BPM system is a Motorola MVME5500 with 512MB of on-board memory.  The processor will run the VxWorks embedded operating system which is widely used throughout the lab.  The ESS department in the Computing Division has a MVME 5500 data sheet  on its web site.  Operation of the TeV BPM DAQ software is described in Beams doc #1949, Main Injector Beam Position Upgrade Software Specification for Data Acquisition.

## Sub Rack Cables

Interconnection cables are an important facet of this upgrade.  The signal cables moving sensor outputs around the electronics must not distort, attenuate or add noise to the signals of interest.  Also the signals from the two outputs of each side of a sensor must travel through cables that have the signal delay matched to better than 50 picoseconds.  This is approximately one degree of phase shift at 53 MHz.  The timing cables also have a similar delay matching specification so that clock and gate delays within a sub rack are uniform.  The chosen default cable is double-shielded RG-400.  The short jumpers between the TGF-II and the Echotek DSR are RG-316 because of their flexibility requirement.  The Cable Specifications are Beams Document #1257.

The connectors chosen for the cables are a mixture of SMA and SMB style connectors which have excellent signal characteristics at 53 MHz.  Connections to existing systems use the SMA style.  Where the connectors are too close together, the SMB style connectors are used because they snap on and off rather than requiring a nut to be tightened.

## Installation

The following is a list of the quantity of components installed in the Main Injector service buildings:

Table 5

| Service Building | BPMs | Analog Transition Card | Transition Controller Card | TGF-IIs | FECs | Echoteks |
| --- | --- | --- | --- | --- | --- | --- |

## Cable Specifications

Beams Document #

## Sub rack Specifications

Beams Document #

## Echotek DSR Module

Data Sheet attached.

ECHOTEK CORPORATION 555 SPARKMAN DRIVE #400 HUNTSVILLE, AL 35816 PHONE 256 721 1911 FAX 256 721 9266 E-MAIL: sales@echotek.com

ECDR-GC814

This analog to digital converter and digital drop receiver board provides eight channels of 14 bit, 105 MHz analog to digital conversion and digital processing suitable for wideband and narrowband down conversion and filtering in a single 6U VME slot.  Data decimation range of each receiver channel is user programmable from 8 to 16384.  The receiver section of each channel may be bypassed to output raw A/D data.

###### Data Inputs

Up to eight analog signals may be input to this board via front panel SMA connectors.  Each signal is converted using Analog Devices AD6644, 14 bit, 65 MHz A/D converters or AD6645 for sampling rates to 105 MHz. Direct digitization for IF’s > 200 MHz are supported.  These high quality analog to digital converters exhibit Spur Free Dynamic Ranges in excess of 90 dB FS.

###### Clocks

The A/D clock is provided by the user via front panel SMA connector.  This clock signal (sine wave into 50 ohms) is buffered and distributed to all eight channels so that all channels are sampled simultaneously.

###### Other Inputs

The user may also input a sync signal through the front panel and up to 16 bit wide digital word - normally used to insert a header into the data stream or for tagging data.

###### Receiver Channels

The digitizer outputs from each A/D converter interfaces to the input crossbar switch on the Graychip GC4016 multi-standard quad digital down converter chip.  A receiver channel block diagram is shown below for clarity - the Graychip GC4016 contains four such channels that can be combined as mentioned below, only one Graychip 4016 channel or channel combination is output per receiver channel on the ECDR-GC814.

The ECDR-GC814 supports the various channel combination configurations allowed by the Graychip 4016, that is: Sample Accumulator:

| A | No channels combined. In this mode, each channel supports decimations from 32 to 16384. |
| --- | --- |
| B | Channel pair combinations supports decimations of 16 to 8192 while in this mode |
| C | Four channel combination resulting in one wideband output channel: |
|  | - supports decimations of 8 to 4096 while in this mode. |

The output of the receivers or the A/D converter (in receiver bypass mode) is directed to an FPGA that outputs the sum of a programmed number of samples.  The output may be the sum of 1, 2, 4, 8, 16, 32, 64, 128, or 256 raw A/D samples or I’s or Q’s.

###### FIFO Buffer

Data is output from the sample accumulator to each channel’s FIFO buffer.  The FIFO buffer is 16K x 32 bits (by default) and may be configured at the factory as large as 128K x 32 bits.

###### Data Output

Data, raw A/D output, complex receiver data, or either of the quadrature components, is output via the RACE++ interface.  The data may also be output via the VME64X interface - but this interface was designed to be used for control in real time and, therefore, is not optimized for high speed data transfer.

###### Operating Modes

The ECDR-814 supports both CW and pulsed system applications.  The board may be controlled in one of two basic operating modes:

Gate Mode - Data acquisition occurs when the gate signal is active.  The gate signal may be provided via external front panel input or a software write.

Counted Burst - A preprogrammed number of samples are acquired and processed with each occurrence of an external trigger pulse.  This trigger pulse also has a software bit counterpart associated with it.

###### Set-Up and Control

All set-up and control registers are accessible via the VME interface.  Additionally, the RACE++ interface may be programmed by either VME or RACE++ and can be a RACE++ master or slave.

###### Driver Support

Drivers are supplied for VxWorks Operating Systems.

| Model | Part Number | Ruggedization | Options |
| --- | --- | --- | --- |
| ECDR-GC814/8 | 12-0050/CC334A | Commercial | 65 MHz A/D, 16K FIFO |
| ECDR-GC814/8 | 12-0050/CC335 | Commercial | 65 MHz A/D, 128K FIFO |
| ECDR-GC814/8-80 | 12-0050/CC450 | Commercial | 80 MHz A/D, 16K FIFO |
| ECDR-GC814/8-80 | 12-0050/CC512 | Commercial | 80 MHz A/D, 128K FIFO |
| ECDR-GC814/8-105 | 12-0050/CC495 | Commercial | 105 MHz A/D, 16K FIFO |
| ECDR-GC814/8-105 | 12-0050/CC496 | Commercial | 105 MHz A/D, 128K FIFO |
| ECDR-GC814/4 | 12-0057/CC339 | Commercial | 65 MHz A/D, 16K FIFO |
| ECDR-GC814/4 | 12-0057/CC337 | Commercial | 65 MHz A/D, 128K FIFO |
| ECDR-GC814/4-80 | 12-0057/CC494 | Commercial | 80 MHz A/D, 16K FIFO |
| ECDR-GC814/4-80 | 12-0057/CC451 | Commercial | 80 MHz A/D, 128K FIFO |
| ECDR-GC814/4-105 | 12-0057/CC477 | Commercial | 105 MHz A/D, 16K FIFO |
| ECDR-GC814/4-105 | 12-0057/CC497 | Commercial | 105 MHz A/D, 128K FIFO |
| ECDR-GC814/2 | 12-0058/CC350 | Commercial | 65 MHz A/D, 16K FIFO |
| ECDR-GC814/2 | 12-0058/CC498 | Commercial | 65 MHz A/D, 128K FIFO |
| ECDR-GC814/2-80 | 12-0058/CC420 | Commercial | 80 MHz A/D, 16K FIFO |
| ECDR-GC814/2-80 | 12-0058/CC340 | Commercial | 80 MHz A/D, 128K FIFO |
| ECDR-GC814/2-105 | 12-0058/CC499 | Commercial | 105 MHz A/D, 16K FIFO |
| ECDR-GC814/2-105 | 12-0058/CC500 | Commercial | 105 MHz A/D, 128K FIFO |
| DC COUPLED OPTION: |  |  |  |
| ECDR-GC814/8-DC | 12-0107/CC426 | Commercial | 65 MHz A/D, 16K FIFO |
| ECDR-GC814/8-DC | 12-0107/CC431 | Commercial | 65 MHz A/D, 128K FIFO |

Model Part Number Ruggedization Options

ECDR-GC814/4-DC 12-0108/CC427 Commercial 65 MHz A/D, 16K FIFO

ECDR-GC814/4-DC 12-0108/CC432 Commercial 65 MHz A/D, 128K FIFO

| Ruggedization Levels | Operating Temp (°C) | Storage Temp (°C) | Vibration | Shock | Humidity | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| Commercial | 0 to +55°C with 300ft./ min. airflow | -50 to +85°C | 0.002g2/Hz from 10 to 2000Hz random and 2g sinusoidal from 5 to 500 Hz | 20g peak saw tooth, 11mS duration | Up to 95% RH | Commercial Grade, cooled by blown air, for use in benign environments and software development applications |

| ECDR-GC814/2-DC | 12-0109/CC428 | Commercial | 65 MHz A/D, 16K FIFO |
| --- | --- | --- | --- |
| ECDR-GC814/2-DC | 12-0109/CC433 | Commercial | 65 MHz A/D, 128K FIFO |

# Footnotes

Requirements for the Main Injector BPM upgrade, D. Capista, A. Marchionni

# Text boxes

Sync0 Output

BES Input

TClk Input

MDat Input

Clk9 Output

Clk8 Output

Clk7 Output

Clk6 Output

Clk5 Output

Clk4 Output

Clk3 Output

Clk2 Output

Clk1 Output

Clk0 Output

RRBSync Input

MIBSync Input

RFClk Input

BES

TClk

RFClk Locked

MDat

RRBSync

MIBSync

RfClk

VME Sel

JTAG

Parallel Data Output

5.0V

Start Event

Decoded

?

Figure 1 BPM Front End Functional Block Diagram

3.3V

N

Y

PMC

UCD

Echotek Digitizers

Timing

Generator

CPU

Decrement

Pretrigger Counter

(500Hz)

Scaler

Mode

?

Closed Orbit

Mode

N

Pretrigger

Counter

= 0

?

Scaler Counter

= 0

?

N

AA

Or

PseudoTurn

Load Scaler

Counter

Turn-by-Turn

Mode

Decrement

Scaler Counter

Y

Turn Event

Decoded

?

Start Event

Mode

?

Y

N

Y

Idle

9

8

FEC

Transition  Card

B

u

f

f

e

r

Daughter Card

BES

RRBS

Trigger Out

(0, 1)

TTL

Driver

24ma

Altera  Stratix

EP1S25

BLM Control

EDB

(Parallel TTL)

SMB

Simplified Flow Diagram of

Main Injector TGF-II

Part 1/2

Simplified Flow Diagram of

Main Injector TGF-II

Part 1/2

MIBS

State Machine

Idle

Y

2.

N

Decrement

Gate Count

Gate

Count = 0

?

Issue Echotek

Sync

Y

N

Decrement

Bucket Delay

Counter

Bucket

Delay = 0

?

Load

Bucket Delay

Counter

1.

2.

1.

BPM VME Timing Generator Fanout (TGF-II)

:

:

:

:

:

9

0

Clock

Multiplier

and

FanOut

:

:

:

:

:

9

0

SYNC

OUT

CLOCK

OUT

MDAT

Load

Pretrigger & Gate

Count

N

Y

RFCLK

TCLK

PECL

R

C

V

R

Calibration

Relay Control (PSI)

V

M

E

B

u

f

f

e

r

s

Sync1 Output

Sync2 Output

Sync3 Output

Sync4 Output

Sync5 Output

Sync6 Output

Sync7 Output

Sync8 Output

Sync9 Output