# Run Plan for Commissioning the 325 MHz RF System for the High Intensity Neutrino Source R&D Facility in the Fermilab Meson Detector Building

Bob Webber

3/14/2007

## Introduction

This note outlines the run plan to commission the 325 MHz klystron RF system used in the High Intensity Neutrino Source (HINS) R&D facility in the Meson Detector Building. The plan includes commissioning the klystron RF power source itself, staged commissioning of the RF power distribution system, commissioning the 325 MHz RF Component Test Cage, and commissioning the 325 MHz Cavity Test Cave.

Successful completion of this plan will establish the 325 MHz RF system as an operational system for powering loads in the 325 MHz RF Component Test Cage and the 325 MHz Cavity Test Cave.

## A Three-Part Plan

325 MHz RF system commissioning consists of three parts; each part includes a subset of the full RF power distribution system. The parts are defined as:

- Klystron Commissioning

- RF Component Test Cage Commissioning

- Cavity Test Cave Commissioning

Klystron commissioning must be certified fully or conditionally complete before Parts 2 or 3 may begin. Completion of Part 2 or Part 3 is not a prerequisite for starting the other; however, configuration controls defined below preclude concurrent activity. Required documentation, configuration, and conditions for completion of each part are defined in the sections that follow.

## Klystron Commissioning

Klystron commissioning is the activity of establishing and verifying klystron RF power operation through the circulator, the waveguide switch, and into a dummy load.

### Required configuration for klystron commissioning

Figure 1 is a block diagram of the klystron and RF power distribution system. The RF distribution system shall be locked in a configuration such that RF power is directed only toward the RF load immediately attached to the waveguide switch. Other sections of the RF distribution system shall be secured in a state that will be safe even if RF energy were to be presented.

- The waveguide switch shall be locked in position to direct RF energy to the nearest RF load.

- The waveguide shutter shall be locked in closed position.

- The coaxial RF port in the RF Component Test Cage shall be blanked off with a conducting plate.

- The coaxial switch in the 3” coaxial line shall be locked in position to direct RF energy to the immediately attached RF load.

- Any section of the 3” coaxial RF distribution system attached to the remaining coaxial switch port shall be blanked off at the end(s) with a conducting plate.

Figure 1. RF Power Distribution System

### Prerequisites to begin klystron commissioning

Documentation:

Approved modulator, klystron, and RF distribution system LOTO procedures

RF and Klystron Equipment Interlock Test Procedure/ Checklists

Klystron commissioning procedure

Tasks the must be completed:

Execution of Equipment Interlock Test Checklist – D. Bollinger

Final Area/Equipment Walkthru – M. Andrews, R. Pasquinelli, J. Reid

Verification of required RF power distribution system configuration – R. Webber or designee

Approvals to proceed by:

Klystron manufacturer (Toshiba) representative

Modulator system engineer – C. Jensen

Klystron and RF Power system engineer – A. Moretti

Klystron equipment interlock system engineer – D. McDowell

RF equipment interlock system engineer – P. Prieto

HINS Meson Program Leader – R. Webber

HINS Meson Safety Coordinator – M. Andrews

AD ES&H – J. Anderson or designee

AD Division Headquarters – R. Dixon or designee

### Personnel required present for klystron commissioning

Chris Jensen and/or Howie Pfeffer – modulator specialists

Al Moretti – RF engineer

Peter Prieto – RF interlocks engineer

AD ES&H Dept. designee – for monitoring radiated electromagnetic energy

AD ES&H Dept. designee – for measuring x-ray levels

### Definition of Completion of Klystron commissioning

Full completion:

Klystron operation has been successfully demonstrated at >2 MW peak pulse power at > 1 millisecond RF pulse length and with maximum duty cycle supported by the presently installed modulator charging power supply.

AND under those conditions:

- RF equipment interlocks have been demonstrated to function satisfactorily.

- RF leakage from klystron and RF power distribution system has been verified as acceptable within guidelines of Fermilab ES&H and Radiological Control Manuals and IEEE Standard C95.1.

- X-ray dose rates from klystron are verified as acceptable within standards established by Fermilab ES&H and Radiological Control Manuals.

AND: Approved klystron/modulator operational procedures are available.

Conditional completion:

Conditional completion, sufficient to begin RF Component Test Cage and Cavity Test Cave Commissioning, may be accepted by Program Leader and AD ES&H when:

Klystron operation has been successfully demonstrated at >1 MW peak pulse power at > 0.1 millisecond RF pulse length and with maximum duty cycle supported by the presently installed modulator charging power supply.

AND under those conditions:

- RF equipment interlocks have been demonstrated to function satisfactorily.

- RF leakage from klystron and RF power distribution system has been verified as acceptable within guidelines of IEEE Standard C95.1.

- X-ray dose rates from klystron are verified as acceptable within standards established by Fermilab ES&H and Radiological Control Manuals.

AND conditions are established and accepted to constrain operations to acceptable peak and average power limits for the purpose of commencing limited RF Component Test Cage and Cavity Test Cave Commissioning activities.

AND: Approved klystron/modulator operational procedures are available.

-----  Similar to above to be included for RF Cage and Cavity Cave commissioning.  -----

# Approval Sheet to Execute

# Run Plan for Commissioning the 325 MHz RF System for the High Intensity Neutrino Source R&D Facility in the Fermilab Meson Detector Building

# Text boxes

Waveguide Shutter

Coax Switch

2.5 MW

250 kW

to Cavity Test Cave

25 kW

10 db Coax Coupler

Load

Load

10 db Waveguide Coupler

to RF Component Test Area Cage

Load

2.5 MW ultimately to RFQ and Linac Cave Switch

Load

Waveguide Switch

Circulator

Load

2.5 MW

Klystron