# Configuration Control of HINS Waveguide Switch for Protection of RF Loads in RFQ Distribution during High Power Operation

## Purpose:

This document describes the procedure for protecting the coaxial loads of the HINS RF distribution system from over-powering. These coaxial loads act as substitutes for the RFQ until it arrives. The klystron is capable of 2.5 MW of RF power, but the coaxial loads can only handle 600kW each, for a total limit of 1.2 MW.  An attenuator on the output of the RF driver amplifier for the klystron limits the power output of the klystron to 1.1 MW.  However, there may be applications where the full power of the klystron is required, such as cavity testing.  This document describes the procedure for increasing the available power of the klystron while protecting the coaxial loads.

## Configuring for High Power:

- Verify that the klystron is off, or follow the procedure for turning the klystron off.

- Move the waveguide switch into the “Test Area” position.

- Verify the waveguide switch is in the correct position by examining the switch shaft.

- Turn off power to the waveguide switch motor.

- Lock the power switch in the off position using the configuration control lock.

- Store the key to the lock in the key tree located on the south wall of the test cage.

- Unplug the RF driver amplifier.

- Change the output attenuation from 18dB to 12dB.

- Plug in the RF driver amplifier.

The klystron may now be operated up to its 2.5 MW limit.

## Return to Power Limit Conditions:

- Verify that the klystron is off, or follow the procedure for turning the klystron off.

- Unplug the RF driver amplifier.

- Change the output attenuator from 12dB to 18dB (see instructions below), or verify that the extra 6dB attenuator is connected.

- Plug in the RF driver amplifier.

- Retrieve the configuration control lock key from the key tree located on the south wall of the test cage.

- Unlock the waveguide switch motor power switch.

- Turn the power switch on.

The klystron may now power the coaxial line and load safely.

## Connection of 18dB Attenutation:

- Connect  RF elbow to output of 12dB attenuator.

- Connect 1W 6dB attenuator to  RF elbow.

- Connect input of 20dB directional coupler to output of 6dB attenuator.