## Slide 1: Electron Beam Profiler for the Main Injector
Randy Thurman-Keup
Instrumentation Department
APT Seminar
17 June 2014

## Slide 2: Fellow Conspirators
Instrumentation
  Amber Johnson, Carl Lundberg, Jim Galloway, Jim Fitzgerald, Peter Prieto, John van Bogaert, Andrea Saewart, Dave Slimmer, Dehong Zhang, Brian Fellenz, Alex Lumpkin
Mechanical Support
  Wade Muranyi, Brad Tennis, Elias Lopez, Debbie Bonifas, Scott McCormick, Ryan Montiel, Sali Sylejmani, Tom Lassiter, James Williams, John Sobolewski, Matt Alvarez, Kevin Duel
Summer Students
  Paul Butkovich, Khalida Hendricks, Danila Nikiforov
APC
  Charles Thangaraj

## Slide 3: Motivation
The long range plan for Fermilab calls for large proton beam power in excess of 2 MW for use in the neutrino program
Higher proton intensities are problematic for profile diagnostics that physically intercept the beam

## Slide 4: Damage Montage
DESY electrons
YAG:Ce
60 keV electrons
SS OTR mirror
Tevatron Collimator
NuMI OTR Al-coated Kapton foil
~ 6.5e19 120 GeV protons
3 mil Ti vacuum window
1020 120 GeV protons.
ZrO2:Al
Ø 30 mm
GSI heavy ions
(from Beata Walasek-Höhne)
Broken Flying Wire micrograph

## Slide 5: Motivation
The long range plan for Fermilab calls for large proton beam power in excess of 2 MW for use in the neutrino program
Higher proton intensities are problematic for profile diagnostics that physically intercept the beam
Hence the goal of non-intercepting profile diagnostics
  Laser Based (need electrons; either e beam or H-)
  Ionization Profile Monitors (IPM)
  Gas Fluorescence Detectors
  Gas Jets
  Probe Beams

## Slide 6: Probe Beam Concept
Deflection vs. Impact parameter provides information about the charge distribution in the direction of the impact parameter
Charge Distribution
Probe beam
Impact parameter
Deflection
Probe beam is deflected by electric and/or magnetic fields of a charge distribution

## Slide 7: Probe Beam History
Beam probe for plasma distribution
  Paul D. Goldan, Collisionless Sheath – An Experimental Investigation, Phys. Fluids 13 1055 (1970).
  C.H. Stallings, Electron Beam as a Method of Finding the Potential Distribution in a Cylindrically Symmetric Plasma, J. Appl. Phys. 42 (1971) 2831. electron beam
  C.W. Mendel Jr., Apparatus for measuring rapidly varying electric fields in plasmas, Rev. Sci. Instrum. 46 847 (1975). He+ ion beam
Beam probes for other beams
  J. Shiloh, et al., Electron beam probe for charge neutralization studies of heavy ion beams, Rev. Sci. Instrum. 54 (1983) 46.
  V. Shestak, et al., Electron Beam Probe for Ion Beam Diagnostics, TRIUMF Design Note, TRI-DN-87-36 (1987).
  P. Gross, et al., An Electron Beam Probe for Ion Beam Diagnosis, in proceedings of the European Particle Accelerator Conference 1990, p. 806, 12 – 16 June 1990, Nice, France.
  J. Bosser, et al., Transverse Profile Monitor using Ion Probe Beams, Nucl. Instrum. Methods Phys. Res. A 484 (2002) 1. Xe+ ion beam curtain
  P.V. Logatchov, et al., Non-Destructive Singlepass Monitor of Longitudinal Charge Distribution in an Ultrarelativistic Electron Bunch, in proceedings of the Particle Accelerator Conference 1999. electron beam @ VEPP-3

## Slide 8: Theory
Assume , no magnetic field,
Assume deflection is very small such that
Assume again that deflection is very small such that and
x
y
b
q(b)
Beam
x profile

## Slide 9: Reality
The beam has magnetic fields
  Sideways deflection of the probe beam
  Sideways deflection varies with longitudinal shape
The bunch does not have infinite length
  Varying longitudinal shape will alter deflection
    Both electrostatically and magnetically
Deflection may not be all that small
External magnetic fields
Measurement artifacts, etc…

## Slide 10: SNS Device
W. Blokland, 9th DITANET Topical Workshop, April 2013

## Slide 11: Techniques
Collaborating with Wim Blokland at SNS who has done simulations of the various techniques
Possible techniques for measuring deflection
  Fast scan through peak of bunch
    Requires fast deflector (< 1 ns sweep time)
  Slow scan, akin to flying wires
    Position the beam and record the maximum deflection as the beam passes by
      Leave the electron beam stationary
      Sweep the beam along the proton direction
        Obtain longitudinal distribution
        Probably what we will start with

## Slide 12: Fast Scan
x
x
x
x
x
x
x
x
x
x
x
x
x
x
x
Y
Z
X
Y
Z
Proton Beam
Electron Beam Above
Electron
Beam Below
x
x
x
x
x
x
x
x
x
x
x
x
x
x
x
Y
Z
If scan time is too slow longitudinal and transverse charge distributions become entangled

## Slide 13: Techniques
Collaborating with Wim Blokland at SNS who has done simulations of the various techniques
Possible techniques for measuring deflection
  Fast scan through peak of bunch
    Requires fast deflector (< 1 ns sweep time)
  Slow scan, akin to flying wires
    Position the beam and record the maximum deflection as the beam passes by
      Leave the electron beam stationary
      Sweep the beam along the proton direction
        Obtain longitudinal distribution
        Probably what we will start with

## Slide 14: Slow Electron Scan
Plots courtesy of Wim Blokland
Stationary Beam
Position the electron beam
Record the deflection of a bunch
Move the electron beam and repeat

## Slide 15: Slow Electron Scan Simulation
Plots courtesy of Wim Blokland
Step the electron beam through the proton beam and record maximum deflections
Derivative of deflection vs. position is nominally beam profile
Derivative

## Slide 16: Techniques
Collaborating with Wim Blokland at SNS who has done simulations of the various techniques
Possible techniques for measuring deflection
  Fast scan through peak of bunch
    Requires fast deflector (< 1 ns sweep time)
  Slow scan, akin to flying wires
    Position the beam and record the maximum deflection as the beam passes by
      Leave the electron beam stationary
      Sweep the beam along the proton direction
        Obtain longitudinal distribution
        Probably what we will start with

## Slide 17: Pseudo-fast plus Slow Scan
Sweep the electron beam along the proton bunch
Sweep duration coincides with the duration of the proton bunch
Magnetic field of beam distorts measurement
Beam Simulated Longitudinal s = 2 ns
Measured Simulated Longitudinal s = 2.3 ns
Better background gives s = 2.1 ns
Background fit not so good
Electron Sweep
Proton Beam
Simulation

## Slide 18: Simulation
Fields of proton beam are evaluated on a grid
Electron beam is steered by electrostatic deflector
  Fields are calculated in 2D via Poisson
Electrons are tracked through the fields
  Initial electron beam parameters taken from test stand measurements
  Tracking is done via MATLAB code

## Slide 19: Reconstruction
Beam Sim. Longitudinal s = 2 ns
Meas. Sim. Longitudinal s = 2.3 ns
Beam Simulated Transverse s = 3 mm
Meas. Simulated Transverse s = 3.5 mm

## Slide 20: Electron Gun
Commercial source: Kimball Physics electron gun
  Model EGH-6210
  Typically designed for electron microscopes
  LaB6 cathode, up to 60 KeV, 6 mA gateable, <100mm spot size

## Slide 21: Phase 1 Test Stand
YAG or OTR Screens
Electron Gun
Lens / Digital Camera Imaging Systems
Faraday Cup

## Slide 22: Gun Tests
Gun has internal solenoid
  Scanned beam through waist at first screen
Scanned beam sizes from Ce:YAG screens (1 A beam)
Scanned beam sizes from OTR screens (1 mA beam)

## Slide 23: Phase 2 Test Stand
Stretched Wires
Single OTR Port
Hoped to simulate beam with stretched wires…

## Slide 24: Wire Test
Wire to simulate proton beam
e Beam pulsed on for 40 ms
Wire pulsed for 20 ms
Half the time the beam is deflected
0V
150V
50V
250V
200V
300V
100V

## Slide 25: Test of Electrostatic Deflector
Deflector Pulse
Deflector Pulse
15 cm long plates
~120 V
~190 V
Deflecting Voltage vs. Deflector Length
500 V
80 ns

## Slide 26: Electrostatic Deflector Test
Short sweep
Effect is similar to proton bunch passing by
Longer sweep
Bright part off screen
Beam size not uniform
  Possibly due to poor pulse quality

## Slide 27: Electron Device
Ion Pump
60 keV Electron Gun
Kimball Physics
Pneumatic Beam Valve
Electrostatic Deflector
Ion Gauge
Ion Gauge
Pneumatic
Insertion Device
with OTR Stainless
Steel Mirror
Phosphor Screen
Optical Breadboard
~ 60 cm x 150 cm
Main Injector
beampipe
Optical components box

## Slide 28: Devices
Solenoid and steering magnets
Cathode
Thermionic Triode Electron Gun
Electrostatic Deflector
  Kimball Physics EGH-6210 up to 60 keV
(we will use up to 15 keV for Nova)
6 mA, pulsed, 2 ms to DC @ 1 kHz
LaB6 cathode, 100 mm spot size
15 cm long ‘circular’ plates
~2.5 cm diameter
Plates

## Slide 29: Devices
Beam Imaging Systems, Phosphor Screen
P47 (Y2SiO5:Ce3+), 400 nm, 60 ns decay, 0.055 quantum yield (photons/eV/electron)
Conductive coating with drain wire
4” Huntington Pneumatic Actuator
SS Mirror for OTR (calibrate electron beam size @ proton beam location)

## Slide 30: OTR Screen
Light yield over the 2 ms electron pulse
Initial beam images determined to be blackbody
  No polarization
  Intensity increased nonlinearly with duration
  Damage to stainless steel mirror observed
Electron energy low
  Broad angular distribution
  Mirror should be 15 instead of 45
(E. Bravin, private communication)

## Slide 31: Optical Acquisition
Calibration OTR
Phosphor
Image Intensifier
Megarad CID camera plus C-mount objective lens
Motorized Stage
Motorized Stage
f = 40 mm
Selectable Neutral Density
Filters (ND 1,2,3) and
Ver / Hor Polarizers
f = 40 mm
f = 125 mm
Mirror on Motorized Stage selects OTR or Phosphor
RS-170 video capture via computer in service building

## Slide 32: Optics

## Slide 33: Full Device

## Slide 34: Install Location
MI 620
Electron Gun

## Slide 35: MI-62 Service Building
Reusing kicker cables to bring electron gun voltages to tunnel
Also reusing flying wire cables

## Slide 36: High Voltage Distribution
e Gun Controller
Service Bldg Transition Box
Has all the fancy controls
Custom Cable
Common (HV)
Filament+
Filament-
Grid
Interlock (not HV)
RG-220
to Tunnel
Vacuum relay
Displays
Manual lockout
p Beam interlock
Interlock
Interlock
in Tunnel
Tunnel Transition Box
RG-220
Interlock
Custom Cable
e Gun
Vacuum relay w/ door switch(?)

## Slide 37: Magnetic Fields are a Problem
No field
5 G along beam, 2 G transverse

## Slide 38: Magnetic Fields in Tunnel
Electron beam
Bhorizontal
Bvertical
2 Gauss
0 G
Quad busses
3500 A
Dipole busses
9000 A
CST Simulation
Lower Dipole bus goes in proton direction
Quad bus closest to beam is defocusing bus and goes in direction of protons

## Slide 39: From e cloud Measurements
From Michael Backfish thesis

## Slide 40: Mumetal Wrapping
Cover “everything” with 1 or more layers of mu metal

## Slide 41: Mumetal Test
Mumetal to enclose Hall Probe
Dipole from A0
With 31 Gauss
3 layers of mumetal reduced the field to 0.2 - 0.4 Gauss
4 - 5 layers knocked it down to 0 - 0.1 Gauss

## Slide 42: CST Simulation of Mumetal
Horizontal B field
  Green is 0 G
2.6 G
-2.6 G
Slice through center of Mu metal transverse to proton beam
B vs H

## Slide 43: CST Simulation of Mumetal
Fields along central electron path

## Slide 44: Optics Simulation
Check
magnification
Outer edge of phosphor
Pattern
Image on
Intensifier
Outer edge of Intensifier
Check
acceptance
Uniform
Source on
phosphor
Uniform
Image on
intensifier

## Slide 45: Summary
Gun mounted in stand
Leak checked (twice)
Cables pulled from MI-62 to device location
  Reused Flying wire cables and Kicker RG-220s
HV Distribution and interlocks being built
Recently reviewed
Plan to install in September shutdown
More studies of magnetic shielding
More studies of measurement systematics

## Slide 46: Questions?

## Slide 47: Vacuum Topology
Differential pumping in gun
Ion pump on cathode side
Nothing on this side except MI
Have another 30 l/s pump

## Slide 48: Gun Internals

## Slide 49: Pneumatics
Input
Solenoid Valves
Beam Valve
OTR Actuator

## Slide 50: Compressed Air

## Slide 52: Simulated Camera Image
Camera frames are ~30 ms
Main Injector cycle is ~1 s
Need to step many times per frame to accumulate data fast enough for measurement
  Complicated to extract each step