## Slide 1
John Kuharik
Booster 2nd Harmonic Cavity Meeting
March 30, 2017
Garnet Witness Sample Testing Lot 62800 Results

## Slide 2: Witness Sample Testing Lot 62800, Samples 1-16
16 samples from lot #62800 tested.
6 runs each.
Sample removed and flipped over between each run.
SG, amplifier, scope at the same settings for every sample/run.
Measurements made over two days.
  Samples 1-12 on 3/15
  Samples 13-16 on 3/16
Samples tested with 19.1 Amp peak excitation current.
Results include corrections previously discussed for air gap, solenoid turn density.
Sample 2, run 1

## Slide 3: Sample 1
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.91 | 0.03 | 0.01 | 4.55 | 0.34 | 0.15
1500 | 1499.91 | 0.03 | 0.01 | 2.13 | 0.06 | 0.03
2000 | 1999.91 | 0.03 | 0.01 | 1.68 | 0.03 | 0.01
2500 | 2499.91 | 0.03 | 0.01 | 1.50 | 0.02 | 0.01
3164 | 3163.91 | 0.03 | 0.01 | 1.35 | 0.02 | 0.01

## Slide 4: Sample 2
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.92 | 0.03 | 0.01 | 4.46 | 0.27 | 0.12
1500 | 1499.92 | 0.03 | 0.01 | 2.13 | 0.05 | 0.02
2000 | 1999.92 | 0.03 | 0.01 | 1.68 | 0.03 | 0.01
2500 | 2499.92 | 0.03 | 0.01 | 1.50 | 0.02 | 0.01
3164 | 3163.92 | 0.03 | 0.01 | 1.36 | 0.02 | 0.01

## Slide 5: Sample 3
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.92 | 0.04 | 0.02 | 4.65 | 0.74 | 0.33
1500 | 1499.92 | 0.04 | 0.02 | 2.15 | 0.10 | 0.04
2000 | 1999.92 | 0.04 | 0.02 | 1.68 | 0.04 | 0.02
2500 | 2499.92 | 0.04 | 0.02 | 1.51 | 0.03 | 0.01
3164 | 3163.92 | 0.04 | 0.02 | 1.36 | 0.02 | 0.01

## Slide 6: Sample 4
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.91 | 0.02 | 0.01 | 4.48 | 0.27 | 0.12
1500 | 1499.91 | 0.02 | 0.01 | 2.13 | 0.04 | 0.02
2000 | 1999.91 | 0.02 | 0.01 | 1.68 | 0.02 | 0.01
2500 | 2499.91 | 0.02 | 0.01 | 1.50 | 0.01 | 0.01
3164 | 3163.91 | 0.02 | 0.01 | 1.36 | 0.01 | 0.005

## Slide 7: Sample 5
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.91 | 0.03 | 0.01 | 4.43 | 0.25 | 0.11
1500 | 1499.91 | 0.03 | 0.01 | 2.11 | 0.05 | 0.02
2000 | 1999.91 | 0.03 | 0.01 | 1.66 | 0.03 | 0.01
2500 | 2499.91 | 0.03 | 0.01 | 1.49 | 0.02 | 0.01
3164 | 3163.91 | 0.03 | 0.01 | 1.34 | 0.02 | 0.01

## Slide 8: Sample 6
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.91 | 0.02 | 0.01 | 4.65 | 0.47 | 0.21
1500 | 1499.91 | 0.02 | 0.01 | 2.15 | 0.07 | 0.03
2000 | 1999.91 | 0.02 | 0.01 | 1.69 | 0.03 | 0.01
2500 | 2499.91 | 0.02 | 0.01 | 1.51 | 0.02 | 0.01
3164 | 3163.91 | 0.02 | 0.01 | 1.36 | 0.02 | 0.01

## Slide 9: Sample 7
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.92 | 0.04 | 0.02 | 4.48 | 0.26 | 0.11
1500 | 1499.92 | 0.04 | 0.02 | 2.13 | 0.04 | 0.02
2000 | 1999.92 | 0.04 | 0.02 | 1.68 | 0.02 | 0.01
2500 | 2499.92 | 0.04 | 0.02 | 1.50 | 0.01 | 0.01
3164 | 3163.92 | 0.04 | 0.02 | 1.36 | 0.01 | 0.005

## Slide 10: Sample 8
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.94 | 0.03 | 0.01 | 4.28 | 0.37 | 0.17
1500 | 1499.94 | 0.03 | 0.01 | 2.10 | 0.06 | 0.03
2000 | 1999.94 | 0.03 | 0.01 | 1.66 | 0.03 | 0.01
2500 | 2499.94 | 0.03 | 0.01 | 1.49 | 0.02 | 0.01
3164 | 3163.94 | 0.03 | 0.01 | 1.35 | 0.01 | 0.01

## Slide 11: Sample 9
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.93 | 0.03 | 0.01 | 4.43 | 0.20 | 0.09
1500 | 1499.93 | 0.03 | 0.01 | 2.12 | 0.04 | 0.02
2000 | 1999.93 | 0.03 | 0.01 | 1.68 | 0.02 | 0.01
2500 | 2499.93 | 0.03 | 0.01 | 1.50 | 0.02 | 0.01
3164 | 3163.93 | 0.03 | 0.01 | 1.36 | 0.01 | 0.01

## Slide 12: Sample 10
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.91 | 0.03 | 0.01 | 4.82 | 0.09 | 0.04
1500 | 1499.91 | 0.03 | 0.01 | 2.17 | 0.01 | 0.01
2000 | 1999.91 | 0.03 | 0.01 | 1.70 | 0.01 | 0.003
2500 | 2499.91 | 0.03 | 0.01 | 1.51 | 0.004 | 0.002
3164 | 3163.91 | 0.03 | 0.01 | 1.36 | 0.003 | 0.002

## Slide 13: Sample 11
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.90 | 0.03 | 0.01 | 5.08 | 0.61 | 0.27
1500 | 1499.90 | 0.03 | 0.01 | 2.21 | 0.08 | 0.04
2000 | 1999.90 | 0.03 | 0.01 | 1.71 | 0.04 | 0.02
2500 | 2499.90 | 0.03 | 0.01 | 1.52 | 0.03 | 0.01
3164 | 3163.90 | 0.03 | 0.01 | 1.37 | 0.02 | 0.01

## Slide 14: Sample 12
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.91 | 0.04 | 0.02 | 4.64 | 0.77 | 0.34
1500 | 1499.91 | 0.04 | 0.02 | 2.15 | 0.11 | 0.05
2000 | 1999.91 | 0.04 | 0.02 | 1.68 | 0.05 | 0.02
2500 | 2499.91 | 0.04 | 0.02 | 1.51 | 0.03 | 0.01
3164 | 3163.91 | 0.04 | 0.02 | 1.36 | 0.02 | 0.01

## Slide 15: Sample 13
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.93 | 0.02 | 0.01 | 4.37 | 0.23 | 0.10
1500 | 1499.93 | 0.02 | 0.01 | 2.11 | 0.04 | 0.02
2000 | 1999.93 | 0.02 | 0.01 | 1.66 | 0.02 | 0.01
2500 | 2499.93 | 0.02 | 0.01 | 1.49 | 0.02 | 0.01
3164 | 3163.93 | 0.02 | 0.01 | 1.35 | 0.01 | 0.01

## Slide 16: Sample 14
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.93 | 0.04 | 0.02 | 4.81 | 1.17 | 0.53
1500 | 1499.93 | 0.04 | 0.02 | 2.16 | 0.14 | 0.06
2000 | 1999.93 | 0.04 | 0.02 | 1.69 | 0.06 | 0.03
2500 | 2499.93 | 0.04 | 0.02 | 1.51 | 0.04 | 0.02
3164 | 3163.93 | 0.04 | 0.02 | 1.36 | 0.03 | 0.01

## Slide 17: Sample 15
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.93 | 0.02 | 0.01 | 4.84 | 0.26 | 0.11
1500 | 1499.93 | 0.02 | 0.01 | 2.18 | 0.04 | 0.02
2000 | 1999.93 | 0.02 | 0.01 | 1.70 | 0.02 | 0.01
2500 | 2499.93 | 0.02 | 0.01 | 1.51 | 0.01 | 0.01
3164 | 3163.93 | 0.02 | 0.01 | 1.36 | 0.01 | 0.01

## Slide 18: Sample 16
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.92 | 0.01 | 0.004 | 4.72 | 0.36 | 0.16
1500 | 1499.92 | 0.01 | 0.004 | 2.16 | 0.05 | 0.02
2000 | 1999.92 | 0.01 | 0.004 | 1.69 | 0.02 | 0.01
2500 | 2499.92 | 0.01 | 0.004 | 1.51 | 0.02 | 0.01
3164 | 3163.92 | 0.01 | 0.004 | 1.36 | 0.01 | 0.01

## Slide 19: Samples 1 – 5 Compared to Lot Average

## Slide 20: Sample 6 – 10 Compared to Lot Average

## Slide 21: Samples 11 – 16 Compared to Lot Average

## Slide 22: Samples All Summary
 | Bmean | Bstd | Bsterr | MuMean | Mustd | Musterr
1000 | 999.92 | 0.03 | 0.01 | 4.61 | 0.50 | 0.22
1500 | 1499.92 | 0.03 | 0.01 | 2.14 | 0.07 | 0.03
2000 | 1999.92 | 0.03 | 0.01 | 1.68 | 0.03 | 0.01
2500 | 2499.92 | 0.03 | 0.01 | 1.50 | 0.02 | 0.01
3164 | 3163.92 | 0.03 | 0.01 | 1.36 | 0.02 | 0.01