# Numi BPM Test Stand Measurements

The Beam Position Monitors (BPMs) used for the Numi beam line incorporate one pair of "split plate" electrodes.  Two sizes with a similar design are used, for the Numi beam line.  A small aperture (2 in Nom) referred to as the "Numi Target BPM" and a larger (3.875 Nom), BPM referred to as the "Numi Large (MI8) BPM". Both have an overall length of 8.0 in.  The two port BPMs are sensitive to beam position in only one plane, and therefore require that they be installed with the mechanical orientation to match the appropriate plane.  The BPMs have a thick (.375 in.) "Alignment plate" attached at both ends, which are used to align the BPM in a test stand and to provide fiducials for alignment in the beam lines. Refer to file "Numi Large BPM Data Sheet" and "Numi Target BPM Data Sheet" for more design details.

The BPM are tested on a "BPM Test Stand" for performance and as a quality control check before they are installed in the beam line.  The Test Stand is a semi-automated fixture that uses an RF signal on a wire in the aperture to simulate beam. It uses stepping motor tables to move the wire through a programmed grid pattern of positions in the aperture.  An HP Network Analyzer is used to measure the relative amplitude of the signal on the two ports at each grid point.  A PC using LabVIEW software operates the test stand controlling the motion tables, the Network Analyzer and generating a text file of the data.  In the case of the Numi BPMs an average of five measurements were preformed by the network analyzer at each point.  The data is then transferred to an Excel spreadsheet for analysis.  BPM sensitivity (or gain) and zero offset values are produced by the spreadsheets.

The test stand uses precision reference plates, holes and pins to locate each style of BPM.  An effort is made before testing, to align the zero position of the test wire to the mechanical center of the BPM.  Misalignment of the test wire or mechanical mounting deviations can be a source of offset measurement error or test stand bias.  Another source of bias can be caused by an impedance mismatch between the test wire and BPM aperture.  Matching transformers are used in an effort to match the 50-ohm signal cable impedance to higher wire-BPM impedance.  However, as the BPMs come in various aperture sizes an optimum match is not always available.  An impedance mismatch can cause RF standing waves on the test wire, which, in the case of the split plate BPM, can cause a signal "RF bias" to one side of the BPM.

In an effort to overcome possible errors produced by any "test stand bias" or an "RF bias" a routine was developed where the BPMs are measured in four orientations.  The BPMs are measured in a standard orientation and then a second time rotated 180 deg.  A mechanical bias caused by the test stand wire position will be a "common mode signal" and when the measurements are combined differentially, it will be canceled.  The RF bias can be produced when the "split plates", which are offset axially from one another, receive a different signal from a signal on a centered wire.  In this case BPM is flipped vertically, on the test stand and re-measured  The relative positions of the plates are reversed and any difference, as before, will be common mode, and when combined differentially will be canceled out.  By combining all four measurements properly, the legitimate offset, the test stand bias, and the RF bias can be separated.  BPMs  numbered 1117, 1135, 3136, 3140, 3141, 3142, 3144, 3147, 3149, 3150, 3160, 3152, 3153, 3154, 3156, 3157, 3159, and 3160 were measured in all four orientations.   BPMs numbered 3134, 3138, 3139, 3148, 3151, 3155 and 3158 were measured only in the normal orientation.  , The summary averages are based on the larger group, which was measured in a qrid pattern, of 5 mm steps, –15 to +15 mm both axis.  Refer to the files "Numi BPM Test Stand Orientations" and "Measurement Summary for more details and reference.

|  | BPM TEST STAND DATA |  |  | Jim Fitzgerald |
| --- | --- | --- | --- | --- |
|  | Numi Large (MI8) BPMs |  |  | Sept. 2004 |  |
| Average of 17 BPMS | Ave. Sen. | Ave. Sen. |  | Ave. BPM | Ave. Table | Ave RF |
| measured in four | db/mm | (mm/db) |  | Offset | Bias | Bias |
| orientations | 0.245 | 4.088 | Ave dB = | -0.005 | -0.001 | -0.036 |
|  |  |  |  | Ave mm = | -0.021 | -0.005 | -0.145 |
| BPM No. |  | Sensitivity | Sensitivity | Raw Meas | BPM OS | Wire OS | RF Bias |
| N-MI8 3140 |  | (dB/mm) | (mm/dB) | dB | No chg | Chg mid 2 | Chg bot 2 |
| 3140GTSN |  | 0.245 | 4.082 | -0.067 | -0.067 | -0.067 | -0.067 |
| 3140GTRN |  | 0.245 | 4.082 | -0.047 | -0.047 | 0.047 | -0.047 |
| 3140GTSF |  | 0.245 | 4.082 | 0.013 | 0.013 | -0.013 | -0.013 |
| 3140GTRF |  | 0.245 | 4.082 | -0.016 | -0.016 | -0.016 | 0.016 |
|  | Ave 4 | 0.245 | 4.082 | Ave dB | -0.029 | -0.012 | -0.028 |
|  |  |  |  | Ave mm | -0.119 | -0.050 | -0.113 |
| 3141GTSN |  | 0.245 | 4.082 | -0.015 | -0.015 | -0.015 | -0.015 |
| 3141GTRN |  | 0.245 | 4.082 | -0.009 | -0.009 | 0.009 | -0.009 |
| 3141GTSF |  | 0.245 | 4.082 | 0.053 | 0.053 | -0.053 | -0.053 |
| 3141GTRF |  | 0.245 | 4.082 | 0.048 | 0.048 | 0.048 | -0.048 |
|  | Ave 4 | 0.245 | 4.082 | Ave dB | 0.019 | -0.003 | -0.031 |
|  |  |  |  | Ave mm | 0.079 | -0.011 | -0.128 |
| 3142GTSN |  | 0.245 | 4.082 | -0.041 | -0.041 | -0.041 | -0.041 |
| 3142GTRN |  | 0.246 | 4.065 | -0.040 | -0.040 | 0.040 | -0.040 |
| 3142GTSF |  | 0.245 | 4.082 | 0.024 | 0.024 | -0.024 | -0.024 |
| 3142GTRF |  | 0.245 | 4.082 | 0.028 | 0.028 | 0.028 | -0.028 |
|  | Ave 4 | 0.245 | 4.076 | Ave dB | -0.007 | 0.001 | -0.033 |
|  |  |  |  | Ave mm | -0.030 | 0.003 | -0.136 |
| 3144GTSN |  | 0.247 | 4.049 | -0.049 | -0.049 | -0.049 | -0.049 |
| 3144GTRN |  | 0.247 | 4.049 | -0.070 | -0.070 | 0.070 | -0.070 |
| 3144GTSF |  | 0.246 | 4.065 | 0.042 | 0.042 | -0.042 | -0.042 |
| 3144GTRF |  | 0.246 | 4.065 | 0.058 | 0.058 | 0.058 | -0.058 |
|  | Ave 4 | 0.247 | 4.057 | Ave dB | -0.005 | 0.009 | -0.055 |
|  |  |  |  | Ave mm | -0.019 | 0.038 | -0.222 |
| 3147GTSN |  | 0.245 | 4.082 | -0.057 | -0.057 | -0.057 | -0.057 |
| 3147GTRN |  | 0.246 | 4.065 | -0.046 | -0.046 | 0.046 | -0.046 |
| 3147GTSF |  | 0.245 | 4.082 | 0.027 | 0.027 | -0.027 | -0.027 |
| 3147GTRF |  | 0.246 | 4.065 | -0.002 | -0.002 | -0.002 | 0.002 |
|  | Ave 4 | 0.246 | 4.073 | Ave dB | -0.020 | -0.010 | -0.032 |
|  |  |  |  | Ave mm | -0.079 | -0.041 | -0.130 |
| 3149GTSN |  | 0.243 | 4.115 | -0.042 | -0.042 | -0.042 | -0.042 |
| 3149GTRN |  | 0.243 | 4.115 | -0.048 | -0.048 | 0.048 | -0.048 |
| 3149GTSF |  | 0.243 | 4.115 | 0.330 | 0.330 | -0.330 | -0.330 |
| 3149GTRF |  | 0.244 | 4.098 | 0.040 | 0.040 | 0.040 | -0.040 |
|  | Ave 4 | 0.243 | 4.111 | Ave dB | 0.070 | -0.071 | -0.115 |
|  |  |  |  | Ave mm | 0.288 | -0.292 | -0.473 |

| BPM No. |  | Sensitivity | Sensitivity | Offset | BPM OS | Wire OS | RF Bias |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | (dB/mm) | (mm/dB) | (dB) | No chg | Chg mid 2 | Chg bot 2 |
| 3150GTSN |  | 0.246 | 4.065 | -0.029 | -0.029 | -0.029 | -0.029 |
| 3150GTRN |  | 0.246 | 4.065 | -0.043 | -0.043 | 0.043 | -0.043 |
| 3150GTSF |  | 0.246 | 4.065 | 0.048 | 0.048 | -0.048 | -0.048 |
| 3150GTRF |  | 0.246 | 4.065 | 0.060 | 0.060 | 0.060 | -0.060 |
|  | Ave 4 | 0.246 | 4.065 | Ave dB | 0.009 | 0.007 | -0.045 |
|  |  |  |  | Ave mm | 0.037 | 0.026 | -0.183 |
| 3152GTSN |  | 0.245 | 4.082 | -0.047 | -0.047 | -0.047 | -0.047 |
| 3152GTRN |  | 0.245 | 4.082 | -0.290 | -0.290 | 0.290 | -0.290 |
| 3152GTSF |  | 0.245 | 4.082 | 0.087 | 0.087 | -0.087 | -0.087 |
| 3152GTRF |  | 0.245 | 4.082 | 0.071 | 0.071 | 0.071 | -0.071 |
|  | Ave 4 | 0.245 | 4.082 | Ave dB | -0.045 | 0.057 | -0.124 |
|  |  |  |  | Ave mm | -0.183 | 0.232 | -0.505 |
| 3153GTSN |  | 0.244 | 4.098 | -0.054 | -0.054 | -0.054 | -0.054 |
| 3153GTRN |  | 0.245 | 4.082 | -0.042 | -0.042 | 0.042 | -0.042 |
| 3153GTSF |  | 0.245 | 4.082 | 0.016 | 0.016 | -0.016 | -0.016 |
| 3153GTRF |  | 0.244 | 4.098 | 0.004 | 0.004 | 0.004 | -0.004 |
|  | Ave 4 | 0.245 | 4.087 | Ave dB | -0.019 | -0.006 | -0.029 |
|  |  |  |  | Ave mm | -0.078 | -0.025 | -0.119 |
| 3154GTSN |  | 0.245 | 4.082 | -0.044 | -0.044 | -0.044 | -0.044 |
| 3154GTRN |  | 0.245 | 4.082 | -0.036 | -0.036 | 0.036 | -0.036 |
| 3154GTSF |  | 0.245 | 4.082 | 0.035 | 0.035 | -0.035 | -0.035 |
| 3154GTRF |  | 0.245 | 4.082 | 0.030 | 0.030 | 0.030 | -0.030 |
|  | Ave 4 | 0.245 | 4.082 | Ave dB | -0.004 | -0.003 | -0.036 |
|  |  |  |  | Ave mm | -0.015 | -0.013 | -0.148 |
| 3157GTSN |  | 0.246 | 4.065 | -0.056 | -0.056 | -0.056 | -0.056 |
| 3157GTRN |  | 0.246 | 4.065 | -0.066 | -0.066 | 0.066 | -0.066 |
| 3157GTSF |  | 0.246 | 4.065 | 0.012 | 0.012 | -0.012 | -0.012 |
| 3157GTRF |  | 0.246 | 4.065 | 0.009 | 0.009 | 0.009 | -0.009 |
|  | Ave 4 | 0.246 | 4.065 | Ave dB | -0.025 | 0.002 | -0.036 |
|  |  |  |  | Ave mm | -0.103 | 0.007 | -0.145 |
| 3159GTSN |  | 0.244 | 4.098 | -0.024 | -0.024 | -0.024 | -0.024 |
| 3159GTRN |  | 0.244 | 4.098 | -0.035 | -0.035 | 0.035 | -0.035 |
| 3159GTSF |  | 0.244 | 4.098 | 0.043 | 0.043 | -0.043 | -0.043 |
| 3159GTRF |  | 0.244 | 4.098 | 0.051 | 0.051 | 0.051 | -0.051 |
|  | Ave 4 | 0.244 | 4.098 | Ave dB | 0.009 | 0.005 | -0.038 |
|  |  |  |  | Ave mm | 0.036 | 0.019 | -0.157 |
| 3160GTSN |  | 0.246 | 4.065 | -0.028 | -0.028 | -0.028 | -0.028 |
| 3160GTRN |  | 0.246 | 4.065 | -0.009 | -0.009 | 0.009 | -0.009 |
| 3160GTSF |  | 0.246 | 4.065 | 0.058 | 0.058 | -0.058 | -0.058 |
| 3160GTRF |  | 0.246 | 4.065 | 0.042 | 0.042 | 0.042 | -0.042 |
|  | Ave 4 | 0.246 | 4.065 | Ave dB | 0.016 | -0.009 | -0.034 |
|  |  |  |  | Ave mm | 0.064 | -0.036 | -0.139 |

| BPM No. |  | Sensitivity | Sensitivity | Offset | BPM OS | Wire OS | RF Bias |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | (dB/mm) | (mm/dB) | (dB) | No chg | Chg mid 2 | Chg bot 2 |
| 1117GTSN |  | 0.244 | 4.098 | 0.033 | 0.033 | 0.033 | 0.033 |
| 1117GTRN |  | 0.244 | 4.098 | 0.024 | 0.024 | -0.024 | 0.024 |
| 1117GTSF |  | 0.244 | 4.098 | -0.045 | -0.045 | 0.045 | 0.045 |
| 1117GTRF |  | 0.244 | 4.098 | -0.036 | -0.036 | -0.036 | 0.036 |
|  | Ave 4 | 0.244 | 4.098 | Ave dB | -0.006 | 0.005 | 0.035 |
|  |  |  |  | Ave mm | -0.025 | 0.018 | 0.141 |
|  |  | (dB/mm) | (mm/dB) | (dB) | No chg | Chg mid 2 | Chg bot 2 |
| 3136GTSN |  | 0.244 | 4.098 | -0.081 | -0.081 | -0.081 | -0.081 |
| 3136GTRN |  | 0.244 | 4.098 | -0.105 | -0.105 | 0.105 | -0.105 |
| 3136GTSF |  | 0.244 | 4.098 | -0.004 | -0.004 | 0.004 | 0.004 |
| 3136GTRF |  | 0.244 | 4.098 | -0.007 | -0.007 | -0.007 | 0.007 |
|  | Ave 4 | 0.244 | 4.098 | Ave dB | -0.049 | 0.005 | -0.044 |
|  |  |  |  | Ave mm | -0.202 | 0.022 | -0.179 |
| 3156GTSN |  | 0.244 | 4.098 | -0.025 | -0.025 | -0.025 | -0.025 |
| 3156GTRN |  | 0.245 | 4.082 | -0.038 | -0.038 | 0.038 | -0.038 |
| 3156GTSF |  | 0.245 | 4.082 | 0.036 | 0.036 | -0.036 | -0.036 |
| 3156GTRF |  | 0.245 | 4.082 | 0.019 | 0.019 | 0.019 | -0.019 |
|  | Ave 4 | 0.245 | 4.086 | Ave dB | -0.002 | -0.001 | -0.030 |
|  |  |  |  | Ave mm | -0.008 | -0.004 | -0.121 |
| 3135GTSN |  | 0.244 | 4.098 | 0.033 | 0.033 | 0.033 | 0.033 |
| 3135GTRN |  | 0.244 | 4.098 | 0.024 | 0.024 | -0.024 | 0.024 |
| 3135GTSF |  | 0.244 | 4.098 | -0.045 | -0.045 | 0.045 | 0.045 |
| 3135GTRF |  | 0.244 | 4.098 | -0.036 | -0.036 | -0.036 | 0.036 |
|  | Ave 4 | 0.244 | 4.098 | Ave dB | -0.006 | 0.005 | 0.035 |
|  |  |  |  | Ave mm | -0.025 | 0.018 | 0.141 |
| 1135GTSN |  | 0.238 | 4.202 | -0.951 | -0.951 | -0.951 | -0.951 |
| 1135GTRN |  | 0.240 | 4.167 | -0.278 | -0.278 | 0.278 | -0.278 |
| 1135GTSF |  | 0.239 | 4.184 | -0.856 | -0.856 | 0.856 | 0.856 |
| 1135GTRF |  | 0.240 | 4.167 | -0.826 | -0.826 | -0.826 | 0.826 |
| Rejected | Ave 4 | 0.240 | 4.172 | Ave dB | -0.728 | -0.161 | 0.113 |
|  |  |  |  | Ave mm | -3.037 | -0.671 | 0.473 |
| The BPMs listed below were measured only in the normal orientation |  |  |
| BPM No. |  | Sensitivity | Sensitivity | BPM Offset | Rot |  |  |
|  |  | (dB/mm) | (mm/dB) | (dB) | Deg. |  |  |
| 3148 |  | 0.247 | 4.049 | -0.016 | -0.180 |  |  |
| 3138 |  | 0.245 | 4.082 | -0.011 | -0.420 |  |  |
| 3139 |  | 0.247 | 4.049 | 0.014 | -0.034 |  |  |
| 3158 |  | 0.244 | 4.098 | 0.012 | 0.900 |  |  |
| 3155 |  | 0.248 | 4.032 | 0.012 | -0.920 |  |  |
| 3151 |  | 0.244 | 4.098 | 0.001 | -0.500 |  |  |
| 3134 |  | 0.244 | 4.098 | 0.044 | -0.500 |  |  |

| BPM TEST STAND DATA |  |  |  | Jim Fitzgerald |
| --- | --- | --- | --- | --- |
| Numi Target BPMs |  |  |  | Sept. 2004 |
| Ave.of 5 Target BPMs |  | Ave. Sen. | Ave. Sen. |  | Ave. OS | Ave. OS | Ave. OS |
| measured in four | db/mm | (mm/db) |  | (dB) | (mm) | (mm) |
| orientations |  |  |  |  |  |  |
|  |  | 0.558 | 1.792 | Ave dB | 0.000 | -0.007 | 0.034 |
|  |  |  |  | Ave mm | 0.000 | -0.013 | 0.061 |
| BPM No. |  | Sensitivity | Sensitivity | Raw Meas | BPM OS | Wire OS | RF Bias |
|  |  | (dB/mm) | (mm/dB) | dB | No chg | Chg mid 2 | Chg bot 2 |
| 111GTSN |  | 0.556 | 1.799 | 0.045 | 0.045 | 0.045 | 0.045 |
| 111GTRN |  | 0.556 | 1.799 | 0.048 | 0.048 | -0.048 | 0.048 |
| 111GTSF |  | 0.556 | 1.799 | -0.011 | -0.011 | 0.011 | 0.011 |
| 111GTRF |  | 0.555 | 1.802 | -0.013 | -0.013 | -0.013 | 0.013 |
|  | Ave 4 | 0.556 | 1.799 | Ave dB | 0.017 | -0.001 | 0.029 |
|  |  |  |  | Ave mm | 0.031 | -0.002 | 0.053 |
| 112GTSN |  | 0.557 | 1.792 | -0.005 | -0.005 | -0.005 | -0.005 |
| 112GTRN |  | 0.558 | 1.799 | -0.001 | -0.001 | 0.001 | -0.001 |
| 112GTSF |  | 0.556 | 1.795 | -0.050 | -0.050 | 0.050 | 0.050 |
| 112GTRF |  | 0.557 | 1.795 | -0.081 | -0.081 | -0.081 | 0.081 |
|  | Ave 4 | 0.557 | 1.795 | Ave dB | -0.034 | -0.009 | 0.031 |
|  |  |  |  | Ave mm | -0.061 | -0.016 | 0.056 |
| 113GTSN |  | 0.557 | 1.795 | 0.069 | 0.069 | 0.069 | 0.069 |
| 113GTRN |  | 0.557 | 1.795 | 0.085 | 0.085 | -0.085 | 0.085 |
| 113GTSF |  | 0.557 | 1.795 | 0.000 | 0.000 | 0.000 | 0.000 |
| 113GTRF |  | 0.556 | 1.799 | -0.013 | -0.013 | -0.013 | 0.013 |
|  | Ave 4 | 0.557 | 1.796 | Ave dB | 0.035 | -0.007 | 0.042 |
|  |  |  |  | Ave mm | 0.063 | -0.013 | 0.075 |
| 114GTSN |  | 0.564 | 1.773 | 0.027 | 0.027 | 0.027 | 0.027 |
| 114GTRN |  | 0.564 | 1.773 | 0.066 | 0.066 | -0.066 | 0.066 |
| 114GTSF |  | 0.564 | 1.773 | -0.022 | -0.022 | 0.022 | 0.022 |
| 114GTRF |  | 0.563 | 1.776 | -0.054 | -0.054 | -0.054 | 0.054 |
|  | Ave 4 | 0.564 | 1.774 | Ave dB | 0.004 | -0.018 | 0.042 |
|  |  |  |  | Ave mm | 0.008 | -0.031 | 0.075 |
| 115GTSN |  | 0.554 | 1.805 | 0.004 | 0.004 | 0.004 | 0.004 |
| 115GTRN |  | 0.554 | 1.805 | 0.000 | 0.000 | 0.000 | 0.000 |
| 115GTSF |  | 0.555 | 1.802 | -0.044 | -0.044 | 0.044 | 0.044 |
| 115GTRF |  | 0.563 | 1.776 | -0.054 | -0.054 | -0.054 | 0.054 |
|  | Ave 4 | 0.557 | 1.797 | Ave dB | -0.024 | -0.002 | 0.026 |
|  |  |  |  | Ave mm | -0.042 | -0.003 | 0.046 |
| BPM No. |  | Sensitivity | Sensitivity | Raw Meas | BPM OS | Wire OS | RF Bias |
|  |  | (dB/mm) | (mm/dB) |  |  |  |  |