Beams-doc-1786

October, 2005

Requirements for the Main Injector BPM upgrade

D. Capista, A. Marchionni, A. Yousif

October 24, 2005

1	Introduction	4

1.1	The present system	5

1.2	Scope of the Project	8

2	Beam time structures, beam modes and dynamic ranges	8

3	Measurement Specifications	10

4	Data Acquisition	13

4.1	Measurement Modes	13

4.2	Measurement Types	13

4.3	Acquisition specifications	13

4.4	Specific types of Turn by Turn Measurements	14

4.4.1	Injection and Extraction Measurements	14

4.4.2	User defined Measurement	14

4.4.3	Safe Mode	15

4.5	Measurement storage	15

4.6	Error Handling	15

5	A suggested implementation	16

5.1	Slip Stacking and NuMI	20

5.2	Pbar Acceleration	20

6	End User Applications	20

6.1	Current System	20

o	I37 BPM Control Parameters	21

o	I38 BPM/BLM Hardware Tests	21

o	I39 BPM/ BLM Plots /List	21

o	I42 Tune From TBT BPM	21

o	I50 MI Orbit Smoothing / C453	21

o	I52 Booster to MI Injection Closure	21

o	I73 MI Measure Chromaticity	21

o	I75 MI Quad Move	21

o	I92 TBT Data Analysis	21

o	T117 Tev Orbit Closure	21

6.2	New applications	21

7	Maintenance and Calibration	22

8	References	22

9	Appendix A	23

10	Appendix B	24

10.1	Reset $29	24

10.2	Reset $21	24

10.3	Reset $23	25

10.4	Reset $2A	26

10.5	Reset $2B	27

10.6	Reset $2D	28

10.7	Resets $20, 2E	29

11	GLOSSARY	30

TABLES

Table 1: Main Injector current Operational Modes	6

Table 2: Protons for anti-proton stacking and fixed target	9

Table 3: Proton transfers to the Tevatron	9

Table 4: Proton transfers for collider operation	10

Table 5: Anti-proton transfers from Accumulator/Recycler Ring	10

Table 6: Anti-proton transfers to the Tevatron	10

Table 7: Measurement specifications for 53 MHz bunch structure, narrow bandwidth mode. Values in parenthesis are for high bandwidth mode, with single turn resolution.	12

Table 8: Measurement specifications for 2.5 MHz bunch structure, narrow bandwidth mode. Values in parenthesis are for high bandwidth mode, with single turn resolution.	12

Table 9: State message definition	19

Table 10: Data acquisition and defition examples.	20

FIGURES

Figure 1: Cross section of Main Injector BPM inside the quadrupole beam pipe.	7

Figure 2: Large aperture BPM detectors compatible with the aperture of the new quadrupoles. The electrodes, 4” long, are positioned in the horizontal and vertical planes	7

Figure 3: data acquisition and memory diagram	16

Figure 4: Data acquisition TBT detail	18

# Introduction

The Fermilab Main Injector (MI) is a rapid cycling accelerator with an injection momentum of 8.9 GeV/c and an extraction momentum of up to 150 GeV/c [1]. It serves multiple purposes, providing protons for anti-proton production, fixed target experiments and Tevatron injections. It receives anti-protons from the Accumulator and accelerates them for Tevatron collider operation. Also it serves as interface between the Recycler Ring and the rest of the accelerator complex. Antiproton transfers in and out of the Recycler Ring take place through two transfer lines connecting the Recycler to the Main Injector.

The Main Injector is seven times in circumference (3319.42 meter) of the Booster (474.2m) from which it accepts the protons, and is slightly larger than half the circumference of the Tevatron (2000( meter) to which it transfers both the protons and the anti-protons. Only six Booster cycles (batches) are utilized to fill the MI, allowing rest of the space for the abort gap. The harmonic numbers of the Booster, MI and the Tevatron are 84, 588 (84(7), and 1113 (84(13.25) respectively. Transfers from the Booster and to the Tevatron are bucket to bucket transfers at a frequency close to 53 MHz. RF frequency sweeps from 52.8 MHz at injection to 53.1 MHz at extraction. Anti-protons transfers from the Accumulator and transfers to or from the Recycler Ring instead happen in 2.5 MHz buckets.

The Main Injector accelerates protons from 8.9 GeV/c to 120 GeV/c for anti-proton production, for fast single turn extraction to dedicated neutrino experiments (i.e. NuMI/MINOS) and for slow resonant extraction to fixed target experimental areas. For collider operations, it accelerates protons from the Booster from 8.9 GeV/c to 150 GeV/c, and anti-protons from the Accumulator or the Recycler Ring from 8.9 GeV/c to 150 GeV/c, injecting them into the Tevatron. Moreover it is used to transfer anti-protons from the Accumulator to the Recycler Ring at 8.9 GeV/c.

Even if the Main Injector operation includes acceleration of both protons and anti-protons, they are never present simultaneously in the ring. As MI is viewed from above, protons and anti-protons circulate in counterclockwise and clockwise directions, respectively.

Each of the different modes of operation of the Main Injector described above, together with possible machine studies, are identified by Tevatron Clock events (TCLK), which act as machine cycle resets (start of cycle) [2][3]. A given MI cycle reset, which essentially defines the magnet ramp, can be associated to multiple Main Injector states. A MI state, through the Low Level RF system and the Timeline Generator (see later), specifies the number of Booster batches injected in the machine and their RF structure. MI states are broadcasted in the MDAT system (I:MDAT56).

A summary of MI TCLK events is shown in Appendix A. Details of the MI cycle resets are given in Appendix B.

The different machine cycles are controlled and synchronized by the Timeline Generator system (TLG). For example, in a 60 s super-cycle, the TLG specifies the sequence of the Main Injector cycle resets, the associated Main Injector states and the duration of the machine cycles. In case of proton transfers from the Booster, it also specifies the number of Booster injections (batches), and, through the associated Booster TCLK event, the number of bunches/batch and their intensity. Multiple Booster injections can happen in a same MI cycle at a maximum rate of 15 Hz.

Single turn extractions of the entire MI beam or of a portion of it are accomplished by the use of fast extraction magnets (kickers). Extraction times are set by the extraction kicker timings as defined by Beam Sync clocks plus a delay. The flattop portion of the kicker magnets determines which portion of the beam is extracted from the machine. Multiple extractions in a same cycle are possible (i.e. when running anti-proton production and beam to neutrino experiments in a same cycle) and extraction times are presently as close as 1 ms.

Numerous different and relatively short machine cycles can be present in a super-cycle at any given time, as defined by the TLG.

In order to support all these different operational modes, multiple beam injection and extraction points are present in the Main Injector.

Table 1 summarizes the Main Injector current operational modes, the corresponding TCLK cycle reset events and the associated MI states, the RF frequencies involved, the approximate cycle lengths and energy sweeps, and the injection/extraction points.

It is worth noting that cycle lengths are relatively short, due to the fast acceleration rate of the Main Injector. Beam injection into the Main Injector can happen as fast as ~ 50 ms after a TCLK reset event.

The need for an upgrade of the MI BPM system was already apparent since a few years ago, but scheduling reasons suggested the project to be delayed after the completion of the Recycler and Tevatron BPM projects. A first requirement document for the MI BPM upgrade was written in 2003 ([4]). The present requirement document benefits from that first document and the experience gained on the Recycler and Tevatron BPM upgrade projects.

## The present system

The Main Injector has four BPM detectors per betatron wavelength in both the horizontal and the vertical planes. There are a total of 208 BPMs for the 3320 m ring. Out of these, 203 are MI style ring BPMs, and are located in the downstream end of every MI quadrupole. The other 5 are wide aperture BPMs located at Q101, Q402, Q522, Q608 and Q620 locations.

The BPM detectors in MI consist of four transmission line strips, or strip-lines, located on the perimeter of the beam pipe. The strip-lines are 250 mm long and 12.7 mm wide. They are positioned at a distance of 46.7 mm horizontally and 44.5 mm vertically, center to center (see

Figure 1). The pick-up has a characteristic impedance of 50(, determined by the gap between the strip and the beam pipe. The current BPM RF module input impedance is matched to 50( within a 5 MHz bandwidth centered at 53MHz. The outputs are combined in pairs externally to form either a horizontal or a vertical detector. Each strip-line is shorted at one end and connected to a ceramic feed-through at the other end, which makes these BPMs non-directional.  At present any BPM measures either the horizontal or the vertical position per cycle at each quadrupole. It can be remotely switched to the orthogonal mode.

The wide aperture BPMs, with 6” long plates and a 4.625” aperture, are located adjacent to the Lambertson magnets and are mounted external to the quadrupole magnets.

The Abort line of the Main Injector is instrumented with 3 large aperture BPMs, each one providing both a horizontal and vertical position measurement.

| Operational Modes | TCLK | Machine State | RF frequency (MHZ) | Cycle Length (s) | Energy sweep (GeV) | Injection region | Extraction region |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Anti-proton production | $29 | 3 | 53 | 1.6 – 2.0 | 8.9 - 120 | Q101 | Q522 |
| Neutrino experiments | $23 | 21 | 53 | 2.0 | 8.9 - 120 | Q101 | Q608 |
| Anti-proton production and Neutrino experiments | $23 | 5 | 53 | 2.0 | 8.9 - 120 | Q101 | Q522, Q608 |
| Slow extraction to fixed-target experiment | $21 | 28 | 53 | 2.6 – 6.0 | 8.9 - 120 | Q101 | Q522 |
| Slow extraction to fixed-target experiment and anti-proton production | $21 | 14 | 53 | 2.6 – 6.0 | 8.9 - 120 | Q101 | Q522, Q522 |
| Proton transfers to the Tevatron | $2B | 4 | 53 | 3.65 | 8.9 - 150 | Q101 | Q522 |
| Anti-proton transfers from Accumulator to Recycler | $2D | 20 | 2.5 | 4.9 | 8.9 | Q522 | Q222 |
| Anti-proton transfers from Accumulator to Tevatron | $2A, $2E( | 23 | 2.5 and 53 | 7.8 – 12 | 8.9 - 150 | Q522 | Q620 |
| Anti-proton transfers from Recycler to Tevatron | $2A, $2E | 11 | 2.5 and 53 | 7.8 – 12 | 8.9 - 150 | Q321 | Q620 |

Table 1: Main Injector current Operational Modes

Figure 1: Cross section of Main Injector BPM inside the quadrupole beam pipe.

Figure 2: Large aperture BPM detectors compatible with the aperture of the new quadrupoles. The electrodes, 4” long, are positioned in the horizontal and vertical planes

## Scope of the Project

The scope of the upgrade includes all the BPM detectors installed in the Main Injector.

The readout of the MI Abort line BPMs will be upgraded utilizing the system recently developed within the transfer line BPM upgrade project and it is outside the scope of this project.

The present switching capability between horizontal and vertical position measurements of the detectors in the MI ring will not be retained in the upgraded system. Horizontal only measurements are required at focusing quadrupole locations and vertical only measurements at defocusing quadrupole locations.

The upgrade will make use of the existing pick-ups and cables. This is true for all locations except at the Q101, Q222, Q321, Q402, Q522, Q608 and Q620 locations in the Main Injector ring, where large aperture quadrupole magnets, presently under construction, will be installed. In these specific locations the present BPM detectors will be replaced by larger aperture ones compatible with the aperture of the new quadrupole magnets (see Figure 2). The new BPM detectors in these locations shall be capable of simultaneous horizontal and vertical position measurements.

The project includes electronics and software necessary to measure beam positions and relative beam intensity utilizing the BPM detectors. Also included are relevant documentation and user guides to describe, use and maintain the system.

The project includes diagnostics to verify the performance of the electronics and to maintain the system within the specification requirements.

# Beam time structures, beam modes and dynamic ranges

At any particular instance of time there will be only one kind of beam in the MI. It will be either protons or anti-protons. The beam energy can vary between 8.9 GeV to 150 GeV. The time structure of the beam can be:

- 53 MHz

  - Protons. From 1 bunch up to a full batch of 84 bunches in successive 53 MHz buckets (19ns apart). The 95% bunch length varies between 1 ns, at transition and at 120 GeV after bunch rotation for anti-proton stacking, to 10 ns for single coalesced bunches. Up to 6 batches, each one of 84 bunches, can be loaded in the MI for anti-proton production and in support of NuMI, presently with two main operational modes:

    - Mixed-mode, where one batch (which is the result of slip-stacking two different batches) is extracted to the anti-proton target and 5 batches are extracted to NuMI. The first bunch of the anti-proton batch is spaced by 122 RF buckets from the first bunch of the adjacent NuMI batches, in order to allow for the rise and fall times of the MI52 extraction kicker. NuMI batches are spaced so that first bunches of adjacent batches are 86 RF buckets apart.

    - NuMI only, where 6 batches are extracted to NuMI. First bunches of adjacent batches are spaced by 86 RF buckets.

  - Anti-protons. Anti-protons are currently accelerated using 53 MHz RF cavities from 8.9 GeV to 150 GeV. With the implementation of the 2.5 MHz acceleration scheme up to 27 GeV, anti-protons would be accelerated in 53 MHz bunches from 27 GeV to 150 GeV. Four consecutive groups of antiprotons, spaced by 396 ns, of typically 5 (up to 9) 53 MHz bunches, are present in MI.

- 2.5 MHz - Protons or anti-protons.

  - Anti-protons transfers from the Accumulator and transfers to or from the Recycler Ring instead happen in four 2.5 MHz bunches in successive 2.5 MHz RF buckets (396 ns spacing). 95% bunch lengths range from 100 to 200 ns at 8GeV. With the implementation of 2.5 MHz anti-proton acceleration, bunch lengths will shrink down to 40 ns at 27 GeV.

For each beam RF structure of the Main Injector operational modes, we will specify bunch arrangements, lengths and intensities.

- Protons for anti-proton stacking and fixed target ($29, $23, $21)

One up to 6 Booster batches, however spaced, are injected in MI at 8.9 GeV and accelerated up to 120 GeV. Beam is extracted in a single turn to the anti-proton production target and to dedicated neutrino experiments or in a slow resonant extraction mode (slow spill) to fixed target experimental areas. Mixed-mode cycles are present in MI where beam is successively extracted to the anti-proton source and to the neutrino experiments.

| Energy | 8 to 120 GeV |
| --- | --- |
| Bunch RF | 53 MHz |
| Bunch structure | 84 bunches/batch, up to 6 batches however spaced |
| Bunch length (at 95%) | 6 ns (8 GeV) 1.5 ns (150 GeV, after bunch rotation) |
| Intensity/bunch | 5E9 to 200E9 |

Table 2: Protons for anti-proton stacking and fixed target

- Proton transfers to the Tevatron ($2B)

  - Tune-up (MI state 12)

| Energy | 8 to 150 GeV |
| --- | --- |
| Bunch RF | 53 MHz |
| Bunch structure | 30 to 84 uncoalesced bunches |
| Bunch length (at 95%) | 6 ns (8 GeV) 2.5 ns (150 GeV) |
| Intensity/bunch | 10E9 to 100E9 |

Table 3: Proton transfers to the Tevatron

  - Proton transfers for collider operation (MI state 4)

| Energy | 8 to 150 GeV |
| --- | --- |
| Bunch RF | 53 MHz |
| Bunch structure | 5-9 bunches, coalesced into one bunch at 150 GeV |
| Bunch length (at 95%) | 6 ns (8 GeV) 2.5 ns (150 GeV uncoalesced) 10ns (150 GeV coalesced) |
| Intensity/bunch | 30E9 to 100E9 (uncoalesced) 50E9 to 400E9 (coalesced) |

Table 4: Proton transfers for collider operation

- Anti-protons transfers from Accumulator/Recycler Ring ($2A, $2E, $2D)

| Energy | 8 to 27 GeV |
| --- | --- |
| Bunch RF | 2.5 MHz |
| Bunch structure | 4 consecutive bunches |
| Bunch length (at 95%) | 100-200 ns (8GeV) 40 ns (27 GeV) |
| Intensity/bunch | 5E9 to 150E9 |

Table 5: Anti-proton transfers from Accumulator/Recycler Ring

- Anti-proton transfers to the Tevatron ($2A, $2E)

| Energy | 8 to 150 GeV |
| --- | --- |
| Bunch RF | 53 MHz |
| Bunch structure | 4 consecutive groups, spaced by 400 ns, of typically 5 (up to 9) 53 MHz bunches, coalesced into a single bunch at 150 GeV |
| Bunch length (at 95%) | 8.5 ns (8 GeV) 3.5 ns (150 GeV uncoalesced) 10 ns (150 GeV coalesced) |
| Intensity/bunch | 5E9 (uncoalesced) 10E9 to 150E9 (coalesced) |

Table 6: Anti-proton transfers to the Tevatron

# Measurement Specifications

The upgraded BPM system will provide a position measurement of the beam present in the machine for each one of the operating modes described above. In particular the system should be capable of measuring beam positions with 6 batches loaded in the MI, however spaced.

The upgraded BPM system will also provide a beam intensity measurement from the magnitude of the sum of the plate signals for each BPM. This measurement will provide a “relative” intensity measurement between different BPMs. BPM-to-BPM scaling capability should be incorporated in the system to equalize the response of the different BPMs.

Beam in the Main Injector will either be bunched in 53 MHz or 2.5 MHz. The BPM system is expected to make measurements with either bunch structure. The system will be able to be switched from one bunch structure to the other structure in a given cycle with a minimum interruption in data acquisition (less than ~ 10 ms).

It is understood that during particular RF manipulations of the beam, like slip-stacking and coalescing, with no steady defined RF structure, the BPM system might not report sensible data.

The system will operate in both narrow and high bandwidth mode.

Narrow bandwidth will be used for measurements that do not need single turn resolution and represent a single ensemble position measurement of the beam present in the machine averaged over many turns (100 turns or more).

High bandwidth is needed for measurements that require single turn resolution.

For 2.5 MHz anti-proton transfers the system will provide a single ensemble position measurement of the beam present in the machine. No bunch-by-bunch position measurement is required in this case.

For 53 MHz multi-batch operation the system should be able to provide, if requested, a batch-by-batch measurement. The batch-by-batch measurements are only possible in high bandwidth mode and will measure one batch at a time.

The system must be able to provide first injection turn and last extraction turn for each portion of beam injected or extracted from the machine. It has to be able to cope with multiple beam injection/extraction regions and with opposite circulating directions of protons and anti-protons in the machine.

The system should also be able to provide, when requested, beam orbits on multiple successive turns.

The system will need to be switched between narrow and high bandwidth operations in a given cycle with a minimum interruption in data acquisition (less than ~ 10 ms).

We will now proceed to define the quantities that will serve as a base for the BPM measurement specifications.

- Measurement range. This is defined as the range of positions, relative to the BPM center, over which the BPM measurement must be valid and meet the accuracy requirements.

- Absolute position accuracy. This determines how well the position of the beam is measured with respect to the survey center-line. This is limited by a number of factors including alignment, calibration, cable attenuation and noise.

- Relative position accuracy. This determines how well the displacement of the beam is measured over the measurement range. This requirement does not include offset errors, but does set limits on the scale errors, non-linearities and random errors.

- Position resolution. This is a requirement on the smallest change in beam position that the BPM can reliably measure.

- Position linearity. This is a requirement on the linearity of the BPM response to orbit changes over the measurement range. The linearity is defined as the difference between the measured BPM position and the slope of the BPM response at the center of the BPM.

- Long-term position stability. This is a requirement on the BPM system ability to give the same position value for the same beam position and intensity over time.

- Relative intensity accuracy. This is a requirement on the relative intensity measurements between BPMs. It is understood that during the acceleration cycle there will be, for a given BPM, apparent variations of beam intensity due to changes in bunch widths.

Table 7 and Table 8 summarize measurement specifications for 53 MHz and 2.5 MHz bunch structures, respectively.

| Measurement specifications | 53 MHz bunch structure |
| --- | --- |
| Measurement range | (25 mm, (30 mm for large aperture BPMs |
| Absolute position accuracy | 1 mm + 5% of actual beam position |
| Relative position accuracy (3 () | 5% of actual beam position |
| Position resolution (3 () | 0.05 mm (0.1 mm) |
| Position linearity | within 1.5% of the slope of the BPM response at the center of the BPM over (15 mm( |
| Long-term position stability | 0.2 mm |
| Relative intensity accuracy | 10% (20%) |

Table 7: Measurement specifications for 53 MHz bunch structure, narrow bandwidth mode. Values in parenthesis are for high bandwidth mode, with single turn resolution.

| Measurement specifications | Anti-protons of nominal intensity (>20E9/bunch), 2.5 MHz bunch structure | Anti-protons at low intensity (<20E9/bunch), 2.5 MHz bunch structure |
| --- | --- | --- |
| Measurement range | (25 mm, (30 mm for large aperture BPMs | (25 mm, (30 mm for large aperture BPMs |
| Absolute position accuracy | 1 mm + 10% of actual beam position | 1 mm + 20% of actual beam position |
| Relative position accuracy (3 () | 10% of actual beam position | 20% of actual beam position |
| Position resolution (3 () | 0.3 mm (0.5 mm) | 0.5 mm |
| Long-term position stability | 0.5 mm | N/A |
| Relative intensity accuracy | 20% (30%) | 30% |

Table 8: Measurement specifications for 2.5 MHz bunch structure, narrow bandwidth mode. Values in parenthesis are for high bandwidth mode, with single turn resolution.

# Data Acquisition

The new BPM must utilize the infrastructure of the Accelerator control system to automate data acquisition as much as reasonably possible. This includes use of

- the accelerator clock system,

- defined ACNET devices to transfer data and control information in/out of the front ends

- accelerator MDAT frames.

## Measurement Modes

- High bandwidth: see section 3.

- Narrow bandwidth: see section 3.

## Measurement Types

There are several different data types that the BPM system is required to acquire. These types are:

- Flash Frame: A single turn orbit measurement, performed in high bandwidth mode, associated with injection or extraction of beam.

- Turn-by-turn: A measurement of the orbit on every turn (588 53 MHz RF buckets) for a specified number of turns, performed in high bandwidth mode.

- Averaged orbit: An average of a specified number of turn-by-turn measurements.

- Display Frame: A narrow bandwidth measurement triggered by the display frame TCLK ($7B). This measurement occurs once per cycle.

- Profile frame: A narrow bandwidth measurement triggered by the profile frame TCLK ($7A). This measurement can occur up to 128 times per cycle.

- Fast time plot (FTP): a narrow bandwidth measurement taken at least at 500 Hz. FTP measurements for all channels in parallel must be supported.

## Acquisition specifications

Each Main Injector cycle is uniquely identified by a machine state. This state information is transmitted on the MDAT system and is valid at the TCLK reset for the main ramp. Each state must have a unique set of readout specifications that is re-configurable by the end-user.

The acquisition specification for a given cycle must specify the particle type (protons or anti-protons) and a series of measurements setups during that cycle. Each setup needs to specify an offset from the start of the cycle (in ms) for the change to take affect. The valid settings are:

- the analog filter (53 MHz or 2.5 MHz),

- the bandwidth mode (narrow or high)

- expected beam intensity (given the large dynamic range required the system will probably use programmable attenuators or amplifiers).

When in high bandwidth mode, the measurement setup must further specify

- if in flash or turn-by-turn mode

- a time offset from the start of the cycle  to arm the system

- Trigger event  (ie, the associated Beam Sync event)

- for the 53 MHz mode, should include the possibility to choose a user selectable portion of the beam.

It is understood that this system will require periods of time when data acquisition is not possible. There are currently three identified periods:

- at the start of cycle (50 ms);

- when the configuration mode is changing (max of 10 ms);

- when the data is being collected at the end of beam (TCLK event $26) on a cycle (max of 500ms).

- Efficient use of memory and programming is encouraged to keep these times to a minimum.

Once a data collection is started at the end of beam on a cycle, the BPM system will not start a new data acquisition or service BPM reads until all the data has been moved to the appropriate buffers accessible by the host level applications.

## Specific types of Turn by Turn Measurements

### Injection and Extraction Measurements

Beam position data for injections and extractions for the Main Injector are key to the successful monitoring of the machine. Injections and extractions for the Main Injector are identified by Beam Sync clock events. The BPM system must be able to collect data based on these events as defined in the state acquisition specification.

The BPM system will be directed (via the acquisition specifications) to take turn-by-turn measurements for particle injection and extraction, triggered by the relevant Beam Sync Clock event. The system must support 512 turns for each measurement. The specific turn with actual injection or extraction must occur in the corresponding dataset and the BPM system must identify the data for the first or last turn in the set.

If there are multiple injections in the same cycle, the system must provide the capability to measure the first turn for each portion of beam injected in the machine. Successive injections from the Booster happen at a maximum rate of 15 Hz.

If there are multiple single turn extractions in the same cycle, the system must provide the capability to measure the extraction turn of at least one portion of beam extracted from the machine.

The BPM system must support up to 20 injection/extraction turn-by-turn data sets for each machine state.

The system must generate an Averaged Orbit by averaging the injection turn-by-turn data for the first injection into the Main Injector in each cycle. The Averaged Orbit will be the average of the first 16 turns with beam. This data will be used for machine injection closure.

### User defined Measurement

The BPMs must provide the capability to select, for a specific machine state, a turn-by-turn measurement at a user specified time with 2048 data points. When a turn-by-turn measurement is requested for a particular machine state, no additional measurements are required for that state. All other requests (Flash, Display, Profile frames) should be ignored.

### Safe Mode

When in high bandwidth mode, it would be desirable, for debugging purposes, if the system would be able to respond to a self-trigger, based on beam over threshold, instead of using hard timing controls.

## Measurement storage

The BPM system must make the following data available for application access for each machine states:

- Up to 20 injection/extraction data sets

- 1 turn by turn buffer with a maximum of 2048 points

- Display frame

- Profile frames

- Average Orbit

The system must support 50 unique machine states. Data from a particular cycle, identified by a machine state, should be collected at the end of beam on that cycle and should remain accessible by host level applications until the end of beam event of the following cycle of that type (i.e. characterized by the same machine state).

In addition to these state buffers, the system will make available the last acquisition data and this data will reside in state zero(0). The “last acquisition” data buffer will make this system compatible with the current BPM End User applications to aid in installation.

Data must be consistent:

- data must be from the last occurred state;

- host level applications must be able to collect display and profile frames, injection and extraction data, and user defined turn-by-turn measurements from ALL BPMs (ie, across houses) for the same event. In particular, injection, extraction and turn-by-turn data must refer to the same turn or the same specified set of turns for all BPMs.

## Error Handling

A digital ACNET device with at least 16 bits has to be defined for each machine state, each bit corresponding to a possible error condition. Detailed error conditions have to be defined at the implementation stage of the system. When an error condition occurs during a state data acquisition, the system should set the corresponding bit in the digital ACNET word associated with that state. If the state executes without an error, the system should reset the corresponding digital ACNET device.

The possibility of alarming on at least a subset of the bits is subsequently left to the user.

# A suggested implementation

We suggest here an implementation of the upgraded BPM system that is compatible with the timing signals of current BPM system. This feature will aid in the installation process during periods of beam operation and help simplify modifications to the applications software.

The new BPM system will be able to respond to the defined TCLK events of the current BPM system. This includes Prepare for Beam, Display Frame and Profile Frame. The Set Bunch Mode events of the current BPM system are not needed for the new system.

Figure 3 shows a block diagram of the suggested implementation of the data acquisition system. This system and the following discussion represent the current understanding of hardware choices, desired measurements, automation, Accelerator control system compatibility, and operational integration.

Figure 3: data acquisition and memory diagram

The new BPM system will need to be configured in the proper mode before data acquisition can occur. The user will need to select 2.5Mhz or 53Mhz bunch structure, select the particle type (protons or anti-protons), and select between high bandwidth (turn-by-turn) and narrow bandwidth (closed orbit) mode.

The choice of hardware for this BPM system means that the system can only acquire data in a single mode at a time.

There are three distinct parts to this diagram: the closed orbit data, the turn-by-turn data, and the state machine that controls the mode of the system.

Selecting the Closed Orbit mode allows the system to acquire narrow bandwidth data. In this mode, data is saved into the profile frame buffer (128 frames) or the display frame buffer (one frame) on the specified TCLK events (see also Appendix A). Because this mode is a narrow bandwidth measurement, timing specifications to the RF bucket level are not required.

The new system will also support fast time plots of the BPM channels at 500 Hz update rate and will not have limits on the number on BPM channels that can be plotted from a given front end. It is understood that fast time plots can only be active when the system is in the Closed Orbit mode.

When the system is in the turn-by-turn mode, Beam Sync events (possibly both clock and hardwired events) plus a delay control both the start of the data acquisition and the RF bucket sample.

For the turn-by-turn Buffer, a new Beam Sync event will need to be defined. When this event occurs in this mode, the system will wait for the specified delay (bucket offset), trigger a flash frame, and continue taking flash frames every turn (588 buckets) for 2048 turns storing this data in the TBT buffer (Figure 4).

The Flash Turn Buffers are triggered by Beam Sync events that cause beam transfers into and out of the Main Injector. The timing card requires that a Beam Sync event be supplied at each mode change command. These transfer events are already part of the MI Beam Sync clock, the Recycler Beam Sync clock, and are present as hardwired signals in case of injections from the Booster (referred to as $xx in Figure 3 Figure 4). When the specified Beam Sync event occurs, the system will wait for a specified delay (bucket offset), trigger a flash frame, and continue taking flash frames every 588 buckets (every turn) for 512 turns, storing this data in the first available Flash Turn Buffer. The next trigger will store its data in the next Flash Turn Buffer.

There will be 20 Flash Turn Buffers to support injections and extractions. The data within the Flash Turn Buffer will contain either an injection or an extraction. If the buffer contains injection data, the first turns acquired will be zero (no beam) until the beam is injected. In this case the first buffer element with beam will be moved to the Flash Buffer. If the data in the Flash Turn Buffer is the extraction data, the first turns will contain data and the final turns will have zero. In this case the last buffer element with beam will be moved to the Flash Buffer. If the Flash Turn Buffer contains all zeros, all beam, or zeros at the start and end of the buffer, a zero will be placed in the flash buffer indicating an unknown condition.

Figure 4: Data acquisition TBT detail

The first Flash Turn Buffer will generate the averaged orbit. This will be the average of the buffer elements with beam. This averaged orbit will be used for closing the injection into the machine. The averaged orbit only needs to occur for the first Flash Turn Buffer.

Each Main Injector cycle is uniquely identified by a machine state. This state information is transmitted on the MDAT system and is valid at the TCLK reset for the main ramp. The BPM system will need to be able to listen to the state information in order to identify the required mode definitions for that State. In addition to the MDAT state information, the BPM system will need to listen for the TCLK Prepare for Beam event ($79) to start the data acquisition.

The user will need to specify information for a given MI state in order to control the analog filters and the mode that the system is operating in during the data acquisition. This information will be in the format of time and command. Using a command structure familiar to accelerator operations, which is currently implemented in the Main Injector Low Level RF (LLRF), one can specify the commands in Table 9.

MI BPM State Messages

Set_BPM_FILTER

Datum1: Sets Protons or Pbars.

- Protons

- Pbars

Datum2: Sets 2.5Mhz or 53Mhz filter.

- 2.5Mhz

- 53Mhz

Datum3: Sets the filter attenuator.

- 0db, -6db, -12db, -18db, -24db, -30db, -36db, -42db, -48db

Set_BPM_Mode

Datum1: Enables and disables message

- Enable

- Disable

Datum2: Sets Wide band TBT, wideband flash, or Narrow band (Closed orbit) mode.

- Flash/TBT

- Closed Orbit

Datum3: Sets trigger event.

- MIBS $TBT, $xx, $7A, $D6, $D8, $A1, $A6, $74, $75, $76, $7B, $7C, $7D, $7E

- RRBS $A0, $A1, $A2,$A3, $A6, $A7

Datum 4: Sets sample bucket.

- 0-588

Table 9: State message definition

It follows that this command message will use the same application page format that is currently used on ACNET page I6. An example of its use is shown in Table 10.

## Slip Stacking and NuMI

| ROW | TYPE | SIGNAL | MESSAGE | DATUM1 | DATUM2 | DATUM3 | DATUM4 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 2 3 4 5 6 | Delay Delay Delay Delay Delay Delay | 0 sec .001 sec .2 sec .45sec 1.2 sec 1.201 sec | Set_BPM_Filter Set_BPM_Mode Set_BPM_Mode Set_BPM_Mode Set_BPM_Mode Set_BPM_Mode | Proton Enable Disable Enable Enable Disable | 53 Mhz Flash/TBT Closed Orbit Closed Orbit Flash/TBT Flash/TBT | -48 Db BES MIBS $74 MIBS $79 | 200 0 |

## Pbar Acceleration

| ROW | TYPE | SIGNAL | MESSAGE | DATUM1 | DATUM2 | DATUM3 | DATUM4 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 2 2 3 4 | Delay Delay Delay Delay Delay | 0 sec .001 sec 1.6 sec 2.5sec 6.9 sec | Set_BPM_Filter Set_BPM_Mode Set_BPM_Mode Set_BPM_Filter Set_BPM_Mode | Pbar Enable Enable Pbar Enable | 2.5 MHz Pbar Flash/TBT Closed Orbit 53 MHz Flash/TBT | 0 Db RRBS $A7 0 Db MIBS $7B | 0 0 |

Table 10: Data acquisition and defition examples.

For each MI State there will be a user specified data acquisition definition. This definition will specify the filter configuration, the data type, and the timing. Due to the multi function nature of the MI, the BPM system is expected to be able to switch between both filter and data type during a given cycle with minimal data loss.

# End User Applications

## Current System

The present BPM library (BPMUTI) needs updated low level access routines, which should include the machine state as an additional readout specification. An updated Configuration Database is necessary to provide the hardware configuration information to the BPM library.

In the following we will list the current application programs and identify the changes that are needed.

## I37 BPM Control Parameters

- rewrite by cloning the Tevatron diagnostic application

## I38 BPM/BLM Hardware Tests

- rewrite by cloning the Tevatron diagnostic application

## I39 BPM/ BLM Plots /List

- addition of state selection

- update to timing information

- addition of support for average orbit

## I42 Tune From TBT BPM

- convert to use of library routines

- addition of state selection

- addition of multiple TBT buffer support

## I50 MI Orbit Smoothing / C453

- addition of state selection

## I52 Booster to MI Injection Closure

- addition of state selection

- addition of support for averaged orbit

## I73 MI Measure Chromaticity

- addition of state selection

## I75 MI Quad Move

- addition of state selection

## I92 TBT Data Analysis

- convert to use of library routines

- addition of state selection

- addition of multiple TBT buffer support

## T117 Tev Orbit Closure

- addition of state selection

- addition of support for averaged orbit

## New applications

As detailed in Section 4.3, a new application program is necessary to provide the user with the possibility to configure the BPM acquisition specifications. An implementation, following closely the LLRF application program I6, is suggested in 5.1 and 5.2.

# Maintenance and Calibration

The new MI BPM system should be equipped with a test system to diagnose problems in the raw signals from the tunnel and at every stage in the electronic front-end. Test signals should be provided to verify the performance of the electronics. They should also perform as a calibration tool able to maintain the system within the specification requirements.

# References

- For general Main Injector information see

http://www-bdnew.fnal.gov/operations/rookie_books/Main_Injector_PDF/Main_Injector.html

- D. G. Beechy and R.J. Ducar ‘Time and Data Distribution Systems at the Fermilab Accelerator’, Nucl. Instr. and meth. A247 (1986) 231

- Clock and MDAT reference

http://www-ad.fnal.gov/controls/hardware_vogel/index.html

- B. Choudhary, ‘Main Injector BPM system upgrade + NuMI beamline BPM requirements’, Beams-doc-471-v1

# Appendix A

###### MI Clock Events

| Tclk event | Beam Sync event | Description |
| --- | --- | --- |
| 20,21,23,29,2A,2B,2D,2E | none | MI Ramp Reset |
| 22 | none | MI Start of Ramp |
| 25 | none | MI Start of Flattop |
| 26 | none | MI End of Beam Operations |
| 27 | none | MI Beam has been aborted |
| 28 | none | MI Injected Beam Sync (Booster) |
| 2C | none | MI Sample Time |
| 2F | none | Abort Cleanup |
| A9 | 74 | MI=>NuMI Proton Transfer |
| 39 | 75 | MI=>Swithcyard Proton Transfer |
| 81 | 79 | MI=>Pbar target |
| 94 | 7A | Acc=>MI Pbar transfer |
| 5B | 7B | MI=>Tev Pbar Transfer |
| 5C | 7C | MI=>Tev Proton Transfer |
| 86 | 7D | MI=>Debuncher Transfer (Protons) |
| 99 | 7E | MI=>Accumulator Transfer (Protons) |
| 98 | A1 | Accumulator=>Recycler Pbar Transfer |
| 97 | A6 | Recycler=>Accumulator Proton Transfer |
| 5F | D6 | Tevatron=>MI Pbar Transfer |
| 55 | D8 | Tevatron=>MI Proton Transfer |

MI BPM Tclock Events

| Tclk event | Beam Sync event | Description |
| --- | --- | --- |
| 79 | None | MI BPM Prepare for Beam |
| 7A | None | MI BPM Write Profile Memory |
| 7B | None | MI BPM Write Display Frame |
| 7C | None | MI BPM Flash Trigger |
| 7D | None | MI BPM set Bunch Mode |

# Appendix B

## Reset $29

Use:

- Stacking and Slip Stacking at 120Gev.

Description:

- Either a single Booster batch or two slip stacked Booster batches are injected. A single batch is accelerated to 120 Gev, bunch rotated and extracted to the Pbar target.

Beam Structure:

- 53Mhz.

- A single batch with 84 bunches.

BPM Measurements

- Injection Flash 53Mhz.

- Averaged or Display at injection.

- Display frame.

- Profile frame.

- TBT

- Last turn for the Pbar transfer.

## Reset $21

Use:

- Switchyard 120Gev, Stacking and SY120, Slip Stacking and SY120.

Description:

- For dedicated SY120, up to 6 batches are injected into the ring, accelerated to 120Gev and slow extracted to the Swithcyard.

- For Stacking and SY120, a single batch is injected for the Pbar target and a short batch is injected for the Switchyard. These two injections are different intensities and placed in various locations in the machine.

Beam Structure:

- 53Mhz.

- For dedicated SY120 operation this could be up to 6 Booster batches. The batches will have between 15 and 84 bunches each.

- For mixed mode operation there is one batch for the Pbar target and a second batch for SY120. These two batches are different intensities. The pbar batch is 84 bunches and the SY120 batch is between 15 and 84 bunches

BPM Measurements

- Injection Flash 53Mhz, all injections.

- Averaged or Display at injection.

- Display frame.

- Profile frame.

- TBT

- Last turn for the Pbar target batch.

## Reset $23

Use:

- NuMI Target 120Gev, Stacking and NuMI, Slip Stacking and NuMI.

Description:

- For dedicated NuMI, up to 6 batches are injected into the ring, accelerated to 120Gev and extracted to the NuMI target.

- For Stacking and NuMI, a single batch is used for the Pbar target and up to 5 batches are injected for NuMI. These two functions are of different intensities and placed in various locations in the machine.

Beam Structure:

- 53Mhz.

- For dedicated NuMI operation this could be up to 6 Booster batches. The batches will have 84 bunches each.

- For mixed mode operation there is one batch for the Pbar target and up to 5 batches for NuMI. These batches are 84 bunches each.

BPM Measurements

- Injection Flash 53Mhz, all injections.

- Averaged or Display at injection.

- Display frame.

- Profile frame.

- TBT

- Last turn for the Pbar transfer.

- Last turn of the NuMI transfer.

## Reset $2A

Use:

- Pbar Acceleration to 150 Gev.

- Tune up of reverse injection from Tevatron with protons.

Description:

- Four 2.5Mhz bunches are transfer into the MI from either the Accumulator or the Recycler. These four bunches are rebunched into four groups of 53Mhz bunches, Accelerated to 150, and coalesced into four single 53Mhz bunches for transfer to the Tevatron.

- Reverse injection tune up from Tevatron involves injection a short batch, 30 bunches, down the Pbar transfer line. Adjustments are made to close orbit in MI.

- In the future we may accelerate in 2.5Mhz up to 27Gev before capture in 53Mhz

Beam Structure:

- 2.5Mhz for 8 Gev injection.

- 53Mhz for reverse injection.

- 53Mhz for acceleration.

- Four 2.5Mhz at injection.

- Four groups of 5-9 bunches for acceleration.

- Four single 53Mhz bunches after coalescing.

- 30x53Mhz bunches for reverse injection.

- Future unknowns.

BPM Measurements

- Injection Flash in both 53Mhz and 2.5Mhz.

- Averaged or Display at injection in both 53Mhz and 2.5Mhz.

- Display frame in both 53Mhz and 2.5Mhz.

- Profile frame in both 53Mhz and 2.5Mhz

- TBT in both 53Mhz and 2.5Mhz

- Last turn for the Pbar transfer.

## Reset $2B

Use:

- Proton Acceleration to 150 Gev.

Description:

- A short Booster batch of 5-9 bunches are injected. These bunches are accelerated to 150, coalesced into a single 53Mhz bunch, and transferred to the Tevatron.

- For Proton tune up of the Tevatron, 30x53Mhz bunches are injected into MI, accelerated and transferred to the Tevatron.

Beam Structure:

- 53Mhz.

- 5-9 high intensity bunches for proton loading of the Tevatron.

- 30 low intensity bunches for tune up of the Tevatron.

- One single 53Mhz bunch after coalescing.

BPM Measurements

- Injection Flash.

- Averaged or Display at injection.

- Display frame.

- Profile frame.

- TBT.

- Last turn for the proton transfer to Tevatron.

## Reset $2D

Use:

- A multipurpose 8Gev cycle for the tune up of beam lines with protons and the transfer of Pbars from the Accumulator to the recycler.

Description:

- For the tune up of beam lines with Protons. A short Booster batch, usually 30x53Mhz bunches, is injected in the MI and extracted to the beam lines.

- For Pbar transfer. Four 2.5Mhz bunches are injected into MI, decelerated to the Recycler energy, cogged to the correct alignment with Recycler, and transferred to the Recycler.

Beam Structure:

- 53Mhz for protons.

- 2.5Mhz for Pbars.

BPM Measurements

- Injection Flash. Both 2.5Mhz and 53Mhz.

- Averaged or Display at injection. Both 2.5Mhz and 53Mhz.

- Display frame. Both 2.5Mhz and 53Mhz.

- Profile frame. Both 2.5Mhz and 53Mhz.

- TBT. Both 2.5Mhz and 53Mhz.

- Last turn for the proton transfers to the beam lines. 53Mhz

- Last turn for Pbar transfers to Recycler. 2.5Mhz

## Resets $20, 2E

Use:

- $20 was originally reserved for deceleration of Pbars from tevatron. Currently used for studies.

- $2E is a study cycle

Description:

- Unknown since this is a study cycle, Timing and  beam structure will vary.

Beam Structure:

- 2.5Mhz and 53Mhz.

- Could be a single 2.5Mhz bunch or 6x53Mhz bunches.

- Could be a single 53Mhz bunch at high energy.

- Could a single 2.5Mhz bunch.

BPM Measurements

- Injection Flash with both 53Mhz and 2.5Mhz RF structure.

- Averaged or Display at injection.

- Display frame.

- Profile frame.

- TBT

# GLOSSARY

- Batch

This refers to a unit of beam that corresponds to the output of one Booster cycle. The Booster can hold 84 bunches of beam in one acceleration cycle. The Main Injector can hold 6 batches from the Booster.

- (Short) Batch

This refers to any batch length shorter than the 84 bunches that the Booster provides. Although each Booster batch length is 84 bunches, fewer than 84 can be extracted from the Booster into the Main Injector.

- Beam Synch

The clocks MIBS, RRBS and TVBS have their basic frequency directly related to the revolution frequency of the beam. They are derived from their respective LLRF systems. They operate at frequencies of about 7.5 MHz and produce a clock “tick” every 7 bunches. Beam Synch clocks are used for all critical timing of beam transfers between accelerators.

- Betatron wavelength

The length of one Betatron oscillation. Betatron oscillations are the transverse oscillations of particles in a circular accelerator about the equilibrium orbit. Focusing components in the magnetic field provide a restoring force for the oscillations and act to bend a particle off the equilibrium orbit back toward it.

- BPMs

Beam Position Monitors. Physical components that are designed to determine the overall position of a particle as it passes through it. Typically single plane (horizontal or vertical) with 2 pickup plates to determine proximity of a particle based on the relative amount of charge induced on the two plates. In some instances, such as the majority of the Main Injector, BPMs can have 4 pickups.

- Bucket

A bucket is the stable region in longitudinal phase space. The bucket width gives the maximum phase error or timing error at the R.F. cavity that a particle may have and still complete the whole acceleration cycle. The bucket height is the corresponding limit on momentum error.

- Bucket to bucket transfer

The process of transferring buckets from one machine to another in a synchronized fashion such that the buckets are injected at the RF zero points of the accepting machine.

- Bunch

A group of particles captured in a bucket.

- Closed orbit

This refers to the ideal orbit of particles in the accelerator. Betatron oscillations are superimposed on this orbit.

- Cycle

This refers to the length of time between successive repetitions of a periodic phenomenon. Main Injector cycles are defined by the magnet ramp.

- Harmonic number

The number of times the RF voltage oscillates during the time the beam makes one revolution around a circular accelerator. It also equals the number of RF buckets in the machine.

- MDAT

Machine data. Machine data distribution is accomplished by means of a separate communications link that is distributed throughout the accelerator complex. The data broadcast on this link is in the form of separate serially encoded data frames, each including an 8 bit code and an associated 16 bit word of data. MDAT is carried on either one of the 19-conductor cables or newly installed fiber at a 720Hz rate. This information is then decoded and used by several devices such as the ramped correction elements. The Main Injector is serviced completely by a fiber system for its MDAT signals.

- Ramp

This means to excite a magnet with a time dependant excitation current. A time varying excitation current for a magnet is used to either track changing beam energy or to reduce average power dissipation by reducing the excitation when the magnetic field is not required. At Fermilab, the Main Injector bending magnetic field varies with the energy of the accelerating protons in order to confine them within the beam pipe.

- Super-cycle

The supercycle time is the amount of time required to complete all the different machine cycles that have been defined to operate periodically as a unit. This implies that the supercycle time is periodic and that its length is dependent on the number and duration of different machine cycles defined to occur within its period. It is currently defined from page D69.

- State

A unique set of conditions that define a mode of operation. A state will dictate the path, RF structure and eventual destination of the accelerated particles. Currently, there are 31 states in the Main Injector. States can be viewed on page I6 of an active console. A given Main Injector cycle reset has a state associated with it. MI states are broadcasted by the MDAT system (I:MDAT56). This state is fixed for the lifetime of the particular cycle. A given MI cycle reset can be associated to multiple Main Injector states.

- TCLK

Tevatron Clock. The master clock for the accelerators. It operates at a 10 MHz rate and is able to transmit encoded events at a 1MHz rate. It uses an encoding scheme called bi-phase or modified Manchester coding. The coding is characterized by a transition at each 100 ns bit cell boundary, with an additional transition at mid-cell if the bit is true and no transition at mid-cell if the bit is false. There are 256 possible events (displayed in 2 digit hexadecimal numbers) that the clock can send out.

- TLG

Time Line Generator, a VME based system controlled from page D69 that defines the start and cycle times for every major system in the accelerator complex. It’s located in the MAC room.

# Footnotes

( $2E is a cycle reset preceding $2A in order to adjust the magnet injection currents to the Recycler momentum, which is about 40 MeV/c lower than the Accumulator momentum. No beam is present on the $2E reset.

( Because of the non-linearity intrinsic to the MI BPM pickups, this specification only applies when the beam is off center on only one of the planes.

# Text boxes

Digital Filter

Mode Switch

Closed Orbit Mode

Profile Frame Buffer

Display

Frame

Buffer

- Switch activated

every 2ms

when system is in closed orbit mode

- Switch

activated

by User

Configurable

Tclock

Event ($7A)

- Switch

Activated by

User configurable event ($7B)

FIFO full generates error.  Pointer reset on next state.

128 pts

All Closed Orbit Mode switches are interrupted by TBT mode.

Flash

Turn

Buffer

Switch activated by MIBS clock $xx, $7A, $D6, $D8 for first turn; $A1, $A6, $74, $75, $76, $7B, $7C, $7D, $7E for last turn and RRBS clock $A3, $A6, $A7 for first turn; $A0, A1, $A2 for last turn. Switch remains activated for  512 turns.

512 pts

20 Buffers

TBT

Mode

(see detail)

Fast

Time

Plot

Buffer

User

Data

Buffer

Data at end of FIFO is discarded

128 pts

- Switch activated

By user

configurable event

First

Turn

Closed

Orbit

Buffer

- Value derived as an average of  points with beam captured in the First Flash Turn Buffer Buffer.

Flash

Buffer

- Switch changed by SET BPM MODE message

-

Turn

by

Turn

Buffer

- Switch activated by delayed beam synch clock $??.  Switch remains activated for specified number of turns.

2048 pts

State System for MI BPMs

2.5 Mhz

53 Mhz

2.5 and 53 Mhz Data select and Filter attenuator changed by SET BPM MODE message

20 pts

First

Flash

Turn

Buffer

Switch activated by MIBS clock $xx, $7A, $D6, $D8 for first turn; $A1, $A6, $74, $75, $76, $7B, $7C, $7D, $7E for last turn and RRBS clock $A3, $A6, $A7 for first turn; $A0, A1, $A2 for last turn. Switch remains activated for  512 turns.

512 pts

TBT

Mode

First

Turn

Closed

Orbit

Buffer

Value derived from an average of the points with beam captured in the First Flash Turn Buffer.

Flash

Buffer

Turn

by

Turn

Buffer

- Switch activated by delayed beam synch clock $??.  Switch remains activated for specified number of turns.

2048 pts

Turn by Turn Mode Detail

20 pts

Second Flash Turn Buffer

512 Pts

512 pts

20Th Flash Turn Buffer

Buffer contains

The first or last turn flash data.