MI BPM COMMISSIONING CHECKLIST

Date: ____________

Service Building: MI40

1. Save an orbit with the old system (I39) 8GeV NuMI bumps, 120GeV NuMI bumps On and Off,

150 GeV NuMI 														         [ ]

Purpose: We need references.

2. Hardware Installed, Connected, Powered On (Marv will have his own checklist.  Someone should double-

check that everything is connected correctly in the service building.)                                                                                     [ ]

Purpose: Physical installation and connection of VME and XMOD hardware.

3. Software Heartbeat Verified (Steve will perform this.)                                                                                                         [ ]

Purpose:  Verify software communication between front-end and Acnet.

4. Cable offsets/polynomials checked (Brian will help with this.)                                                                                             [ ]

Purpose: These must be entered correctly to get an accurate measurement.

5. Save orbits with the new system (I39) and compare to previous orbits                                                                                [ ]

Purpose:  Quick check to identify major problems/differences from the reference.

6. Tweak beam energy to quickly see if there are any obvious problems.                                                                                [ ]

Purpose:  Another quick sign check.

7. Closed Orbit Measurements/ 3-Bumps

Purpose: Verify that position and intensity signals are reading sensibly, that no cables have been swapped either A & B or connected to the wrong channel, that sign is correct, that FTPs work, that FTP agrees with I39 flashes.

|  | Bump | FTP | FTP | FTP | FTP | Display Frm. | Display Frm. | Display Frm. | Display Frm. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Change | Pre-bump | During Bump | Post-bump | Delta | Pre-bump | During Bump | Post-bump | Delta |
| BPM Name | (mrad) | (mm) | (mm) | (mm) | (mm) | (mm) | (mm) | (mm) | (mm) |
| HP324 |  |  |  |  |  |  |  |  |  |
| VP325 |  |  |  |  |  |  |  |  |  |
| HP326 |  |  |  |  |  |  |  |  |  |
| VP327 |  |  |  |  |  |  |  |  |  |
| HP328 |  |  |  |  |  |  |  |  |  |
| VP329 |  |  |  |  |  |  |  |  |  |
| HP330 |  |  |  |  |  |  |  |  |  |
| VP331 |  |  |  |  |  |  |  |  |  |
| HP332 |  |  |  |  |  |  |  |  |  |
| VP333 |  |  |  |  |  |  |  |  |  |
| HP334 |  |  |  |  |  |  |  |  |  |
| VP335 |  |  |  |  |  |  |  |  |  |
| HP336 |  |  |  |  |  |  |  |  |  |
| VP337 |  |  |  |  |  |  |  |  |  |
| HP338 |  |  |  |  |  |  |  |  |  |
| VP339 |  |  |  |  |  |  |  |  |  |
| HP340 |  |  |  |  |  |  |  |  |  |
| VP341 |  |  |  |  |  |  |  |  |  |
| HP400 |  |  |  |  |  |  |  |  |  |
| VP401 |  |  |  |  |  |  |  |  |  |
| HP402 |  |  |  |  |  |  |  |  |  |
| VP403 |  |  |  |  |  |  |  |  |  |
| HP404 |  |  |  |  |  |  |  |  |  |
| VP405 |  |  |  |  |  |  |  |  |  |
| HP406 |  |  |  |  |  |  |  |  |  |
| VP407 |  |  |  |  |  |  |  |  |  |
| HP408 |  |  |  |  |  |  |  |  |  |
| VP409 |  |  |  |  |  |  |  |  |  |
| HP410 |  |  |  |  |  |  |  |  |  |
| VP411 |  |  |  |  |  |  |  |  |  |
| HP412 |  |  |  |  |  |  |  |  |  |
| VP413 |  |  |  |  |  |  |  |  |  |

7. Check timing for BPM where signal arrives first…might need to adjust the house delay

8. Set all other Channel delays so that the beam falls in the same position in the digitization window as the short BPM

9. Verify that the Turn-By-Turn measurement is working properly with protons from the Booster