Fermilab/BD/TEV

Beams-doc-1276

October 9, 2006

Version 02

Main Injector Beam Position Monitor Front End Software User’s Guide

Luciano Piccoli, Steve Foulkes

Fermilab, Computing Division, CEPA

Abstract

This document is geared to help past developers remember what they did, future developers to know how to build and distribute the software, commissioners to understand how to configure a system, and maintainers to help diagnose software problems. Contact mi-bpm-project@fnal.gov for comments/updates.

1	Overview	2

2	Commissioning	3

2.1	Crate Controller Configuration	3

2.1.1	Registration	3

2.1.2	Boot Parameters	3

2.1.3	Startup Script	4

2.1.4	Changing the vxWorks startup script for all houses	5

2.1.5	Changing the vxWorks startup script for a particular house	5

2.2	Echotek Board Configuration	6

2.3	Timing Module Configuration	6

3	Developing	6

3.1	Build Environment	6

3.2	Build Instructions	7

3.3	Unit Tests	7

3.3.1	GBPM Unit Tests	8

3.3.2	Echotek Driver Unit Test	9

4	Diagnostics	10

4.1	Software Statistics	10

4.1.1	bpmChangeBool	11

4.1.2	bpmChangeConfig	11

4.1.3	bpmChangeFloat	12

4.1.4	bpmClearBuffers	12

4.1.5	bpmDumpBuffer	12

4.1.6	bpmMode	13

4.1.7	bpmPeek	14

4.1.8	bpmStart	15

4.1.9	bpmShow	15

4.1.10	bpmShowAlarms	17

4.1.11	bpmShowBuffers	18

4.1.12	bpmShowConfig	19

4.1.13	bpmShowModes	21

4.1.14	bpmShowStateChanges	21

4.1.15	bpmShowTCLK	22

4.2	Trace Diagnostics	22

4.2.1	Trace Levels	24

4.3	VME Backplane Diagnostics	24

4.4	Looking at ACNET variables	31

5	Appendix: Node name/ACNET addresses	31

6	Appendix: Full VME trace	32

# Overview

This document is intended to be a guide for the front end users of the Main Injector BPM system. It is meant to help

- developers build system and verify the builds

- commissioners now how install and commission boards and crates

- maintainers diagnose problems

# Commissioning

## Crate Controller Configuration

The crate controller in a MIBPM house is a Motorola 5500 running vxWorks. Node names are selected with the convention of ibpm<id> where <id> is a two character identifier of location in the ring, e.g., ibpm10.

### Registration

General instructions can be found at:

http://www-bd.fnal.gov/controls/micro_p/rom_install.html.

IP addresses must be obtained through the Acceleration Division networking group. Before requesting an IP address, you must know the Fermilab ID and Ethernet MAC address of the board. Once that has been identified, register the board at:

http://www-bdnew.fnal.gov/Netwebrequests/net-connection.asp

Once the IP address is obtained, each host will also need an ACNET address, which consists of the above node name plus an assigned trunk and node id. Contact Brian Hendricks to procure a trunk and node id.

### Boot Parameters

Boards are currently running VxWorks v5.5. Each crate controller should have the following boot parameters:

boot device          : gei0

processor number     : 0

host name            : fecode-bd

file name            : vxworks_boot/kernel/mv5500/vxWorks

inet on ethernet (e) : 131.225.<xxx>.<yyy>:ffff0000

inet on backplane (b):

host inet (h)        : 131.225.121.145

gateway inet (g)     : 131.225.<xxx>.200

user (u)             : vxworks_boot

ftp password (pw) (blank = use rsh):

flags (f)            : 0x8

target name (tn)     : <nodename>_0x<trunkid><nodeid>

startup script (s)   : vxworks_boot/fe/mibpm/mibpmstartup

other (o)            :

Where <xxx>, <yyy>, <trunkid>, and <nodeid> are obtained in the registration step above.

### Startup Script

The Main Injector BPM software is designed to detect EchoTek board addresses at run time, so all BPMs can share the same startup script. The startup script pointed to in the boot parameters currently contains:

#------------------------------------------------------------------------

#--- General initialization

#------------------------------------------------------------------------

shellPromptSet( "Booting incomplete->" )

taskPrioritySet( taskIdFigure( "tExcTask" ), 1 )

taskPrioritySet( taskIdFigure( "tLogTask" ), 250 )

# get target name in 6 characters

targetName=malloc(20)

gethostname(targetName,20)

targetName[6]=0

hostAdd "quasar.fnal.gov","131.225.121.36"

#------------------------------------------------------------------------

# mount NFS file systems - 1217, 5143 is user vxworks_boot group bdmicrop

#------------------------------------------------------------------------

topLevel="vxworks_write/fe/mibpm"

outputDir=malloc (strlen(topLevel) + strlen(targetName))

strcpy(outputDir,topLevel)

strcat(outputDir,targetName)

nfsAuthUnixSet( "fecode-bd", 1217, 5143, 0, 0 )

nfsMount( "fecode-bd", "vxworks_boot/module/PPC604", "/controls"  )

nfsMount( "fecode-bd", "vxworks_boot/fe", "/fe"  )

nfsMount( "fecode-bd", outputDir, "/write" )

# Mount the scratch disk

nfsMount( "fecode-bd", "vxworks_write/fe/fccts1", "/write1")

# Setup VME Mappings and load the MVME5500 support libraries

ld < /fe/mibpm/mv5500sup.out

ld < /fe/mibpm/trace.mv5500.o

vmeMasterMap(0x0b, 0xa0000000, 0xa0000000, 0, 0, 1)

vmeMasterMap(0x0b, 0xe0000000, 0xeff00000, 0x44000000, 0, 2)

sysUniverseCapt

# Load up the PMC UCD Card, and all the ACNET libraries

ld (1, 1, "vxworks_boot/fe/pmcucd/mv5500/pmcclassLib-latest.o");

ld (1, 1, "vxworks_boot/module/PPC604/libpmctrig-latest.o");

set_rbf_file ("vxworks_boot/fe/pmcucd/pmcucd5500.rbf")

pmcucd_start(0)

pmcucd_set_adr (get_alteramem(0))

pmcucd_set_madr (get_alteramem(0)+0x1000)

ld < /controls/libssm-2.4.o

ld < /fe/mibpm/acnet-1.025-no-trace.o

ld < /controls/libmooc-3.7.o

# Load database libraries

ld < /controls/errlog-latest.o

ld < /controls/dbdown-latest.out

errlogInit()

ACNET_SSM_avail=0

# Load the MIBPM libraries and EchoTek config files

ld < /fe/tbpm/libculite.out

ld < /fe/echotek_mibpm/libechotek_mibpm.out

ld < /fe/gbpm/libgbpm.out-v1_57

ld < /fe/mibpm/libmibpm.out

cd( "/fe/mibpm/ini" )

ecdr814gcReadSetup("prieto-051209.ini",     15);

ecdr814gcReadSetup("count.ini",             19);

# Start Mooc

MoocNew()

MoocDeviceDownLoad(1)

# Initialize trace

clk1Stop()

clk1Start()

traceInit 50000,10,0

traceMode 1

traceGlobalOff 0,31,0

traceGlobalOn 6,6,0

traceReset

# Start the BPM

bpmStart

bpmChangeBoolConfig 37,1

# Add the injection events so that the flash task knows

# where to look for the first/last turn. BES is 0x200.

# All recycler events get ORed with 0x100, so event

# 0x7A would become 0x17A.

bpmAddInjectionEvent(0x200);

bpmAddInjectionEvent(0x7A);

bpmAddInjectionEvent(0xD6);

bpmAddInjectionEvent(0xD8);

bpmAddInjectionEvent(0x1A3);

bpmAddInjectionEvent(0x1A7);

# Finish up

prompt = malloc(30)

strcpy (prompt,targetName)

strcat (prompt,":mibpm> ")

shellPromptSet ( prompt )

free(prompt)

free(targetName)

free(outputDir)

period 600, set_date

### Changing the vxWorks startup script for all houses

The Main Injector BPM houses share a common startup script that resides in nova.fnal.gov:/fecode-bd/vxworks_boot/fe/mibpm. If the startup scripts for all houses needs to change, please follow the build instructions in subsequent section.

Note that we are loading fixed versions of the ACNET/MOOC libraries. These versions need to match the versions of ACNET/MOOC header files that are compiled into both gbpm and tbpm. When reloading a different version of mooc, you will also need to edit the _acnetheaders.h file in both packages and recompile (see build instructions below).

### Changing the vxWorks startup script for a particular house

While debugging problems at a specific house, there will sometimes be a need to have a startup script that is specific to that particular house. Since this is only a temporary state, there is no need to track the changes. Simply:

nova> cd /fecode-bd/vxworks_boot/fe/mibpm

nova> cp mibpmstartup temp_startup

Edit the temp_startup script as necessary, being sure to change the vxworks boot parameters to use this script. Please remember to reset the boot parameters to use the canonical script when done.

## Echotek Board Configuration

The configuration of the Echoteks is handled differently between the Tevatron BPMs and the Main Injector BPMs.  For the Main Injector, one setup file is used to specificy all the gray chip parameters for running with all four gray chip channels enabled.  The front end software takes care of modifiying the Echotek registers as it deems necessary.  Echotek configuration files are not reloaded to change modes.

Each board setup file contains another pointer to a file that contains information regarding a channel setup. Unless you really know what you’re doing, all channels should be configured with the same file. By convention, the channel setup files have an extension of .ch. By convention, the <abc>.ini file will read in a channel file named <abc>.ch.

The driver reads the associated files from the current working directory, so it’s important to cd to this directory in the VxWorks startup script.

## Timing Module Configuration

The timing board is described in detail in the document #2038. The base address defined for the timing board is 0x2000. If the board is set to any other address the software will not be able to contact it.

# Developing

There are 3 packages that constitute the Main Injector BPM software, they are:

- gbpm – generic bpm classes

- mibpm – bpm implementation specific to the Main Injector

- echotek_mibpm – underlying echotek driver

## Build Environment

The Main Injector BPM software uses the RFI build methodology. Please refer to Beams Document #1271 for a description of the tools and, for first time users, how to configure your account on nova.fnal.gov. Following these instructions please set your work group to rfiinst.  All code is compiled on nova.fnal.gov.

## Build Instructions

All packages are built in a similar fashion

nova-> cd ~/esd/src		# only do once; create if needed

nova-> cvs checkout gbpm	# only do this once

nova-> setup gbpm

nova-> cvs update –d 		# cvs checkout gbpm if first time

nova-> make clean

nova-> make

nova-> make development		# put in development location

nova-> make test			# put in test location

nova-> make production		# put in production location (will

# be installed at next reboot)

Library locations are:

nova.fnal.gov:/fecode-bd/vxworks_boot/fe/gpm/devgpm.out  # development

nova.fnal.gov:/fecode-bd/vxworks_boot/fe/gpm/testgpm.out # test

nova.fnal.gov:/fecode-bd/vxworks_boot/fe/gpm/libgpm.out  # production

There is a similar set of commands to build mibpm and echotek_mibpm.

Note that the Tevatron and Main Injector share the libgbpm library.  As it is not always convienent to test and updated the front ends on both machines at the same time the startup script for the Main Injector BPMs has the version of gbpm it is supposed to use hard coded.  The Tevatron BPM startup script will just load libgbpm.out, which is symlinked to the appropriate version for the Tevatron BPMs.

## Unit Tests

The three mibpm packages have unit or module tests associated with them. The unit tests are designed to verify the core functionality of the packages. The tests should be used only when adding new functionality to the software in order to check whether new code has affected existing functions. The gbpm and mibpm packages use an underlying software package called culite.

The culite product is based on CppUnitLite, developed by Michael Feathers (http://c2.com/cgi/wiki?CppUnitLite). The source code is available from the Computing Division repository: cvsuser@cdcvs.fnal.gov:/cvs/cd. Version v1_0 is installed on nova on /fecode-bd/vxworks_boot/fe/culite.

### GBPM Unit Tests

The GBPM unit tests can be executed by running the command GBPMUnitTest, which should generate the following output:

fccts0:tbpm-dev> GBPMUnitTest

traceLevel array at 0x1da1ba50, traceTid array at 0x1da1bcd8

sizeof(ENTRY_T)=32

callocing 32000 bytes

traceCircQueBegin=0x1d2561f0, traceCircQueEnd=0x1d25deef (last byte of

allocation)

cfreeing

*** Starting Unit Tests ***

1) DataEntry

There were no test failures

2) TestDataEntry

There were no test failures

3) Packer

There were no test failures

4) TestPacker

There were no test failures

5) Event

There were no test failures

6) EventListener

There were no test failures

7) EventGenerator

There were no test failures

8) InterruptEventGenerator

There were no test failures

9) HardwareInterruptEventGenerator

There were no test failures

10) TimeEventGenerator

There were no test failures

11) AuxTimeEventGenerator

There were no test failures

12) TCLKEventGenerator

There were no test failures

13) AlarmGenerator

There were no test failures

14) DataAcquisitionTask

There were no test failures

15) DataBuffer

There were no test failures

16) CircularDataBuffer

There were no test failures

17) TestControl

There were no test failures

18) TestControlTask

There were no test failures

19) AlarmTask

There were no test failures

20) StateChangeEventGenerator

Please set V:PING to 5 or 10

Failure: "Listener didn't receive ping event in time" line 77 in

StateChangeEventGeneratorUnitTest.cpp

Failure: "For succeeding in this unit test you must set V:PING to 5 or

10" line 78 in StateChangeEventGeneratorUnitTest.cpp

Failure: "Listener didn't receive ping event in time" line 82 in

StateChangeEventGeneratorUnitTest.cpp

Failure: "For succeeding in this unit test you must set V:PING to 5 or

10" line 83 in StateChangeEventGeneratorUnitTest.cpp

There were 4 failures

21) ConfigStateChangeListener

There were no test failures

22) ConfigStateChangeTask

There were no test failures

23) FTPClass

There were no test failures

*** Unit Tests End ***

value = 2543116 = 0x26ce0c = cout

Each test (from 1 to 23) can be executed individually. The entry point is the name of the test followed by UnitTest (e.g. AlarmTaskUnitTest).

The StateChangeEventGenerator unit test expects the state device V:PING to be changed to 10 after the test is started. It waits for 5 seconds. If the state device does not change within this period the test fails. In order to change the value of V:PING one should use the ACL script window available through the ACNET pages. The command ‘set V:PING=10’ should be issued.

### Echotek Driver Unit Test

The Echotek driver uses a different testing structure. The command ecdr814gcUnitTest can be used to test the driver.

The addresses used for the Main Injector BPM project always start at 0x6000, increasing by 0x1000. The Echotek board closer to the processor board always has the lowest address. Neighbor boards to the right have higher base addresses.

The setup file index available can be showed by using the command ecdr814gcShowSetupAll:

mibpmt:mibpm-dev> ecdrecdr814gcShowSetupAll

15: prieto-051209.ini

19: count.inivalue = 0 = 0x0

mibpmt:mibpm-dev>

The setup index 19 contains the most basic setup for testing an Echotek board. It sets the board to a simple counting mode and the unit test checks if the data generated is correct. The following is a sample output from the unit test:

ibpmt:mibpm-dev> ecdr814gcUnitTest

[452a739433756bb4,VXWORKS,W,0,tShell] Looking for boards starting at 0x4000, using setup 19

[452a739433756bb4,VXWORKS,W,0,tShell] Candiate board at: 0x6000, local address: 0xf0f06000

[452a739435720c08,VXWORKS,W,0,tShell] Testing the old driver, all channels...

[452a73943a699cda,VXWORKS,W,0,tShell] Transfered data from 1 boards in 79635 micro seconds, at a rate of 51434 KB/s

[452a73943b67ed04,VXWORKS,W,0,tShell] Count sequences are correct.

[452a739500cb732e,VXWORKS,W,0,tShell] Testing the new driver, all channels...

[452a739604c4b3fe,VXWORKS,W,0,tShell] Transfered data from 1 boards in 79623 micro seconds, at a rate of 51442 KB/s

[452a739606c15452,VXWORKS,W,0,tShell] Count sequences are correct.

[452a739606c15452,VXWORKS,W,0,tShell] Testing the new driver, odd channels...

[452a739708bdf4a6,VXWORKS,W,0,tShell] Transfered data from 1 boards in 39831 micro seconds, at a rate of 51417 KB/s

[452a739709bc44d0,VXWORKS,W,0,tShell] Count sequences are correct.

[452a73970aba94fa,VXWORKS,W,0,tShell] Testing the new driver, even channels...

[452a73980cb7354e,VXWORKS,W,0,tShell] Transfered data from 1 boards in 39831 micro seconds, at a rate of 51417 KB/s

[452a73980db58578,VXWORKS,W,0,tShell] Count sequences are correct.

[452a73980eb3d5a2,VXWORKS,W,0,tShell] Testing the new driver, multi chain (closed orbit)...

[452a73990eb3d5a2,VXWORKS,W,0,tShell] Transfered data from 1 boards in 114 micro seconds, at a rate of 0 KB/s

[452a73990fb225cc,VXWORKS,W,0,tShell] Count sequences are correct.

[452a73990fb225cc,VXWORKS,W,0,tShell] Testing the new driver, multi chain (turn by turn)...

[452a739a0fb225cc,VXWORKS,W,0,tShell] Transfered data from 1 boards in 14649 micro seconds, at a rate of 51198 KB/s

[452a739a11aec620,VXWORKS,W,0,tShell] Count sequences are correct.

value = 0 = 0x0

mibpmt:mibpm-dev>

# Diagnostics

Once the system is up and running, how do you start analyzing problems? There are many diagnostic tools available. Choosing the right one to most expediently analyze the problem is a combination of experience and instinct. The following sections describe various diagnostic information that are available in the tbpm software.

## Software Statistics

The bpm software internally keeps structures that one can peek at through the vxworks shell. The bpmHelp command will list the most current set of commands that are available:

mibpmt:mibpm-dev> bpmHelp

BPMHelp> help

list - display list of available BPM commands.

help - print this help.

quit - exit BPM help.

BPMHelp> list

Commands available:

bpmChangeConfig         bpmClearBuffers

bpmDumpBuffer           bpmEnableState

bpmEnableStates         bpmEnableTBT

bpmFilter               bpmMIFilter

bpmHelp                 bpmMode

bpmPeek                 bpmDumpState

bpmEchoTekPoolBC        bpmDelayScan

bpmShow                 bpmStart

bpmStop                 bpmShowAlarms

bpmShowBuffers          bpmShowCommands

bpmShowConfig           bpmShowStateChanges

bpmShowModes            bpmShowTCLK

BPMHelp> quit

value = 0 = 0x0

mibpmt:mibpm-dev>

The following are the description for each of the commands available through the bpmHelp command and a sample output where applicable.

### bpmChangeBool

BPMHelp> bpmChangeBool

----------------------------------------------------

bpmChangeBool (index, value)

Description:

Assign a new value to the boolean variable at the given index

Parameters:

index: entry number in the config manager to be changed

value: boolean value to be assigned (0 or 1)

----------------------------------------------------

This command can be used to change any booleans shown by the bpmShowConfig command. There are only a few boolean variables defined in the system, most of them used for development and debugging purposes.

+-------- Booleans -----------------------------------+

| 20    - TEST_SOURCE                   false

| 21    - DIAGNOSTIC_MODE               false

| 24    - DATA LOGGER CLOSED ORBIT      false

| 25    - DATA LOGGER PROFILE           false

| 26    - DATA LOGGER DISPLAY           false

| 27    - DATA LOGGER USER              false

| 28    - DATA LOGGER FLASH             false

| 29    - DATA LOGGER FLASH AVERAGE     false

| 30    - DATA LOGGER FLASH TURN        false

| 31    - DATA LOGGER TURN BY TURN      false

| 32    - DATA LOGGER CALIBRATION       false

| 33    - IGNORE PREPARE FOR BEAM       false

| 34    - ADD_OFFSETS                   true

| 37    - ENABLE_CABLE_OFFSETS          false

+-------------   End of configuration   -------------+

### bpmChangeConfig

BPMHelp> bpmChangeConfig

----------------------------------------------------

bpmChangeConfig (index, value)

Description:

Assign a new value to the integer variable at the given index

Parameters:

index: entry number in the config manager to be changed

value: integer value to be assigned

----------------------------------------------------

This command can be used to change any integers shown by the bpmShowConfig command.

Be careful when using this command. You may change by mistake important system parameters.

+-------- Integers ----------------------------------+

| 20    - CLOSED_ORBIT_RATE             500

| 21    - ECHOTEK_CONFIG                15

| 22    - STATE                         21

| 24    - TBT_TRIGGERS                  512

| 23    - TCLK_RESET                    0x23

| 25    - FILTER_PARTICLE               0

| 26    - FILTER_FREQUENCY              0

| 27    - FILTER_ATTENUATION            -26107

| 28    - TBT_ARM_EVENT                 0xda

| 29    - TBT_TURN_DELAY                0

| 30    - TBT_BUCKET_DELAY              0

| 31    - FLASH_ARM_EVENT               512

| 32    - FLASH_TURN_DELAY              0

| 33    - FLASH_BUCKET_DELAY            0

| 34    - ACTIVE_MEASUREMENT            0

| 38    - DIAG_CHANNELA                 0

| 39    - DIAG_CHANNELB                 1

| 36    - DIAG_DATA_MODE                11

| 37    - DIAG_STATE                    0

| 40    - ACQSPEC_MODE                  0

| 46    - BURST_COUNT                   1

| 47    - RECV_DATA_SKIP                21

| 48    - CO_GC_ACTIVE                  2

| 49    - TBT_GC_ACTIVE                 4

| 50    - CO_DEC                        30207

| 51    - TBT_DEC                       28679

| 53    - BEAM_THRESHOLD                        250

| 61    - DATA LOGGER BPM START                 0

| 62    - DATA LOGGER BPM STOP                  1

| 65    - CLOSED ORBIT CLEAR                    28

### bpmChangeFloat

BPMHelp> bpmChangeFloat

----------------------------------------------------

bpmChangeFloat (index, value)

Description:

Assign a new value to the float variable at the given index

Parameters:

index: entry number in the config manager to be changed

value: float value to be assigned

----------------------------------------------------

This command can be used to change any floats shown by the bpmShowConfig command. There are only a few float variables defined in the system. They are used for defining thresholds, default values and parameters.

### bpmClearBuffers

BPMHelp> bpmClearBuffers

----------------------------------------------------

bpmClearBuffers (state)

Description:

Clear all buffers in the system

Parameters:

state: MI state (0 to 255)

----------------------------------------------------

All buffers for a particular state are cleared with this command. Pointers are reset and data is zeroed.

### bpmDumpBuffer

BPMHelp> bpmDumpBuffer

----------------------------------------------------

bpmDumpBuffer (bpm, buffer, state)

Description:

Display all contents of the buffer for a given bpm. The position/intensities and I/Q pairs are shown

Parameters:

bpm: 0 through 11

buffer:

closed orbit (11)

profile (12)

display (13)

user buffer (14)

turn-by-turn (15)

flash (16)

flash average (17)

flash turn (18 - 38)

state: MI state (0 to 255)

----------------------------------------------------

The buffer index starts at 11 instead of 0. Lower buffer indexes are reserved for internal use and tests.

Data from all BPMs for the specified buffer is written to the standard output. The output can be redirected to a file. This command is an alternative for ACNET when reading data from the front-ends. The following example shows how to read all fast abort data for all front-end BPMs:

mibpmt:mibpm-dev> bpmDumpBuffer 0,11 > /write1/closedOrbit.data

value = 0 = 0x0

mibpmt:mibpm-dev>

The output generated has the following pattern:

### BPM 0

1010>[-0.845604, 115.784] - [-57,-18] [-56,1]

1011>[-0.864243, 115.444] - [-49,34] [-33,45]

1012>[-0.744931, 114.862] - [-3,59] [14,54]

1013>[-0.883092, 116.123] - [30,-52] [11,-55]]

1014>[-0.803378, 116.031] - [-21,-56] [-35,-44]

…

The first number (e.g. 1011>) can be ignored. It is the entry number within the buffer. The following numbers are the [proton position, proton intensity][anti-proton position, anti-proton intensity]. The raw I and Q values are printed after the dash. The first two [I,Q] pairs are from the proton channels and the last two are from the anti-proton channels.

Even though the front-end may have only a few BPMs (e.g. 60S has 20), the system allocates buffers for the maximum of 40 BPMs. Values for the non-existing BPMs are still printed out with the bpmDumpBuffer command, they will be all zeros.

Please check section 4.1.17 for potential problems regarding this command.

for potential problems regarding this command.

### bpmMode

BPMHelp> bpmMode

----------------------------------------------------

bpmMode (mode, turns)

Description:

Change system the mode of operation

Parameters:

mode: 0 for displaying list of modes

turns: number of turns if in turn-by-turn mode

----------------------------------------------------

The bpmMode command can be used to force the system to enter a specific mode of operation. The system changes mode automatically according to TCLKs received.

If the system is not in the expected mode of operation, e.g. in turn-by-turn or abort mode when it should be in the closed orbit mode, one can use the command 'bpmMode 10' to force the front-end to get back to the closed orbit mode.

### bpmPeek

BPMHelp> bpmPeek

----------------------------------------------------

bpmPeek (bpm, buffer, position, state)

Description:

Display bpm entry

Parameters:

bpm:

buffer:

closed orbit (11)

profile (12)

display (13)

user buffer (14)

turn-by-turn (15)

flash (16)

flash average (17)

flash turn (18 - 38)

position: entry position

state: MI state (0 to 255)

The bpmPeek command can be used to check the most recent data from any of the system buffers. If no parameters are passed to the command, the most recent value from the first bpm on the closed orbit buffer (index 20) is displayed. The following is an example of a bpmPeek  command.

ibpm10:mibpm> bpmPeekbpmPeek

=== BPM 0 [4095] == ClosedOrbitBuffer [0] ===

Status: 0x0

Position=1.5332 intensity=7747.06

I/Q Pairs: [-1576 -3765] [-2928 -3201]

Turn Timestamp: FRI OCT 13 18:20:37 2006

Time in Cycle: 1.2

====================================

value = 0 = 0x0

ibpm10:mibpm>

The status shows the intensity for the BPM is above the predefined threshold. The intensity thresholds are available through the bpmShowConfig command.

The next line shows the calculated positions and intensities, followed by the raw I and Q values.

There are two timestamps displayed. The first (Turn Timestamp) is a course time of the day timestamp. The Time In Cycle is a high resolution timestamp that is zeroed when the cycle begins.

Please check section 4.1.17 for potential problems regarding this command.

### bpmStart

BPMHelp> bpmStart

----------------------------------------------------

bpmStart (rate, source, useTCLK, bsyncBase)

Description:

Start up the TBPM system

Parameters:

rate: closed orbit rate (in Hz)

source: 0 (EchoTek) - 1 (TestEchoTek)

useTCLK: 0 (enable) - 1 (disable)

bsyncBase: BSYNC decoder base address

----------------------------------------------------

The bpmStart command is invoked by the production startup script. When using the development startup script the user should use this command to start the system. Usually the ACNET/MOOC processes are already started by the script.

If no parameters are specified to the command the system is started with the closed orbit rate set to 50 Hz. For starting it at 500 Hz the first parameter should be specified. The next three parameters should be used for development and debugging only. Source defines whether the Echotek boards are used or data are generated by software (TestEchoTek); useTCLK enables or disables the decoding of TCLK signals; and bsyncBase defines the base address of the timing board.

### bpmShow

BPMHelp> bpmShow

----------------------------------------------------

bpmShow ()

Description:

Display general system information

----------------------------------------------------

The bpmShow command displays general information about the system. A typical output of this command consists of three parts. The first shows the current mode of operation (for all available modes use the command bpmMode without arguments), the state (usually the same as the mode), the closed orbit frequency, status of the diagnostics (on or off) and the versions of the GBPM and MIBPM libraries.

If the mode of operation shown by bpmShow is not in the expected mode of operation, e.g. in turn-by-turn or abort mode when it should be in the closed orbit mode, one can use the command 'bpmMode 10' to force the front-end to get back to the closed orbit mode.

The second part of the report contains information regarding the Echotek boards and their configuration. The number of boards is shows, followed by the current burst count (1 for closed orbit mode and 2048 for turn-by-turn mode), number of triggers, samples to be skipped and channel pair delays. The versions of the driver as well as the FPGA code are displayed. The following fields show information about calibration. Currently the upper portion (Survey offsets used (house)) is being used, but the system is expected to start using the data shown on the “Calibration Data” section. That calibration data is read from a database and pushed down to the front-ends.

The last part contains information about the timing system (please refer to document 2083 for more details about the timing system).

If the system is not taking data, e.g. the values on Fast Time Plots are stuck on the same value, please use the bpmShow command and check if the mode of operation is correct of if there is some missing signal reported by the timing system report, e.g. 53MHz: not present.

ibpm10:mibpm> bpmShow

MIBPM system started on FRI OCT 13 11:00:15 2006

Mode:   0 (IDLE)

Time in cycle: 1.105e+09 ms

GBPM version: v1_58

MIBPM version: v2_89

Flash Buffers free: 398

MIBPMEchoTekPool information:

Boards: 9

Burst:  1

Trig:   512

Skip:   0

Proton delays:  13 11 11 12

21 13 16 10

16 13 12 10

12 10 9 8

10 10 9 7

12 9 16 26

10 18 24 31

10 19 29 41

10 10 10 10

Pbar delays:    32 25 17 9

47 30 25 10

42 31 21 10

36 26 16 9

34 27 20 10

25 15 14 15

12 13 13 13

12 12 14 17

10 12 12 12

Driver version: v1_16

FPGA version:   2000007

BPM offsets used (MI-10) [ENABLED]:

[0/0/0].I:HP624=-0.88, [1/1/0].I:VP625=-1.48, [2/0/0].I:HP626=0.55, [3/1/0].I:VP627=0.57,

[4/0/0].I:HP628=-0.5, [5/1/0].I:VP629=-0.59, [6/0/0].I:HP630=0.33, [7/1/0].I:VP631=-0.33,

[8/0/0].I:HP632=-0.18, [9/1/0].I:VP633=0.09, [10/0/0].I:HP634=-0.24, [11/1/0].I:VP635=0.41,

[12/0/0].I:HP636=0.07, [13/1/0].I:VP637=-0.66, [14/0/0].I:HP638=0.78, [15/1/0].I:VP639=0.9,

[16/0/0].I:HP640=0.2, [17/1/0].I:VP641=-0.77, [18/0/0].I:HP100=-0.3, [19/1/2].I:VP101=-0.6137,

[20/0/2].I:HP101=-0.0377, [21/0/1].I:HP102=-0.44, [22/1/0].I:VP103=0.23, [23/0/0].I:HP104=1.64,

[24/1/0].I:VP105=-0.22, [25/0/0].I:HP106=-0.44, [26/1/0].I:VP107=0.32, [27/0/0].I:HP108=-1.66,

[28/1/0].I:VP109=2.1, [29/0/0].I:HP110=0.2, [30/1/0].I:VP111=0.35, [31/0/0].I:HP112=0.89,

[32/1/0].I:VP113=-0.12,

Cable scale factor (2.5 MHz)[ENABLED]:

[0].I:HP624=0.9946, [1].I:VP625=0.9946, [2].I:HP626=1.01, [3].I:VP627=1,

[4].I:HP628=0.9942, [5].I:VP629=1.002, [6].I:HP630=1.019, [7].I:VP631=0.9954,

[8].I:HP632=1.002, [9].I:VP633=0.9925, [10].I:HP634=1, [11].I:VP635=0.998,

[12].I:HP636=1.002, [13].I:VP637=0.999, [14].I:HP638=0.9997, [15].I:VP639=0.9949,

[16].I:HP640=0.9999, [17].I:VP641=1, [18].I:HP100=0.9981, [19].I:VP101=0.9989,

[20].I:HP101=0.9993, [21].I:HP102=0.9979, [22].I:VP103=0.9983, [23].I:HP104=0.999,

[24].I:VP105=1, [25].I:HP106=0.9984, [26].I:VP107=1.001, [27].I:HP108=0.9967,

[28].I:VP109=1.002, [29].I:HP110=0.9968, [30].I:VP111=1.001, [31].I:HP112=1.005,

[32].I:VP113=0.9993,

Cable scale factor (53 MHz)[ENABLED]:

[0].I:HP624=0.8352, [1].I:VP625=0.9764, [2].I:HP626=1.273, [3].I:VP627=0.9906,

[4].I:HP628=0.9302, [5].I:VP629=1.022, [6].I:HP630=1.181, [7].I:VP631=0.9341,

[8].I:HP632=1.008, [9].I:VP633=0.9828, [10].I:HP634=0.9976, [11].I:VP635=1.018,

[12].I:HP636=0.9926, [13].I:VP637=1.008, [14].I:HP638=1.019, [15].I:VP639=0.9883,

[16].I:HP640=0.9996, [17].I:VP641=0.9934, [18].I:HP100=0.9905, [19].I:VP101=0.9907,

[20].I:HP101=0.9907, [21].I:HP102=0.9932, [22].I:VP103=0.99, [23].I:HP104=0.9953,

[24].I:VP105=0.9934, [25].I:HP106=0.9929, [26].I:VP107=0.995, [27].I:HP108=0.9925,

[28].I:VP109=1.003, [29].I:HP110=0.985, [30].I:VP111=0.9839, [31].I:HP112=1.026,

[32].I:VP113=1.024,

53Mhz Closed Orbit Threshold: 250

2.5Mhz Closed Orbit Threshold: 250

53Mhz Turn By Turn Threshold: 250

2.5Mhz Turn By Turn Threshold: 250

Timing System information:

Proton delays:

House delay:  [BES 1288], [MIBS 1147]

Board delays: 26 17 11 8 5 0 33 77 133 0

Pbar delays:

House delay:  3048

Board delays: 234 173 118 68 22 0 5 10 18 0

Transition Board 53Mhz Gains: 69 53 37 21 5 0 11 21 43 0

Transition Board 2.5Mhz Gains: 0 0 0 0 0 0 0 0 0 0

Timing Signal Generator information:

TB Control Registers: 0x0, 0x10, 0x0, 0x43

Acquisition disabled

BSYNC: disabled TCLK: enabled   53MHz: present

Turn event:   0xaa      present

Start event:  0x0       present

Trig. delay:  0

Global delay: 0

House delay:  1288

Board delays: 26 17 11 8 5 0 33 77 133 0

Turns counter: 0

Missing turn: -1593180160

FPGA version: 3.2.0

BSYNC received: 0x55 0x74 0x77 0xaa 0xbf 0xed 0xff

TCLK received: 0x0 0x2 0x7 0x8 0x9 0xa 0xb 0xc 0xd 0xe

0xf 0x10 0x11 0x12 0x17 0x18 0x19 0x1a 0x1b 0x1c

0x1d 0x1e 0x1f 0x22 0x23 0x24 0x25 0x26 0x27 0x28

0x2c 0x2d 0x2f 0x4c 0x50 0x52 0x53 0x57 0x72 0x79

0x7a 0x7b 0x7c 0x7d 0x7e 0x8f 0xa4 0xa5 0xa6 0xa8

0xa9 0xad 0xb8 0xe3 0xeb 0xed 0xef 0xf3 0xfa

value = 0 = 0x0

ibpm10:mibpm>

### bpmShowAlarms

BPMHelp> bpmShowAlarms

----------------------------------------------------

bpmShowAlarms (enabled)

Description:

Show current alarms in the system

Parameters:

enabled: 0 (only enabled alarms) - 1 (all alarms)

----------------------------------------------------

This command displays a summary of the alarms detected by the front-end system. The current version (GBPM v1_58 and MIBPM v2_89) defines the following alarms:

0 – Missing TCLK signal

1 – Missing MIBSYNC signal

2 – Missing RRBSYNC signal

3 – Missing RFClock Signal

4 – Missing MDAT Signal

8 – Timing Board Error

9 – Echotek 1 Error

10 – Echotek 2 Error

11 – Echotek 3 Error

12 – Echotek 4 Error

13 – Echotek 5 Error

14 – Echotek 6 Error

15 – Echotek 7 Error

16 – Echotek 8 Error

17 – Echotek 9 Error

18 – Echotek 10 Error

19 – Transition Board 1 Error

20 – Transition Board 2 Error

21 – Transition Board 3 Error

22 – Transition Board 4 Error

23 – Transition Board 5 Error

24 – Transition Board 6 Error

25 – Transition Board 7 Error

26 – Transition Board 8 Error

27 – Transition Board 9 Error

28 – Transition Board 10 Error

### bpmShowBuffers

BPMHelp> bpmShowBuffers

----------------------------------------------------

bpmShowBuffers (state)

Description:

Display current status/info of system buffers

Parameters:

state: MI state (0 to 255)

----------------------------------------------------

The bpmShowBuffers command can be used to monitor the current number of elements in all the buffers available in the system. An example of output is shown below. For each buffer the current number of elements and the maximum number of elements is shown. The output also informs if there are any blocked tasks trying to access the buffer. Buffers are shared resources between the system tasks and can be accessed only after the task has acquired a lock associated with the buffer. If the task is waiting for the buffer lock then its name will be printed out as a blocked task.

mibpmt:mibpm-dev> bpmShowBuffers

System Buffers (state=0):

Elemts  Size    Blocked Tasks

VGATestBuffer    0       6       None

ClosedOrbitBuffer    4096    4096    None

CalibrationBuffer    0       2048    None

ProfileBuffer    7       128     None

DisplayBuffer    1       1       None

UserBuffer    0       128     None

FlashBuffer    6       20      None

FlashAverageBuffer    1       1       None

TurnByTurnBuffer    0       2048    None

FlashTurnBuffers [6] :

FlashTurnBuffer    512     512     None

FlashTurnBuffer    512     512     None

FlashTurnBuffer    512     512     None

FlashTurnBuffer    512     512     None

FlashTurnBuffer    512     512     None

FlashTurnBuffer    512     512     None

value = 0 = 0x0

mibpmt:mibpm-dev>

The bpmPeek and bpmDump commands all need to access to get the buffer lock before accessing the data. If those commands are killed in the middle of the processing (after acquiring the lock) then the buffer will not be accessible by other tasks. For example, if there was a bpmPeek command invoked to check the latest closed orbit value and the user killed it using ^C before it finished, then the system will not be able to add new values to the buffer and bpmShowBuffer will show that the ClosedOrbit task is blocked.

### bpmShowConfig

BPMHelp> bpmShowConfig

----------------------------------------------------

bpmShowConfig ()

Description:

Display detailed system configuration

----------------------------------------------------

The front-end software has several configuration parameters that can be changed by the user while the system is running (bpmChangeConfig, bpmChangeFloat, bpmChangeBool). Some of those parameters are used at initialization time and event if the user changes them later the system will not be affected by its change. For example the CLOSED_ORBIT_RATE (index 21) value is used only when the system is started.

The command shows all variables defined in the system at once. One can always redirect the command output to a file if it is not possible to see all values. The following is a sample output:

mibpmt:mibpm-dev> bpmShowConfig

+--------------- BPM Configuration      ---------------+

| Index Variable                        Value

+----------------------------------------------------+

|         Integers

| 0     - DEFAULT_EVENT_QUEUE_SIZE      100

| 1     - TCLK_REG_BASE                 0x6000

| 2     - TCKL_MEM_BASE                 0x8000000

| 3     - BSYNC_REG_BASE                0x2000

| 4     - CONTROL_TASK_PRIORITY         15

| 5     - ALARM_TASK_PRIORITY           42

| 5     - ALARM_TASK_PRIORITY           42

| 7     - BSYNC_START_EVENT             0x400

| 8     - BSYNC_TURN_EVENT              0xaa

| 9     - CONTROL_TASK                  0x10ae7bc0

| 10    - MODE                          0

| 11    - MODE_CHANGE_TIME              1160410958

| 12    - START_TIME                    1160410159

| 13    - CYCLE_START_SEC               0

| 14    - CYCLE_START_NANO              0

|        Booleans

| 0     - DEFAULT_READOUT               false

| 1     - USE_TEST_TCLK                 false

| 2     - USE_TEST_TIMING               false

+-------------   End of configuration   -------------+

+-------------   MI BPM Configuration   --------------+

| Index Variable                        Value

+----------------------------------------------------+

+-------- Integers ----------------------------------+

| 20    - CLOSED_ORBIT_RATE             500

| 21    - ECHOTEK_CONFIG                15

| 22    - STATE                         5

| 24    - TBT_TRIGGERS                  512

| 23    - TCLK_RESET                    0x23

| 25    - FILTER_PARTICLE               0

| 26    - FILTER_FREQUENCY              0

| 27    - FILTER_ATTENUATION            5

| 28    - TBT_ARM_EVENT                 0xda

| 29    - TBT_TURN_DELAY                0

| 30    - TBT_BUCKET_DELAY              0

| 31    - FLASH_ARM_EVENT               512

| 32    - FLASH_TURN_DELAY              0

| 33    - FLASH_BUCKET_DELAY            0

| 34    - ACTIVE_MEASUREMENT            0

| 38    - DIAG_CHANNELA                 0

| 39    - DIAG_CHANNELB                 1

| 36    - DIAG_DATA_MODE                11

| 37    - DIAG_STATE                    0

| 40    - ACQSPEC_MODE                  0

| 46    - BURST_COUNT                   1

| 47    - RECV_DATA_SKIP                21

| 48    - CO_GC_ACTIVE                  2

| 49    - TBT_GC_ACTIVE                 4

| 50    - CO_DEC                        30207

| 51    - TBT_DEC                       28679

| 53    - BEAM_THRESHOLD                        250

| 61    - DATA LOGGER BPM START                 0

| 62    - DATA LOGGER BPM STOP                  1

| 65    - CLOSED ORBIT CLEAR                    28

+--------  Floats -------------------------------------+

+-------- Booleans -----------------------------------+

| 20    - TEST_SOURCE                   false

| 21    - DIAGNOSTIC_MODE               false

| 24    - DATA LOGGER CLOSED ORBIT      false

| 25    - DATA LOGGER PROFILE           false

| 26    - DATA LOGGER DISPLAY           false

| 27    - DATA LOGGER USER              false

| 28    - DATA LOGGER FLASH             false

| 29    - DATA LOGGER FLASH AVERAGE     false

| 30    - DATA LOGGER FLASH TURN        false

| 31    - DATA LOGGER TURN BY TURN      false

| 32    - DATA LOGGER CALIBRATION       false

| 33    - IGNORE PREPARE FOR BEAM       false

| 34    - ADD_OFFSETS                   true

| 37    - ENABLE_CABLE_OFFSETS          false

+-------------   End of configuration   -------------+

value = 0 = 0x0

mibpmt:mibpm-dev>

### bpmShowModes

BPMHelp> bpmShowModes

----------------------------------------------------

bpmShowModes ()

Description:

Show the last N modes of operation of the system

----------------------------------------------------

This command shows the last 500 modes of operation of the system. Whenever the system enters a new mode of operation (e.g. when a turn-by-turn is requested) a new entry is added. Each entry contains shows the operating mode and the time when the mode switch completed.

mibpmt:mibpm-dev> bpmShowModes

Mode History (elements = 500)

0) Mode ? - FRI OCT 13 13:48:08 2006

1) Mode FLASH - FRI OCT 13 13:48:08 2006

2) Mode CLOSED ORBIT - FRI OCT 13 13:48:08 2006

3) Mode IDLE - FRI OCT 13 13:48:09 2006

4) Mode ? - FRI OCT 13 13:48:10 2006

5) Mode FLASH - FRI OCT 13 13:48:10 2006

6) Mode CLOSED ORBIT - FRI OCT 13 13:48:10 2006

7) Mode IDLE - FRI OCT 13 13:48:11 2006

8) Mode ? - FRI OCT 13 13:48:12 2006

9) Mode FLASH - FRI OCT 13 13:48:12 2006

10) Mode CLOSED ORBIT - FRI OCT 13 13:48:12 2006

…

491) Mode IDLE - FRI OCT 13 13:52:13 2006

492) Mode ? - FRI OCT 13 13:52:14 2006

493) Mode FLASH - FRI OCT 13 13:52:14 2006

494) Mode CLOSED ORBIT - FRI OCT 13 13:52:14 2006

495) Mode IDLE - FRI OCT 13 13:52:15 2006

496) Mode ? - FRI OCT 13 13:52:16 2006

497) Mode FLASH - FRI OCT 13 13:52:16 2006

498) Mode CLOSED ORBIT - FRI OCT 13 13:52:16 2006

499) Mode IDLE - FRI OCT 13 13:52:17 2006

The output of the commands bpmShowModes, bpmShowStates, bpmShowTCLK and bpmShowConfig can be used to generate valuable information when the system is not responding correctly (e.g. a system is stuck in injection mode). The output of those commands should be saved for further analysis. The command output can be saved to a file by redirecting the output to a file, e.g. “bpmShowMode > /write1/A0_MODES”.

### bpmShowStateChanges

BPMHelp> bpmShowStateChanges

----------------------------------------------------

bpmShowStateChanges ()

Description:

Show the list of last N state changes

----------------------------------------------------

The bpmShowStateChanges displays a history of the state changes recorded by the front-end. The number of state change history is set to 1000. Every line shows the device identifier and the new state value along with the time stamp when the change was received. The state id can be translated to a device name using the application page D80.

mibpmt:mibpm-dev> bpmShowStateChanges

State Change History (elements = 1000)

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:46 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:48 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:50 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:52 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:54 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:56 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:20:58 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:21:00 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:21:02 2006

State 21 - TCLK 35 - Commands 3 - Last 2 - Completed 1 - FRI OCT 13 18:21:04 2006

…

### bpmShowTCLK

BPMHelp> bpmShowTCLK

----------------------------------------------------

bpmShowTCLK ()

Description:

Display the TCLK history

----------------------------------------------------

The bpmShowTCLK command shows a history of the last 500 TCLKs received by the front-end. Along with the TCLK received there are three timestamps. The first one, preceded by (T) is the timing board TCLK timestamp, which records the timestamp at the last TCLK 0x4B. The second timestamp (B) is also from the timing boards, however it is the timestamp of the last BYNC 0x7C (last injection). Finally a time of the day timestamp at the time the TCLK was received is displayed.

tbpma0:tbpm> bpmShowTCLK

TCLK History (elements = 500)

Last TCLK received on TUE JUL 05 05:53:21 2005

0) TCLK 0x75 - (T) 0xa967c48c553 - (B) 0xaf6197f7 - SUN JUL 03 23:51:21 2005

1) TCLK 0x75 - (T) 0xa967c48c553 - (B) 0xaf6197f7 - SUN JUL 03 23:51:27 2005

2) TCLK 0x75 - (T) 0xa967c48c553 - (B) 0xaf6197f7 - SUN JUL 03 23:51:33 2005

3) TCLK 0x75 - (T) 0xa967c48c553 - (B) 0xaf6197f7 - SUN JUL 03 23:51:41 2005

4) TCLK 0x4b - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:53:15 2005

5) TCLK 0xc1 - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:53:45 2005

6) TCLK 0x75 - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:53:56 2005

7) TCLK 0x75 - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:54:18 2005

8) TCLK 0x75 - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:54:32 2005

9) TCLK 0x75 - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:54:41 2005

10) TCLK 0x75 - (T) 0xa9becbfed61 - (B) 0xaf6197f7 - SUN JUL 03 23:54:48 2005

…

491) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:51:52 2005

492) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:51:57 2005

493) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:52:02 2005

494) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:52:07 2005

495) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:52:12 2005

496) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:52:17 2005

497) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:52:22 2005

498) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:52:27 2005

499) TCLK 0x75 - (T) 0x29a10d691c - (B) 0x75d0f0e7 - TUE JUL 05 05:53:21 2005

## Trace Diagnostics

The Main Injector BPM software packages are using a trace facility to monitor software behavior. This document is not meant to be a tutorial for that facility - Refer to the output of the traceHelp command (sample below) for an overview of the trace commands that are available.

fccts1:tbpm-dev> traceHelp

Trace Facility Help

traceMode(mode)          Sets the target(s) of the trace messages.

'mode' can be:

0 = tracing is disabled

1 = trace into circular buffer

2 = trace to VxWorks log facility

3 = trace to both

traceModeGet()           Returns the current trace mode.

traceInfo()              Shows information associated with the

trace facility. Each trace source is

also shown along with its level mask.

traceShow(opts, lines, skip, fd)

Dumps the contents of the trace buffer.

The arguments have the following effects:

'opts' is a bit mask:

bits 0,1:

0 = Relative timestamps are shown

1 = Absolute, but zeroed, timestamps

2 = Absolute timestamps

bit 2:

0 = do not show trace level

1 = show trace level

'lines' can be:

0 = Display all entries (default)

Otherwise only display the first

'lines' entries

'skip' can be:

0 = Don't skip entries

Otherwise display every 'skip'

entry

'fd' can be an open file descriptor. If

it is zero, then standard output is used.

traceReset()             Clears the circular trace buffer.

traceOn(id, lo, hi, f)      Turns on the specified levels of trace

source 'id'. This value can be found

using traceInfo(). 'lo' and 'hi' define

the range of levels to enable (0 - 31).

f is the "function" (currently -

0=circular, 1=logMsg.

traceOff(id, lo, hi, f)     Turns off the specified levels of trace

source 'id'. 'lo' and 'hi' define

the range of levels to disable (0 - 31).

f is same as above.

traceGlobalOn(lo, hi, f)    Like traceOn(), but it affects every

trace source.

f is same as above.

traceGlobalOff(lo, hi, f)   Like traceOff(), but it affects every

trace source.

f is same as above.

traceFreezeOn(lo, hi)    Enables a range of "freeze" levels. If

your code calls traceFreeze(lvl) and the

lvl'th freeze bit was set, the circular

trace queue is frozen (similar to calling

traceMode(0).) traceInfo() shows the

current freeze mask.

traceFreezeOff(lo, hi)   Disables a range of "freeze" levels.

As a quick overview, we typically use the tracing facility in a circular buffer mode where all traced processes write into that buffer. The software is exploiting many levels of trace that can be turned on/off to fine tune which information is stored in the buffer.

As an example, the following use of the traceShow command would display (on standard output) the most recent 100 entries in the trace buffer without skipping any entries, and each entry would be shown with an absolute timestamp and its trace level.

traceShow 6, 100, 0, 0

### Trace Levels

The GBPM library defines in the file Defs.h the trace levels used for the different parts of the system. The following table defines the levels used throughout the code.

| Level | Start | End |
| --- | --- | --- |
| ERROR_LEVEL | 0 | 0 |
| DRIVER_LEVEL | 1 | 8 |
| EVENT_LEVEL | 9 | 11 |
| READOUT_LEVEL | 12 | 15 |
| BUFFER_LEVEL | 16 | 17 |
| CONTROL_LEVEL | 18 | 18 |
| ACNET_LEVEL | 19 | 20 |
| ALARM_LEVEL | 21 | 21 |
| TIMING_LEVEL | 22 | 23 |

## VME Backplane Diagnostics

Bus analyzers are a powerful tool in understanding what is actually happening on the backplane. Using a Vanguard VME bus analyzer and the companion Bus Analyzer software version 4 the following traces where obtained to show VME timing of typical transfers:

| AbsTime | Transfer | AM/XAM | Address | Data | Size | Cycle | Status | Iack | IRQ[7:1]* |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Timing board interrupt (500 Hz rate) |
| 0.0ns | Single | A16nD | IACK=6 | ......A0 | LBYTE | Rd | Data | IACK6 | 1011111 |
| Echotek interrupt – First channel of first board is done |
| 1.358ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |
| Interrupt handler checks the status of all channels |
| 1.362ms | Single | A16nD | 8314 | FF000000 | D32 | Rd | Data | ----- | 111111 |
| 1.363ms | Single | A16nD | 8314 | 1000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.367ms | Single | A16nD | 8314 | FE000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.368ms | Single | A16nD | 8314 | FE000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.369ms | Single | A16nD | 9314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.370ms | Single | A16nD | 9314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.370ms | Single | A16nD | A314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.371ms | Single | A16nD | A314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.372ms | Single | A16nD | B314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.373ms | Single | A16nD | B314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.373ms | Single | A16nD | C314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.374ms | Single | A16nD | C314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.375ms | Single | A16nD | D314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.376ms | Single | A16nD | D314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.377ms | Single | A32nD | 18000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.378ms | Single | A32nD | 18000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.379ms | Single | A32nD | 19000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.380ms | Single | A32nD | 19000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.381ms | Single | A32nD | 1A000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.382ms | Single | A32nD | 1A000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.383ms | Single | A32nD | 1B000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.384ms | Single | A32nD | 1B000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.385ms | Single | A32nD | 1C000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.386ms | Single | A32nD | 1C000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.387ms | Single | A32nD | 1D000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.388ms | Single | A32nD | 1D000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| Board 0 – data read out (8 single transfers) |
| 1.389ms | Single | A32nD | 18100000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.391ms | Single | A32nD | 18180000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.392ms | Single | A32nD | 18200000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.394ms | Single | A32nD | 18280000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.396ms | Single | A32nD | 18300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.397ms | Single | A32nD | 18380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.399ms | Single | A32nD | 18400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.400ms | Single | A32nD | 18480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| Board 1 – data read out (8 single transfers) |
| 1.402ms | Single | A32nD | 19100000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.404ms | Single | A32nD | 19180000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.406ms | Single | A32nD | 19200000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.407ms | Single | A32nD | 19280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.409ms | Single | A32nD | 19300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.410ms | Single | A32nD | 19380000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.412ms | Single | A32nD | 19400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.414ms | Single | A32nD | 19480000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| Board 2 – data read out (8 single transfers) |
| 1.415ms | Single | A32nD | 1A100000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.417ms | Single | A32nD | 1A180000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.419ms | Single | A32nD | 1A200000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.420ms | Single | A32nD | 1A280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.422ms | Single | A32nD | 1A300000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.423ms | Single | A32nD | 1A380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.425ms | Single | A32nD | 1A400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.426ms | Single | A32nD | 1A480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| Board 3 – data read out (8 single transfers) |
| 1.428ms | Single | A32nD | 1B100000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.430ms | Single | A32nD | 1B180000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.432ms | Single | A32nD | 1B200000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.433ms | Single | A32nD | 1B280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.435ms | Single | A32nD | 1B300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.436ms | Single | A32nD | 1B380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.438ms | Single | A32nD | 1B400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.440ms | Single | A32nD | 1B480000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| Board 4 – data read out (8 single transfers) |
| 1.441ms | Single | A32nD | 1C100000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.443ms | Single | A32nD | 1C180000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.445ms | Single | A32nD | 1C200000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.446ms | Single | A32nD | 1C280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.448ms | Single | A32nD | 1C300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.449ms | Single | A32nD | 1C380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.451ms | Single | A32nD | 1C400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.452ms | Single | A32nD | 1C480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| Board 5 – data read out (8 single transfers) |
| 1.454ms | Single | A32nD | 1D100000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.456ms | Single | A32nD | 1D180000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.458ms | Single | A32nD | 1D200000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.459ms | Single | A32nD | 1D280000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.461ms | Single | A32nD | 1D300000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.462ms | Single | A32nD | 1D380000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.464ms | Single | A32nD | 1D400000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.466ms | Single | A32nD | 1D480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| Reading timing boad registers (timestamp and turns count) |
| 1.469ms | Single | A16nD | 203A | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.471ms | Single | A16nD | 203C | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.472ms | Single | A16nD | 203E | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.473ms | Single | A16nD | 2040 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.475ms | Single | A16nD | 2054 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.476ms | Single | A16nD | 2052 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| Board 0 arming (38 ms – Total of 35 VME read and write operations, first and last shown only) |
| 1.525ms | Single | A32nD | 18010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.563ms | Single | A32nD | 18000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| Board 1 arming (38 us) |
| 1.564ms | Single | A32nD | 19010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.602ms | Single | A32nD | 19000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| Board 2 arming (37 us) |
| 1.603ms | Single | A32nD | 1A010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.640ms | Single | A32nD | 1A000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| Board 3 arming (38 us) |
| 1.641ms | Single | A32nD | 1B010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.679ms | Single | A32nD | 1B000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| Board 4 arming (37 us) |
| 1.680ms | Single | A32nD | 1C010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.717ms | Single | A32nD | 1C000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| Board 5 arming (38 us) |
| 1.718ms | Single | A32nD | 1D010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.756ms | Single | A32nD | 1D000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| Waiting for next timing board interrupt… |
| 2.012ms | Single | A16nD | IACK=6 | ......A0 | LBYTE | Rd | Data | IACK6 | 1011111 |
| 3.370ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |

| AbsTime | Transfer | AM/XAM | Address | Data | Size | Cycle | Status | Iack | IRQ[7:1]* |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Timing board interrupt – BSYNC 0x7C or 0xDA received |
| 0.0ns | Single | A16nD | IACK=6 | ......A1 | LBYTE | Rd | Data | IACK6 | 1111111 |
| Echotek interrupt – First channel of first board is done. 8192 turns sampled |
| 171.817ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |
| Interrupt handler checks the status of all channels |
| 171.822ms | Single | A16nD | 8314 | FF000000 | D32 | Rd | Data | ----- | 111111 |
| 171.824ms | Single | A16nD | 8314 | 1000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.833ms | Single | A16nD | 8314 | FE000000 | D32 | Rd | Data | ----- | 1111111 |
| 171.834ms | Single | A16nD | 8314 | FE000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.835ms | Single | A16nD | 9314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 171.835ms | Single | A16nD | 9314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.836ms | Single | A16nD | A314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 171.837ms | Single | A16nD | A314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.838ms | Single | A16nD | B314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 171.838ms | Single | A16nD | B314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.839ms | Single | A16nD | C314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 171.840ms | Single | A16nD | C314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.841ms | Single | A16nD | D314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 171.842ms | Single | A16nD | D314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 171.844ms | Single | A32nD | 18000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 171.845ms | Single | A32nD | 18000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 171.846ms | Single | A32nD | 19000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 171.847ms | Single | A32nD | 19000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 171.848ms | Single | A32nD | 1A000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 171.850ms | Single | A32nD | 1A000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 171.851ms | Single | A32nD | 1B000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 171.852ms | Single | A32nD | 1B000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 171.853ms | Single | A32nD | 1C000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 171.854ms | Single | A32nD | 1C000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 171.856ms | Single | A32nD | 1D000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 171.857ms | Single | A32nD | 1D000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 171.859ms | Single | A32nD | 18000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 171.859ms | Single | A32nD | 18000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 171.860ms | Single | A16nD | 8220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 171.860ms | Single | A16nD | 8218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 171.861ms | Single | A16nD | 8204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 171.861ms | Single | A16nD | 8220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 171.864ms | Start | A32sMBLT | 25079800 FFF | E000100000001 | D64 | Wr | Data | ----- | 1111111 |
| 32702 VME operations omitted |
| 176.859ms | Start | A32sMBLT | 24980800 FFF | F0001FFFD0005 | D64 | Wr | Data | ----- | 1111111 |
| DMA ready interrupt for board 0 |
| 176.870ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |
| 176.874ms | Single | A16nD | 8314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 176.875ms | Single | A16nD | 8314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 176.970ms | Single | A32nD | 19000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 176.971ms | Single | A32nD | 19000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 176.971ms | Single | A16nD | 9220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 176.972ms | Single | A16nD | 9218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 176.972ms | Single | A16nD | 9204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 176.972ms | Single | A16nD | 9220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 176.976ms | Start | A32sMBLT | 24878800 000 | 0FFFF00010000 | D64 | Wr | Data | ----- | 1111111 |
| 181.963ms | Start | A32sMBLT | 2417F800 FFF | F000000010002 | D64 | Wr | Data | ----- | 1111111 |
| DMA ready interrupt for board 1 |
| 181.974ms | Single | A16nD | IACK=7 | ......E3 | LBYTE | Rd | Data | IACK7 | 111111 |
| 181.978ms | Single | A16nD | 9314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 181.979ms | Single | A16nD | 9314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 182.068ms | Single | A32nD | 1A000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 182.069ms | Single | A32nD | 1A000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 182.069ms | Single | A16nD | A220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 182.069ms | Single | A16nD | A218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 182.070ms | Single | A16nD | A204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 182.070ms | Single | A16nD | A220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 182.074ms | Start | A32sMBLT | 24077800 FFF | F000100000004 | D64 | Wr | Data | ----- | 1111111 |
| 32704 VME operations omitted |
| 187.064ms | Start | A32sMBLT | 2397E800 FFF | E0003FFFF0003 | D64 | Wr | Data | ----- | 1111111 |
| DMA ready interrupt for board 2 |
| 187.075ms | Single | A16nD | IACK=7 | ......E5 | LBYTE | Rd | Data | IACK7 | 111111 |
| 187.079ms | Single | A16nD | A314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 187.080ms | Single | A16nD | A314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 187.168ms | Single | A32nD | 1B000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 187.169ms | Single | A32nD | 1B000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 187.169ms | Single | A16nD | B220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 187.170ms | Single | A16nD | B218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 187.170ms | Single | A16nD | B204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 187.171ms | Single | A16nD | B220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 187.174ms | Start | A32sMBLT | 23876800 000 | 10001 | D64 | Wr | Data | ----- | 1111111 |
| 32704 VME operations omitted |
| 192.162ms | Start | A32sMBLT | 2317D800 000 | 00001FFFF0001 | D64 | Wr | Data | ----- | 1111111 |
| DMA ready interrupt for board 3 |
| 192.173ms | Single | A16nD | IACK=7 | ......E7 | LBYTE | Rd | Data | IACK7 | 111111 |
| 192.177ms | Single | A16nD | B314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 192.178ms | Single | A16nD | B314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 192.267ms | Single | A32nD | 1C000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 192.268ms | Single | A32nD | 1C000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 192.268ms | Single | A16nD | C220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 192.269ms | Single | A16nD | C218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 192.269ms | Single | A16nD | C204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 192.270ms | Single | A16nD | C220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 192.273ms | Start | A32sMBLT | 23075800 000 | 1FFFFFFFE0001 | D64 | Wr | Data | ----- | 1111111 |
| 32704 VME operations omitted |
| 197.349ms | Start | A32sMBLT | 2297C800 FFF | D0002FFFF0003 | D64 | Wr | Data | ----- | 1111111 |
| DMA ready interrupt for board 4 |
| 197.362ms | Single | A16nD | IACK=7 | ......E9 | LBYTE | Rd | Data | IACK7 | 111111 |
| 197.365ms | Single | A16nD | C314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 197.366ms | Single | A16nD | C314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 197.470ms | Single | A32nD | 1D000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 197.470ms | Single | A32nD | 1D000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 197.471ms | Single | A16nD | D220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 197.471ms | Single | A16nD | D218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 197.472ms | Single | A16nD | D204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 197.472ms | Single | A16nD | D220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 197.476ms | Start | A32sMBLT | 22874800 000 | 0FFFD00000000 | D64 | Wr | Data | ----- | 1111111 |
| 32704 VME operations omitted |
| 202.650ms | Start | A32sMBLT | 2217B800 000 | 0FFFE0002FFFE | D64 | Wr | Data | ----- | 1111111 |
| DMA ready interrupt for board 5 |
| 202.663ms | Single | A16nD | IACK=7 | ......EB | LBYTE | Rd | Data | IACK7 | 111111 |
| 202.667ms | Single | A16nD | D314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 202.668ms | Single | A16nD | D314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| Timing information from timing board |
| 202.762ms | Single | A16nD | 203A | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.763ms | Single | A16nD | 203C | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.765ms | Single | A16nD | 203E | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.766ms | Single | A16nD | 2040 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.769ms | Single | A16nD | 2044 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.771ms | Single | A16nD | 2042 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.773ms | Single | A16nD | 2048 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.774ms | Single | A16nD | 2046 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| Configure timing board to start taking closed orbits |
| 202.818ms | Single | A16nD | 2030 | ....1603 | WORD | Rd | Data | ----- | 1111111 |
| 202.819ms | Single | A16nD | 2030 | ....1000 | WORD | Wr | Data | ----- | 1111111 |
| 202.820ms | Single | A16nD | 2030 | ....1000 | WORD | Rd | Data | ----- | 1111111 |
| 202.821ms | Single | A16nD | 2030 | ....1040 | WORD | Wr | Data | ----- | 1111111 |
| 202.822ms | Single | A16nD | 2060 | ....09A1 | WORD | Rd | Data | ----- | 1111111 |
| 202.823ms | Single | A16nD | 2060 | ....09A1 | WORD | Wr | Data | ----- | 1111111 |
| 202.826ms | Single | A16nD | 2066 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 202.829ms | Single | A16nD | 2034 | ....00AA | WORD | Wr | Data | ----- | 1111111 |
| 202.830ms | Single | A16nD | 2036 | ....0010 | WORD | Wr | Data | ----- | 1111111 |
| 202.830ms | Single | A16nD | 2010 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.831ms | Single | A16nD | 2012 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.832ms | Single | A16nD | 2014 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.833ms | Single | A16nD | 2016 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.834ms | Single | A16nD | 2018 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.835ms | Single | A16nD | 201A | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.835ms | Single | A16nD | 201C | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.836ms | Single | A16nD | 201E | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 202.837ms | Single | A16nD | 2030 | ....0040 | WORD | Wr | Data | ----- | 1111111 |
| 202.838ms | Single | A16nD | 2026 | ....005F | WORD | Wr | Data | ----- | 1111111 |
| 202.839ms | Single | A16nD | 2030 | ....0040 | WORD | Rd | Data | ----- | 1111111 |
| 202.840ms | Single | A16nD | 2030 | ....00C0 | WORD | Wr | Data | ----- | 1111111 |
| 202.843ms | Single | A16nD | 2060 | ....01A0 | WORD | Wr | Data | ----- | 1111111 |
| 202.844ms | Single | A16nD | 2070 | ....0009 | WORD | Wr | Data | ----- | 1111111 |
| Load closed orbit configuration – board 0 (7216 VME transfers – 10 ms) |
| 205.776ms | Single | A32nD | 18000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| Load closed orbit configuration – board 1 (7216 VME transfers – 10 ms) |
| 215.294ms | Single | A32nD | 19000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| Load closed orbit configuration – board 2 (7216 VME transfers – 10 ms) |
| 224.810ms | Single | A32nD | 1A000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| Load closed orbit configuration – board 3 (7216 VME transfers – 10 ms) |
| 234.311ms | Single | A32nD | 1B000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| Load closed orbit configuration – board 4 (7216 VME transfers – 10 ms) |
| 243.808ms | Single | A32nD | 1C000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| Load closed orbit configuration – board 5 (7216 VME transfers – 10 ms) |
| 253.653ms | Single | A32nD | 1D000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| Board 0 arming (36 us) |
| 263.317ms | Single | A16nD | 8310 | 1010000 | D32 | Rd | Data | ----- | 1111111 |
| Board 1 arming (29 us) |
| 263.353ms | Single | A16nD | 9310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| Board 2 arming (28 us) |
| 263.382ms | Single | A16nD | A310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| Board 3 arming (29 us) |
| 263.410ms | Single | A16nD | B310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| Board 4 arming (28 us) |
| 263.439ms | Single | A16nD | C310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| Board 5 arming (30 us) |
| 263.467ms | Single | A16nD | D310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 263.497ms | Single | A32nD | 18000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 263.497ms | Single | A32nD | 19000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 263.498ms | Single | A32nD | 1A000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 263.499ms | Single | A32nD | 1B000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 263.499ms | Single | A32nD | 1C000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 263.500ms | Single | A32nD | 1D000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 263.501ms | Single | A32nD | 18000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 263.528ms | Single | A32nD | 19000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 263.553ms | Single | A32nD | 1A000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 263.577ms | Single | A32nD | 1B000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 263.601ms | Single | A32nD | 1C000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 263.626ms | Single | A32nD | 1D000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 263.658ms | Single | A32nD | 18010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 263.699ms | Single | A32nD | 19010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 263.737ms | Single | A32nD | 1A010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 263.776ms | Single | A32nD | 1B010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 263.814ms | Single | A32nD | 1C010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 263.853ms | Single | A32nD | 1D010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 263.905ms | Single | A16nD | 2060 | ....01A0 | WORD | Rd | Data | ----- | 1111111 |
| 263.906ms | Single | A16nD | 2060 | ....11A0 | WORD | Wr | Data | ----- | 1111111 |
| 263.916ms | Single | A16nD | 2030 | ....00C0 | WORD | Rd | Data | ----- | 1111111 |
| 263.917ms | Single | A16nD | 2030 | ....02C0 | WORD | Wr | Data | ----- | 1111111 |
| 263.918ms | Single | A16nD | 2062 | ....10C6 | WORD | Rd | Data | ----- | 1111111 |
| 263.919ms | Single | A16nD | 2062 | ....10C6 | WORD | Wr | Data | ----- | 1111111 |
| 263.921ms | Single | A16nD | 2030 | ....02C0 | WORD | Rd | Data | ----- | 1111111 |
| 263.922ms | Single | A16nD | 2030 | ....02C0 | WORD | Wr | Data | ----- | 1111111 |
| 263.923ms | Single | A16nD | 2062 | ....10C6 | WORD | Rd | Data | ----- | 1111111 |
| 263.924ms | Single | A16nD | 2062 | ....10C6 | WORD | Wr | Data | ----- | 1111111 |
| 263.926ms | Single | A16nD | 2030 | ....02C0 | WORD | Rd | Data | ----- | 1111111 |
| 263.928ms | Single | A16nD | 2030 | ....02C0 | WORD | Wr | Data | ----- | 1111111 |
| 263.929ms | Single | A16nD | 2062 | ....10C6 | WORD | Rd | Data | ----- | 1111111 |
| 263.930ms | Single | A16nD | 2062 | ....10C6 | WORD | Wr | Data | ----- | 1111111 |
| 263.931ms | Single | A16nD | 2030 | ....02C0 | WORD | Rd | Data | ----- | 1111111 |
| 263.932ms | Single | A16nD | 2030 | ....0280 | WORD | Wr | Data | ----- | 1111111 |
| 263.933ms | Single | A16nD | 2030 | ....0280 | WORD | Rd | Data | ----- | 1111111 |
| 263.934ms | Single | A16nD | 2030 | ....0683 | WORD | Wr | Data | ----- | 1111111 |
| Waiting for next closed orbit trigger |
| 265.294ms | Single | A16nD | IACK=6 | ......A0 | LBYTE | Rd | Data | IACK6 | 1011111 |
| 266.653ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |

## Looking at ACNET variables

From an ACNET console page, select the “ACL Edit/Run” option from the Pgm_Tools menu which is located near the top right corner of the page. In the Action section of the ACL Edit Window, type ‘read <variableName>’ and type <ctrl>-r to run the command. When the command is run, the ACNET window will split into two parts, the command source (which you typed) and the output. As an example, ‘read V:PING’ followed by <ctrl-r> should return a result like “V:PING =  5    ping” in the output pane.

Setting ACNET variables uses the same ACL Edit/Run application, but first you need to configure your ACNET console to allow variables to be set.  This is done from the Utilities window. Click on the Setting option in the Utilities window and set the value to one hour. Back on the ACL Edit/Run page, you can now set variable values using commands of the form ‘set <variableName>=<value>’.  For example, to set the value of V:PING, you could use a command like ‘set V:PING=5’ (followed by <ctrl>-r).

# Appendix: Node name/ACNET addresses

| node name | ip address | ACNET address | Location |
| --- | --- | --- | --- |
| mibpmts | 131.225.137.100 | 0x0C8D | FCC3 |
| ibpm10 | 131.225.137.80 | 0x0C94 | MI10 |
| ibpm20 | 131.225.137.81 | 0x0C95 | MI20 |
| ibpm30 | 131.225.137.78 | 0x0A0F | MI30 |
| ibpm40 | 131.225.137.70 | 0x0A0E | MI40 |
| ibpm50 | 131.225.137.82 | 0x0C96 | MI50 |
| ibpm6n | 131.225.137.84 | 0x0C98 | MI60 South |
| ibpm6s | 131.225.137.83 | 0x0C97 | Mi60 North |

# Appendix: Full VME trace

| AbsTime | Transfer | AM/XAM | Address | Data | Size | Cycle | Status | Iack | IRQ[7:1]* |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0.0ns | Single | A16nD | IACK=6 | ......A0 | LBYTE | Rd | Data | IACK6 | 1011111 |
| 1.358ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |
| 1.362ms | Single | A16nD | 8314 | FF000000 | D32 | Rd | Data | ----- | 111111 |
| 1.363ms | Single | A16nD | 8314 | 1000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.367ms | Single | A16nD | 8314 | FE000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.368ms | Single | A16nD | 8314 | FE000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.369ms | Single | A16nD | 9314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.370ms | Single | A16nD | 9314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.370ms | Single | A16nD | A314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.371ms | Single | A16nD | A314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.372ms | Single | A16nD | B314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.373ms | Single | A16nD | B314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.373ms | Single | A16nD | C314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.374ms | Single | A16nD | C314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.375ms | Single | A16nD | D314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 1.376ms | Single | A16nD | D314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 1.377ms | Single | A32nD | 18000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.378ms | Single | A32nD | 18000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.379ms | Single | A32nD | 19000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.380ms | Single | A32nD | 19000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.381ms | Single | A32nD | 1A000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.382ms | Single | A32nD | 1A000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.383ms | Single | A32nD | 1B000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.384ms | Single | A32nD | 1B000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.385ms | Single | A32nD | 1C000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.386ms | Single | A32nD | 1C000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.387ms | Single | A32nD | 1D000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 1.388ms | Single | A32nD | 1D000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 1.389ms | Single | A32nD | 18100000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.391ms | Single | A32nD | 18180000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.392ms | Single | A32nD | 18200000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.394ms | Single | A32nD | 18280000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.396ms | Single | A32nD | 18300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.397ms | Single | A32nD | 18380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.399ms | Single | A32nD | 18400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.400ms | Single | A32nD | 18480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.402ms | Single | A32nD | 19100000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.404ms | Single | A32nD | 19180000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.406ms | Single | A32nD | 19200000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.407ms | Single | A32nD | 19280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.409ms | Single | A32nD | 19300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.410ms | Single | A32nD | 19380000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.412ms | Single | A32nD | 19400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.414ms | Single | A32nD | 19480000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.415ms | Single | A32nD | 1A100000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.417ms | Single | A32nD | 1A180000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.419ms | Single | A32nD | 1A200000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.420ms | Single | A32nD | 1A280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.422ms | Single | A32nD | 1A300000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.423ms | Single | A32nD | 1A380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.425ms | Single | A32nD | 1A400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.426ms | Single | A32nD | 1A480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.428ms | Single | A32nD | 1B100000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.430ms | Single | A32nD | 1B180000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.432ms | Single | A32nD | 1B200000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.433ms | Single | A32nD | 1B280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.435ms | Single | A32nD | 1B300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.436ms | Single | A32nD | 1B380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.438ms | Single | A32nD | 1B400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.440ms | Single | A32nD | 1B480000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.441ms | Single | A32nD | 1C100000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.443ms | Single | A32nD | 1C180000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.445ms | Single | A32nD | 1C200000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.446ms | Single | A32nD | 1C280000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.448ms | Single | A32nD | 1C300000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.449ms | Single | A32nD | 1C380000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.451ms | Single | A32nD | 1C400000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.452ms | Single | A32nD | 1C480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.454ms | Single | A32nD | 1D100000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.456ms | Single | A32nD | 1D180000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.458ms | Single | A32nD | 1D200000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.459ms | Single | A32nD | 1D280000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.461ms | Single | A32nD | 1D300000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.462ms | Single | A32nD | 1D380000 | 1 | D32 | Rd | Data | ----- | 1111111 |
| 1.464ms | Single | A32nD | 1D400000 | 0000FFFF | D32 | Rd | Data | ----- | 1111111 |
| 1.466ms | Single | A32nD | 1D480000 | 0 | D32 | Rd | Data | ----- | 1111111 |
| 1.469ms | Single | A16nD | 203A | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.471ms | Single | A16nD | 203C | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.472ms | Single | A16nD | 203E | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.473ms | Single | A16nD | 2040 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.475ms | Single | A16nD | 2054 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.476ms | Single | A16nD | 2052 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 1.525ms | Single | A32nD | 18010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.526ms | Single | A32nD | 18010010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.528ms | Single | A32nD | 18011440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.529ms | Single | A32nD | 18011440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.530ms | Single | A32nD | 18011C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.532ms | Single | A32nD | 18011C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.533ms | Single | A32nD | 18010000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.534ms | Single | A32nD | 18010010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.534ms | Single | A32nD | 18020000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.535ms | Single | A32nD | 18020010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.536ms | Single | A32nD | 18021440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.538ms | Single | A32nD | 18021440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.539ms | Single | A32nD | 18021C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.541ms | Single | A32nD | 18021C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.542ms | Single | A32nD | 18020000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.542ms | Single | A32nD | 18020010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.543ms | Single | A32nD | 18030000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.544ms | Single | A32nD | 18030010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.545ms | Single | A32nD | 18031440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.547ms | Single | A32nD | 18031440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.548ms | Single | A32nD | 18031C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.550ms | Single | A32nD | 18031C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.550ms | Single | A32nD | 18030000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.551ms | Single | A32nD | 18030010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.552ms | Single | A32nD | 18040000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.553ms | Single | A32nD | 18040010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.554ms | Single | A32nD | 18041440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.556ms | Single | A32nD | 18041440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.557ms | Single | A32nD | 18041C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.559ms | Single | A32nD | 18041C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.559ms | Single | A32nD | 18040000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.560ms | Single | A32nD | 18040010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.561ms | Single | A32nD | 18000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.562ms | Single | A32nD | 18000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.563ms | Single | A32nD | 18000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| 1.564ms | Single | A32nD | 19010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.565ms | Single | A32nD | 19010010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.566ms | Single | A32nD | 19011440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.568ms | Single | A32nD | 19011440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.569ms | Single | A32nD | 19011C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.571ms | Single | A32nD | 19011C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.571ms | Single | A32nD | 19010000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.572ms | Single | A32nD | 19010010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.573ms | Single | A32nD | 19020000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.574ms | Single | A32nD | 19020010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.575ms | Single | A32nD | 19021440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.577ms | Single | A32nD | 19021440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.578ms | Single | A32nD | 19021C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.579ms | Single | A32nD | 19021C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.580ms | Single | A32nD | 19020000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.581ms | Single | A32nD | 19020010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.582ms | Single | A32nD | 19030000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.583ms | Single | A32nD | 19030010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.584ms | Single | A32nD | 19031440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.585ms | Single | A32nD | 19031440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.587ms | Single | A32nD | 19031C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.588ms | Single | A32nD | 19031C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.589ms | Single | A32nD | 19030000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.590ms | Single | A32nD | 19030010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.591ms | Single | A32nD | 19040000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.591ms | Single | A32nD | 19040010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.593ms | Single | A32nD | 19041440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.594ms | Single | A32nD | 19041440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.596ms | Single | A32nD | 19041C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.597ms | Single | A32nD | 19041C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.598ms | Single | A32nD | 19040000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.599ms | Single | A32nD | 19040010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.600ms | Single | A32nD | 19000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.601ms | Single | A32nD | 19000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.602ms | Single | A32nD | 19000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| 1.603ms | Single | A32nD | 1A010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.603ms | Single | A32nD | 1A010010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.605ms | Single | A32nD | 1A011440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.606ms | Single | A32nD | 1A011440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.607ms | Single | A32nD | 1A011C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.609ms | Single | A32nD | 1A011C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.610ms | Single | A32nD | 1A010000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.611ms | Single | A32nD | 1A010010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.611ms | Single | A32nD | 1A020000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.612ms | Single | A32nD | 1A020010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.614ms | Single | A32nD | 1A021440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.615ms | Single | A32nD | 1A021440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.616ms | Single | A32nD | 1A021C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.618ms | Single | A32nD | 1A021C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.619ms | Single | A32nD | 1A020000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.620ms | Single | A32nD | 1A020010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.620ms | Single | A32nD | 1A030000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.621ms | Single | A32nD | 1A030010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.622ms | Single | A32nD | 1A031440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.624ms | Single | A32nD | 1A031440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.625ms | Single | A32nD | 1A031C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.627ms | Single | A32nD | 1A031C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.628ms | Single | A32nD | 1A030000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.628ms | Single | A32nD | 1A030010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.629ms | Single | A32nD | 1A040000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.630ms | Single | A32nD | 1A040010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.631ms | Single | A32nD | 1A041440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.633ms | Single | A32nD | 1A041440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.634ms | Single | A32nD | 1A041C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.636ms | Single | A32nD | 1A041C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.636ms | Single | A32nD | 1A040000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.637ms | Single | A32nD | 1A040010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.638ms | Single | A32nD | 1A000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.639ms | Single | A32nD | 1A000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.640ms | Single | A32nD | 1A000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| 1.641ms | Single | A32nD | 1B010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.642ms | Single | A32nD | 1B010010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.643ms | Single | A32nD | 1B011440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.645ms | Single | A32nD | 1B011440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.646ms | Single | A32nD | 1B011C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.648ms | Single | A32nD | 1B011C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.648ms | Single | A32nD | 1B010000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.649ms | Single | A32nD | 1B010010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.650ms | Single | A32nD | 1B020000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.651ms | Single | A32nD | 1B020010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.652ms | Single | A32nD | 1B021440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.654ms | Single | A32nD | 1B021440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.655ms | Single | A32nD | 1B021C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.656ms | Single | A32nD | 1B021C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.657ms | Single | A32nD | 1B020000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.658ms | Single | A32nD | 1B020010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.659ms | Single | A32nD | 1B030000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.660ms | Single | A32nD | 1B030010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.661ms | Single | A32nD | 1B031440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.662ms | Single | A32nD | 1B031440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.664ms | Single | A32nD | 1B031C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.665ms | Single | A32nD | 1B031C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.666ms | Single | A32nD | 1B030000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.667ms | Single | A32nD | 1B030010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.668ms | Single | A32nD | 1B040000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.668ms | Single | A32nD | 1B040010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.670ms | Single | A32nD | 1B041440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.671ms | Single | A32nD | 1B041440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.672ms | Single | A32nD | 1B041C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.674ms | Single | A32nD | 1B041C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.675ms | Single | A32nD | 1B040000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.676ms | Single | A32nD | 1B040010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.676ms | Single | A32nD | 1B000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.678ms | Single | A32nD | 1B000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.679ms | Single | A32nD | 1B000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| 1.680ms | Single | A32nD | 1C010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.680ms | Single | A32nD | 1C010010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.682ms | Single | A32nD | 1C011440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.683ms | Single | A32nD | 1C011440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.684ms | Single | A32nD | 1C011C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.686ms | Single | A32nD | 1C011C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.687ms | Single | A32nD | 1C010000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.688ms | Single | A32nD | 1C010010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.688ms | Single | A32nD | 1C020000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.689ms | Single | A32nD | 1C020010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.690ms | Single | A32nD | 1C021440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.692ms | Single | A32nD | 1C021440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.693ms | Single | A32nD | 1C021C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.695ms | Single | A32nD | 1C021C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.696ms | Single | A32nD | 1C020000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.696ms | Single | A32nD | 1C020010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.697ms | Single | A32nD | 1C030000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.698ms | Single | A32nD | 1C030010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.699ms | Single | A32nD | 1C031440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.701ms | Single | A32nD | 1C031440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.702ms | Single | A32nD | 1C031C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.704ms | Single | A32nD | 1C031C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.704ms | Single | A32nD | 1C030000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.705ms | Single | A32nD | 1C030010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.706ms | Single | A32nD | 1C040000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.707ms | Single | A32nD | 1C040010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.708ms | Single | A32nD | 1C041440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.710ms | Single | A32nD | 1C041440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.711ms | Single | A32nD | 1C041C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.712ms | Single | A32nD | 1C041C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.713ms | Single | A32nD | 1C040000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.714ms | Single | A32nD | 1C040010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.715ms | Single | A32nD | 1C000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.716ms | Single | A32nD | 1C000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.717ms | Single | A32nD | 1C000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |
| 1.718ms | Single | A32nD | 1D010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.719ms | Single | A32nD | 1D010010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.720ms | Single | A32nD | 1D011440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.722ms | Single | A32nD | 1D011440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.723ms | Single | A32nD | 1D011C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.725ms | Single | A32nD | 1D011C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.725ms | Single | A32nD | 1D010000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.726ms | Single | A32nD | 1D010010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.727ms | Single | A32nD | 1D020000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.728ms | Single | A32nD | 1D020010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.729ms | Single | A32nD | 1D021440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.731ms | Single | A32nD | 1D021440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.732ms | Single | A32nD | 1D021C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.733ms | Single | A32nD | 1D021C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.734ms | Single | A32nD | 1D020000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.735ms | Single | A32nD | 1D020010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.736ms | Single | A32nD | 1D030000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.737ms | Single | A32nD | 1D030010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.738ms | Single | A32nD | 1D031440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.739ms | Single | A32nD | 1D031440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.741ms | Single | A32nD | 1D031C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.742ms | Single | A32nD | 1D031C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.743ms | Single | A32nD | 1D030000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.744ms | Single | A32nD | 1D030010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.745ms | Single | A32nD | 1D040000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.745ms | Single | A32nD | 1D040010 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 1.747ms | Single | A32nD | 1D041440 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.748ms | Single | A32nD | 1D041440 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.749ms | Single | A32nD | 1D041C40 | FF001E08 | D32 | Rd | Data | ----- | 1111111 |
| 1.751ms | Single | A32nD | 1D041C40 | FF001E08 | D32 | Wr | Data | ----- | 1111111 |
| 1.752ms | Single | A32nD | 1D040000 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.753ms | Single | A32nD | 1D040010 | 0E738F06 | D32 | Wr | Data | ----- | 1111111 |
| 1.753ms | Single | A32nD | 1D000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.755ms | Single | A32nD | 1D000000 | 010FF000 | D32 | Rd | Data | ----- | 1111111 |
| 1.756ms | Single | A32nD | 1D000000 | 010FF001 | D32 | Wr | Data | ----- | 1111111 |

| TRIG: | ---- | 0.0ns | Single | A16nD | IACK=6 | ......A1 | LBYTE | Rd | Data | IACK6 | 1111111 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | ---- | 171.817ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |
| 2 | ---- | 171.822ms | Single | A16nD | 8314 | FF000000 | D32 | Rd | Data | ----- | 111111 |
| 3 | ---- | 171.824ms | Single | A16nD | 8314 | 1000000 | D32 | Wr | Data | ----- | 1111111 |
| 4 | ---- | 171.833ms | Single | A16nD | 8314 | FE000000 | D32 | Rd | Data | ----- | 1111111 |
| 5 | ---- | 171.834ms | Single | A16nD | 8314 | FE000000 | D32 | Wr | Data | ----- | 1111111 |
| 6 | ---- | 171.835ms | Single | A16nD | 9314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 7 | ---- | 171.835ms | Single | A16nD | 9314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 8 | ---- | 171.836ms | Single | A16nD | A314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 9 | ---- | 171.837ms | Single | A16nD | A314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 10 | ---- | 171.838ms | Single | A16nD | B314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 11 | ---- | 171.838ms | Single | A16nD | B314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 12 | ---- | 171.839ms | Single | A16nD | C314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 13 | ---- | 171.840ms | Single | A16nD | C314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 14 | ---- | 171.841ms | Single | A16nD | D314 | FF000000 | D32 | Rd | Data | ----- | 1111111 |
| 15 | ---- | 171.842ms | Single | A16nD | D314 | FF000000 | D32 | Wr | Data | ----- | 1111111 |
| 16 | ---- | 171.844ms | Single | A32nD | 18000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 17 | ---- | 171.845ms | Single | A32nD | 18000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 18 | ---- | 171.846ms | Single | A32nD | 19000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 19 | ---- | 171.847ms | Single | A32nD | 19000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 20 | ---- | 171.848ms | Single | A32nD | 1A000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 21 | ---- | 171.850ms | Single | A32nD | 1A000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 22 | ---- | 171.851ms | Single | A32nD | 1B000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 23 | ---- | 171.852ms | Single | A32nD | 1B000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 24 | ---- | 171.853ms | Single | A32nD | 1C000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 25 | ---- | 171.854ms | Single | A32nD | 1C000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 26 | ---- | 171.856ms | Single | A32nD | 1D000000 | 010FF001 | D32 | Rd | Data | ----- | 1111111 |
| 27 | ---- | 171.857ms | Single | A32nD | 1D000000 | 010FF000 | D32 | Wr | Data | ----- | 1111111 |
| 28 | ---- | 171.859ms | Single | A32nD | 18000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 29 | ---- | 171.859ms | Single | A32nD | 18000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 30 | ---- | 171.860ms | Single | A16nD | 8220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 31 | ---- | 171.860ms | Single | A16nD | 8218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 32 | ---- | 171.861ms | Single | A16nD | 8204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 33 | ---- | 171.861ms | Single | A16nD | 8220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 34 | 3--- | 171.864ms | Start | A32sMBLT | 25079800 FFF | E000100000001 | D64 | Wr | Data | ----- | 1111111 |
| 290 | 3--- | 171.903ms | Start | A32sMBLT | 2507A000 FFF | F00010000FFFF | D64 | Wr | Data | ----- | 1111111 |
| 32738 | 3--- | 176.859ms | Start | A32sMBLT | 24980800 FFF | F0001FFFD0005 | D64 | Wr | Data | ----- | 1111111 |
| 32802 | ---- | 176.870ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |
| 32803 | ---- | 176.874ms | Single | A16nD | 8314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 32804 | ---- | 176.875ms | Single | A16nD | 8314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 32805 | ---- | 176.970ms | Single | A32nD | 19000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 32806 | ---- | 176.971ms | Single | A32nD | 19000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 32807 | ---- | 176.971ms | Single | A16nD | 9220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 32808 | ---- | 176.972ms | Single | A16nD | 9218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 32809 | ---- | 176.972ms | Single | A16nD | 9204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 32810 | ---- | 176.972ms | Single | A16nD | 9220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 32811 | 3--- | 176.976ms | Start | A32sMBLT | 24878800 000 | 0FFFF00010000 | D64 | Wr | Data | ----- | 1111111 |
| 65515 | 3--- | 181.963ms | Start | A32sMBLT | 2417F800 FFF | F000000010002 | D64 | Wr | Data | ----- | 1111111 |
| 65579 | ---- | 181.974ms | Single | A16nD | IACK=7 | ......E3 | LBYTE | Rd | Data | IACK7 | 111111 |
| 65580 | ---- | 181.978ms | Single | A16nD | 9314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 65581 | ---- | 181.979ms | Single | A16nD | 9314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 65582 | ---- | 182.068ms | Single | A32nD | 1A000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 65583 | ---- | 182.069ms | Single | A32nD | 1A000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 65584 | ---- | 182.069ms | Single | A16nD | A220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 65585 | ---- | 182.069ms | Single | A16nD | A218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 65586 | ---- | 182.070ms | Single | A16nD | A204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 65587 | ---- | 182.070ms | Single | A16nD | A220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 65588 | 3--- | 182.074ms | Start | A32sMBLT | 24077800 FFF | F000100000004 | D64 | Wr | Data | ----- | 1111111 |
| 98292 | 3--- | 187.064ms | Start | A32sMBLT | 2397E800 FFF | E0003FFFF0003 | D64 | Wr | Data | ----- | 1111111 |
| 98356 | ---- | 187.075ms | Single | A16nD | IACK=7 | ......E5 | LBYTE | Rd | Data | IACK7 | 111111 |
| 98357 | ---- | 187.079ms | Single | A16nD | A314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 98358 | ---- | 187.080ms | Single | A16nD | A314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 98359 | ---- | 187.168ms | Single | A32nD | 1B000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 98360 | ---- | 187.169ms | Single | A32nD | 1B000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 98361 | ---- | 187.169ms | Single | A16nD | B220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 98362 | ---- | 187.170ms | Single | A16nD | B218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 98363 | ---- | 187.170ms | Single | A16nD | B204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 98364 | ---- | 187.171ms | Single | A16nD | B220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 98365 | 3--- | 187.174ms | Start | A32sMBLT | 23876800 000 | 10001 | D64 | Wr | Data | ----- | 1111111 |
| 131069 | 3--- | 192.162ms | Start | A32sMBLT | 2317D800 000 | 00001FFFF0001 | D64 | Wr | Data | ----- | 1111111 |
| 131133 | ---- | 192.173ms | Single | A16nD | IACK=7 | ......E7 | LBYTE | Rd | Data | IACK7 | 111111 |
| 131134 | ---- | 192.177ms | Single | A16nD | B314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 131135 | ---- | 192.178ms | Single | A16nD | B314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 131136 | ---- | 192.267ms | Single | A32nD | 1C000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 131137 | ---- | 192.268ms | Single | A32nD | 1C000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 131138 | ---- | 192.268ms | Single | A16nD | C220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 131139 | ---- | 192.269ms | Single | A16nD | C218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 131140 | ---- | 192.269ms | Single | A16nD | C204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 131141 | ---- | 192.270ms | Single | A16nD | C220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 131142 | 3--- | 192.273ms | Start | A32sMBLT | 23075800 000 | 1FFFFFFFE0001 | D64 | Wr | Data | ----- | 1111111 |
| 163846 | 3--- | 197.349ms | Start | A32sMBLT | 2297C800 FFF | D0002FFFF0003 | D64 | Wr | Data | ----- | 1111111 |
| 163910 | ---- | 197.362ms | Single | A16nD | IACK=7 | ......E9 | LBYTE | Rd | Data | IACK7 | 111111 |
| 163911 | ---- | 197.365ms | Single | A16nD | C314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 163912 | ---- | 197.366ms | Single | A16nD | C314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 163913 | ---- | 197.470ms | Single | A32nD | 1D000020 | 400 | D32 | Wr | Data | ----- | 1111111 |
| 163914 | ---- | 197.470ms | Single | A32nD | 1D000024 | 8800000C | D32 | Wr | Data | ----- | 1111111 |
| 163915 | ---- | 197.471ms | Single | A16nD | D220 | 007F0000 | D32 | Wr | Data | ----- | 1111111 |
| 163916 | ---- | 197.471ms | Single | A16nD | D218 | 40000 | D32 | Wr | Data | ----- | 1111111 |
| 163917 | ---- | 197.472ms | Single | A16nD | D204 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 163918 | ---- | 197.472ms | Single | A16nD | D220 | 0C000088 | D32 | Wr | Data | ----- | 1111111 |
| 163919 | 3--- | 197.476ms | Start | A32sMBLT | 22874800 000 | 0FFFD00000000 | D64 | Wr | Data | ----- | 1111111 |
| 196623 | 3--- | 202.650ms | Start | A32sMBLT | 2217B800 000 | 0FFFE0002FFFE | D64 | Wr | Data | ----- | 1111111 |
| 196687 | ---- | 202.663ms | Single | A16nD | IACK=7 | ......EB | LBYTE | Rd | Data | IACK7 | 111111 |
| 196688 | ---- | 202.667ms | Single | A16nD | D314 | 10000 | D32 | Rd | Data | ----- | 111111 |
| 196689 | ---- | 202.668ms | Single | A16nD | D314 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 196690 | ---- | 202.762ms | Single | A16nD | 203A | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196691 | ---- | 202.763ms | Single | A16nD | 203C | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196692 | ---- | 202.765ms | Single | A16nD | 203E | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196693 | ---- | 202.766ms | Single | A16nD | 2040 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196694 | ---- | 202.769ms | Single | A16nD | 2044 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196695 | ---- | 202.771ms | Single | A16nD | 2042 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196696 | ---- | 202.773ms | Single | A16nD | 2048 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196697 | ---- | 202.774ms | Single | A16nD | 2046 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196698 | ---- | 202.818ms | Single | A16nD | 2030 | ....1603 | WORD | Rd | Data | ----- | 1111111 |
| 196699 | ---- | 202.819ms | Single | A16nD | 2030 | ....1000 | WORD | Wr | Data | ----- | 1111111 |
| 196700 | ---- | 202.820ms | Single | A16nD | 2030 | ....1000 | WORD | Rd | Data | ----- | 1111111 |
| 196701 | ---- | 202.821ms | Single | A16nD | 2030 | ....1040 | WORD | Wr | Data | ----- | 1111111 |
| 196702 | ---- | 202.822ms | Single | A16nD | 2060 | ....09A1 | WORD | Rd | Data | ----- | 1111111 |
| 196703 | ---- | 202.823ms | Single | A16nD | 2060 | ....09A1 | WORD | Wr | Data | ----- | 1111111 |
| 196704 | ---- | 202.826ms | Single | A16nD | 2066 | ....0000 | WORD | Rd | Data | ----- | 1111111 |
| 196705 | ---- | 202.829ms | Single | A16nD | 2034 | ....00AA | WORD | Wr | Data | ----- | 1111111 |
| 196706 | ---- | 202.830ms | Single | A16nD | 2036 | ....0010 | WORD | Wr | Data | ----- | 1111111 |
| 196707 | ---- | 202.830ms | Single | A16nD | 2010 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196708 | ---- | 202.831ms | Single | A16nD | 2012 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196709 | ---- | 202.832ms | Single | A16nD | 2014 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196710 | ---- | 202.833ms | Single | A16nD | 2016 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196711 | ---- | 202.834ms | Single | A16nD | 2018 | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196712 | ---- | 202.835ms | Single | A16nD | 201A | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196713 | ---- | 202.835ms | Single | A16nD | 201C | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196714 | ---- | 202.836ms | Single | A16nD | 201E | ....0001 | WORD | Wr | Data | ----- | 1111111 |
| 196715 | ---- | 202.837ms | Single | A16nD | 2030 | ....0040 | WORD | Wr | Data | ----- | 1111111 |
| 196716 | ---- | 202.838ms | Single | A16nD | 2026 | ....005F | WORD | Wr | Data | ----- | 1111111 |
| 196717 | ---- | 202.839ms | Single | A16nD | 2030 | ....0040 | WORD | Rd | Data | ----- | 1111111 |
| 196718 | ---- | 202.840ms | Single | A16nD | 2030 | ....00C0 | WORD | Wr | Data | ----- | 1111111 |
| 196719 | ---- | 202.843ms | Single | A16nD | 2060 | ....01A0 | WORD | Wr | Data | ----- | 1111111 |
| 196720 | ---- | 202.844ms | Single | A16nD | 2070 | ....0009 | WORD | Wr | Data | ----- | 1111111 |
| 196721 | ---- | 205.776ms | Single | A32nD | 18000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 203937 | ---- | 215.294ms | Single | A32nD | 19000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 211153 | ---- | 224.810ms | Single | A32nD | 1A000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 218369 | ---- | 234.311ms | Single | A32nD | 1B000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 225585 | ---- | 243.808ms | Single | A32nD | 1C000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 232801 | ---- | 253.653ms | Single | A32nD | 1D000000 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240017 | ---- | 263.317ms | Single | A16nD | 8310 | 1010000 | D32 | Rd | Data | ----- | 1111111 |
| 240018 | ---- | 263.318ms | Single | A16nD | 8310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240019 | ---- | 263.319ms | Single | A16nD | 8314 | FFFFFFFF | D32 | Wr | Data | ----- | 1111111 |
| 240020 | ---- | 263.319ms | Single | A16nD | 8310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240021 | ---- | 263.320ms | Single | A16nD | 8310 | 1010000 | D32 | Wr | Data | ----- | 1111111 |
| 240022 | ---- | 263.321ms | Single | A32nD | 18000008 | 000000FF | D32 | Wr | Data | ----- | 1111111 |
| 240064 | ---- | 263.353ms | Single | A16nD | 9310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240065 | ---- | 263.354ms | Single | A16nD | 9310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240066 | ---- | 263.354ms | Single | A16nD | 9314 | FFFFFFFF | D32 | Wr | Data | ----- | 1111111 |
| 240067 | ---- | 263.355ms | Single | A16nD | 9310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240068 | ---- | 263.356ms | Single | A16nD | 9310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240069 | ---- | 263.356ms | Single | A32nD | 19000008 | 000000FF | D32 | Wr | Data | ----- | 1111111 |
| 240111 | ---- | 263.382ms | Single | A16nD | A310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240112 | ---- | 263.383ms | Single | A16nD | A310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240113 | ---- | 263.383ms | Single | A16nD | A314 | FFFFFFFF | D32 | Wr | Data | ----- | 1111111 |
| 240114 | ---- | 263.384ms | Single | A16nD | A310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240115 | ---- | 263.385ms | Single | A16nD | A310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240116 | ---- | 263.385ms | Single | A32nD | 1A000008 | 000000FF | D32 | Wr | Data | ----- | 1111111 |
| 240158 | ---- | 263.410ms | Single | A16nD | B310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240159 | ---- | 263.411ms | Single | A16nD | B310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240160 | ---- | 263.412ms | Single | A16nD | B314 | FFFFFFFF | D32 | Wr | Data | ----- | 1111111 |
| 240161 | ---- | 263.412ms | Single | A16nD | B310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240162 | ---- | 263.413ms | Single | A16nD | B310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240163 | ---- | 263.414ms | Single | A32nD | 1B000008 | 000000FF | D32 | Wr | Data | ----- | 1111111 |
| 240205 | ---- | 263.439ms | Single | A16nD | C310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240206 | ---- | 263.440ms | Single | A16nD | C310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240207 | ---- | 263.440ms | Single | A16nD | C314 | FFFFFFFF | D32 | Wr | Data | ----- | 1111111 |
| 240208 | ---- | 263.441ms | Single | A16nD | C310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240209 | ---- | 263.442ms | Single | A16nD | C310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240210 | ---- | 263.442ms | Single | A32nD | 1C000008 | 000000FF | D32 | Wr | Data | ----- | 1111111 |
| 240252 | ---- | 263.467ms | Single | A16nD | D310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240253 | ---- | 263.468ms | Single | A16nD | D310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240254 | ---- | 263.469ms | Single | A16nD | D314 | FFFFFFFF | D32 | Wr | Data | ----- | 1111111 |
| 240255 | ---- | 263.469ms | Single | A16nD | D310 | 10000 | D32 | Rd | Data | ----- | 1111111 |
| 240256 | ---- | 263.470ms | Single | A16nD | D310 | 10000 | D32 | Wr | Data | ----- | 1111111 |
| 240257 | ---- | 263.471ms | Single | A32nD | 1D000008 | 000000FF | D32 | Wr | Data | ----- | 1111111 |
| 240299 | ---- | 263.497ms | Single | A32nD | 18000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240300 | ---- | 263.497ms | Single | A32nD | 19000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240301 | ---- | 263.498ms | Single | A32nD | 1A000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240302 | ---- | 263.499ms | Single | A32nD | 1B000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240303 | ---- | 263.499ms | Single | A32nD | 1C000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240304 | ---- | 263.500ms | Single | A32nD | 1D000018 | 0 | D32 | Wr | Data | ----- | 1111111 |
| 240305 | ---- | 263.501ms | Single | A32nD | 18000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 240324 | ---- | 263.528ms | Single | A32nD | 19000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 240343 | ---- | 263.553ms | Single | A32nD | 1A000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 240362 | ---- | 263.577ms | Single | A32nD | 1B000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 240381 | ---- | 263.601ms | Single | A32nD | 1C000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 240400 | ---- | 263.626ms | Single | A32nD | 1D000004 | FFFF8100 | D32 | Rd | Data | ----- | 1111111 |
| 240419 | ---- | 263.658ms | Single | A32nD | 18010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 240454 | ---- | 263.699ms | Single | A32nD | 19010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 240489 | ---- | 263.737ms | Single | A32nD | 1A010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 240524 | ---- | 263.776ms | Single | A32nD | 1B010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 240559 | ---- | 263.814ms | Single | A32nD | 1C010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 240594 | ---- | 263.853ms | Single | A32nD | 1D010000 | 016B9006 | D32 | Wr | Data | ----- | 1111111 |
| 240629 | ---- | 263.905ms | Single | A16nD | 2060 | ....01A0 | WORD | Rd | Data | ----- | 1111111 |
| 240630 | ---- | 263.906ms | Single | A16nD | 2060 | ....11A0 | WORD | Wr | Data | ----- | 1111111 |
| 240631 | ---- | 263.916ms | Single | A16nD | 2030 | ....00C0 | WORD | Rd | Data | ----- | 1111111 |
| 240632 | ---- | 263.917ms | Single | A16nD | 2030 | ....02C0 | WORD | Wr | Data | ----- | 1111111 |
| 240633 | ---- | 263.918ms | Single | A16nD | 2062 | ....10C6 | WORD | Rd | Data | ----- | 1111111 |
| 240634 | ---- | 263.919ms | Single | A16nD | 2062 | ....10C6 | WORD | Wr | Data | ----- | 1111111 |
| 240635 | ---- | 263.921ms | Single | A16nD | 2030 | ....02C0 | WORD | Rd | Data | ----- | 1111111 |
| 240636 | ---- | 263.922ms | Single | A16nD | 2030 | ....02C0 | WORD | Wr | Data | ----- | 1111111 |
| 240637 | ---- | 263.923ms | Single | A16nD | 2062 | ....10C6 | WORD | Rd | Data | ----- | 1111111 |
| 240638 | ---- | 263.924ms | Single | A16nD | 2062 | ....10C6 | WORD | Wr | Data | ----- | 1111111 |
| 240639 | ---- | 263.926ms | Single | A16nD | 2030 | ....02C0 | WORD | Rd | Data | ----- | 1111111 |
| 240640 | ---- | 263.928ms | Single | A16nD | 2030 | ....02C0 | WORD | Wr | Data | ----- | 1111111 |
| 240641 | ---- | 263.929ms | Single | A16nD | 2062 | ....10C6 | WORD | Rd | Data | ----- | 1111111 |
| 240642 | ---- | 263.930ms | Single | A16nD | 2062 | ....10C6 | WORD | Wr | Data | ----- | 1111111 |
| 240643 | ---- | 263.931ms | Single | A16nD | 2030 | ....02C0 | WORD | Rd | Data | ----- | 1111111 |
| 240644 | ---- | 263.932ms | Single | A16nD | 2030 | ....0280 | WORD | Wr | Data | ----- | 1111111 |
| 240645 | ---- | 263.933ms | Single | A16nD | 2030 | ....0280 | WORD | Rd | Data | ----- | 1111111 |
| 240646 | ---- | 263.934ms | Single | A16nD | 2030 | ....0683 | WORD | Wr | Data | ----- | 1111111 |
| 240647 | ---- | 265.294ms | Single | A16nD | IACK=6 | ......A0 | LBYTE | Rd | Data | IACK6 | 1011111 |
| 240648 | ---- | 266.653ms | Single | A16nD | IACK=7 | ......E1 | LBYTE | Rd | Data | IACK7 | 111111 |