# Text boxes

1300 → 325

+5 db gain

+5 dbm nom.

1300/325 MHz

Frequency Converter

Module

5 ¼” high 19” rack chassis

Type N connectors (rear mounted)

-8 dbm nom.

( 0 dbm max. (

325 → 1300

+8 db gain

In Shielding Pit

- Cables are Andrews 3/8” Heliax or better with Type N connectors

- Nominal power levels for 100 watt incident power are green nom.

- Maximum useful operating power levels are ~ orange max. ~

- DO NOT EXCEED damage power levels are ( red max. (

- Attenuator adjustments may be required at commissioning

In Cryostat

+27 dbm nom. (TBM)

1.3 GHz LLRF Drive from VTS

+6 dbm nom.

OPHIR #5083-003

200 W Amplifier

+58 db gain

Cavity Transmitted Power Signal

Y

Ferrit-Quasar

2CCH32-1

Circulator

Meca 720-30-2 Coupler

-33 db coupling at 325 MHz

-3 db

50 dbm nom.

200 W load

Incident

+17 dbm nom.

3

2

1

Reflected

Input

Output

325 → 1300

+8 db gain

325 → 1300

+8 db gain

50 dbm nom.

HINS / VTS System Diagram

Bob Webber – December 4, 2007

Notes:

-6 db

-20 db

+7 dbm nom.

~ +10 dbm max. ~

( +17 dbm max. (

-10 db

Cavity Incident Power Signal

Cavity Reflected Power Signal

-10 db

+7 dbm nom.

~ +10 dbm max. ~

( +17 dbm max. (

-10 db

325 MHz LLRF Drive

-10 db

0 dbm nom.

~ +10 dbm max. ~

( +17 dbm max. (

Cavity Transmitted Power Signal to VTS +16 dbm nom.

Cavity Reflected  Power Signal to VTS +16 dbm nom.

Cavity Incident Power Signal to VTS +16 dbm nom.

VTS Switch Box

ZHL-42W

1 watt

+37 dB