Booster Bunch Length Monitor

November 22, 2004

Updated October 23, 2009

Craig Drennan

Table of Contents

I. Introduction	3

II. The Bunch Length Formula	3

III. The Bunch Length Analog Signal Measurement	4

IV. Injection Bunch Length Measurement Details	7

# Figures

Figure 1  Block diagram of the analog signal processing.	4

Figure 2  Plot of the response of the tunnel diode detector at four frequencies.	5

Figure 3  Extraction Bunch Length Measurement Assembly (not actually implemented).	6

Figure 4  Injection Bunch Length Measurement Assembly.	6

Figure 5  Data from October 2009 Calibration.	8

Figure 6  ACNET Plot after installing the new module in October 2009.	9

# I. Introduction

This note describes the bunch length measurements made in the Booster at injection and extraction.  Section II derives the formula for determining bunch length as a function of the signal power at a fundamental frequency and the third harmonic of this frequency.

Section III provides a description of the measurement is made in general terms.  Section IV gives more specific information on how the signals for the Injection bunch length are processed and details such as ACNET device names and calculations.

# II. The Bunch Length Formula

The protons in the beam are assumed to be in bunches with a Gaussian longitudinal distribution, separated by a fixed interval ().  At injection  is 200 MHz and at extraction it is 52 MHz.  The signal produced by such a train of bunches is expressed as

Eq. 1

where,

T. Leiri and G. Jackson describe the determination of the bunch length in Fermi TM-1600.  The Fourier transform of Eq. 1 is

Eq. 2

where  is the error function.  The beam spectrum is dominated by harmonics of the RF frequency, .  Therefore the magnitude of  is evaluated at those frequencies.  This becomes

Eq. 3

For the region of interest where   Eq. 3 reduces to

Eq. 4

Consider the ratio

Eq. 5

Solving for  we have,

Eq.6

The bunch length measurement in the Booster is made using tunnel diode detectors.  These devices produce an output voltage proportional to the power of a signal in dBm.  Using some math identities and some algebra we can derive

Eq.7

For

Eq.7

# III. The Bunch Length Analog Signal Measurement

As mentioned previously, the power of the beam pickup signal is measured using tunnel diode detectors, at the fundamental frequency of the beam signal and the third harmonic of the fundamental.  Figure 1 illustrates the analog signal processing that is performed.

Figure 1  Block diagram of the analog signal processing.

As illustrated, the signal is split and the power of the two frequency components,  and , are measured by the tunnel diode detectors.  The transfer function of the detectors is shown in Figure 2.

Figure 2  Plot of the response of the tunnel diode detector at four frequencies.

The curves in Figure 2 are those for the four frequencies of interest in the Booster.  Frequencies 200 MHz and 600 MHz are  and  for the bunch length measurement at injection.  Frequencies 52 MHz and 156 MHz are  and  for the bunch length measurement at extraction.  We can also see that the Volts output of the detectors is most linear with respect to power between -15 dBm and –5 dBm.  A line fit through this range gives that

Eq. 8

This simple fit may not provide sufficient precision, but it allows us to begin considering the range of the measurement.

At injection : 		Eq. 9

At extraction : 		Eq. 10

For a maximum , the largest bunch length we expect to measure is

At injection :

At extraction :

Figure 3  Extraction Bunch Length Measurement Assembly (not actually implemented).

Figure 4  Injection Bunch Length Measurement Assembly.

# IV. Injection Bunch Length Measurement Details

The Injection bunch length is currently measured from an RF signal generated by the “Griffin” detector in the 400 MeV beam line (Booster side) connecting the Linac to the Booster.  The connection to this RF signal is in Rack G24-RR6-3 at a BNC feedthru labeled “Booster Griffin Cable #4 Output”.

Note: This Griffin detector output cable cannot be left un terminated.

The power diode circuit is in a NIM module in Rack G24-RR6-2.  The power diode output voltages are digitized by a Comet VME, 5 MHz digitizer module like those used to digitize BPM position voltages.  Digitization is triggered by a signal from the Linac chopper, “400 MeV Gate”.

The ACNET names associated with this measurement are

B:INJBL1 Voltage between 0 and -0.2 Volts, proportional to the power in the fundamental (~200 MHz).  This isdigitized by a 10 MHz Comet Digitizer (Sharon Lackey).

B:INJBL2 Voltage between 0 and -0.2 Volts, proportional to the power in the 3rd Harmonic (~600 MHz).  This isdigitized by a 10 MHz Comet Digitizer (Sharon Lackey).

B:AVEXT1 Average voltage over a beam pulse between 0 and-0.2 Volts, proportional to the power in the fundamental (~200 MHz).

B:AVEXT2 Average voltage over a beam pulse between 0 and-0.2 Volts, proportional to the power in the 3rd Harmonic (~600 MHz).

B:BL400M This is a calculated value derived from an ACNET MACALC OAC process that employs the measured voltage averages B:AVEXT1 and B:AVEXT2.  Dennis Nicklaus from AD/Controls (x6410) manages the OAC processes.

Eq. 11

Measurements on the power diodes in use are shown below.

Figure 5  Data from October 2009 Calibration.

Figure 6  ACNET Plot after installing the new module in October 2009.