## BLM SCALING TO RADS
Computation of Dynamic Range and Conversion Results
| Log Integ. Calibration Coefficients |  | ACNET Log to Linear conv. coefficients |  | New Integrator Gain Term |  | Coefficients for Equivalent Match |  | New Integrator Gain Term |  | Conversion to Rads (70nC/Rad) |  | Numerator/Denominator for Integer Math
| c1 = | 1.079 | d1 = | 0.007212 | G1 = | 15817100000000 | m = | 3535.6672 | G1 = | 15817100000000 | H1 = 1/(G1_*70e-9) = | 9.0318e-07 | N = | 15
| c2 = | 23.344 | d2 = | 1.0578 |  |  | b = | -30962.7319 |  |  |  |  | D = | 4096
|  | Log Integrator Conversions |  |  |  | New Integrator Conversions (Equivalent) |  |  |  | New Integrator Conversions (To Rads)
| Q=EXP((V/c1)-(c2/c1)) | V=c1*LN(Q)+c2 | Bits=V*3276.8 | RS=d1*EXP(d2*V) | Copied | S(k) = (1/G1)*Qsum | Y=m*LN(S)+b | RS=d1*EXP(d2*Y/3276.8) | Copied | S(k) = (G1)*Qsum | Rads=S*H1 | RS=d1_*EXP(d2_*Y/3276.8)
given | compute | compute | compute | compute | from column 'C' | compute | compute | compute | from column 'C' | compute | compute | compute
Volts | Qsum | Volts | Bits | Rad/Sec | Qsum | S(k), Bits | equivalent Log | Rad/Sec | Qsum | S(k), Bits | Rads | Rad/Sec | =S(k)*N (Hex) | =S(k)*N/D (Hex) | =S(k)*N/D (Dec.) | Rads = S*N/D/4000 | %error | Qsum | Rad/Sec per bit | Rads per bit
0.0002 | 4.0196e-10 | 0.0002 | 0.65536 | 0.0072135 | 4.0196e-10 | 6357 | 0 | 0.007212 | 4.0196e-10 | 6357 | 0.0057415 | 0.007212 | 0001747B | 00000017 | 23 | 0.00575 | -0.14749 | 4.0196e-10 | 0.011007 | 0.00025
0.1 | 4.4092e-10 | 0.1 | 327.68 | 0.0080166 | 4.4092e-10 | 6973 | 327 | 0.0080149 | 4.4092e-10 | 6973 | 0.0062979 | 0.0080149 | 00019893 | 00000019 | 25 | 0.00625 | 0.76325 | 4.4092e-10 | 2.4465e-05 | 0.00025
0.2 | 4.8373e-10 | 0.2 | 655.36 | 0.0089111 | 4.8373e-10 | 7651 | 655 | 0.00891 | 4.8373e-10 | 7651 | 0.0069102 | 0.00891 | 0001C04D | 0000001C | 28 | 0.007 | -1.2905 | 4.8373e-10 | 1.3597e-05 | 0.00025
0.3 | 5.3071e-10 | 0.3 | 983.04 | 0.0099053 | 5.3071e-10 | 8394 | 982 | 0.009902 | 5.3071e-10 | 8394 | 0.0075813 | 0.009902 | 0001EBD6 | 0000001E | 30 | 0.0075 | 1.0782 | 5.3071e-10 | 1.0076e-05 | 0.00025
0.4 | 5.8224e-10 | 0.4 | 1310.72 | 0.011011 | 5.8224e-10 | 9209 | 1310 | 0.011008 | 5.8224e-10 | 9209 | 0.0083174 | 0.011008 | 00021B97 | 00000021 | 33 | 0.00825 | 0.81364 | 5.8224e-10 | 8.4003e-06 | 0.00025
0.5 | 6.3878e-10 | 0.5 | 1638.4 | 0.012239 | 6.3878e-10 | 10103 | 1638 | 0.012237 | 6.3878e-10 | 10103 | 0.0091248 | 0.012237 | 00024FF9 | 00000024 | 36 | 0.009 | 1.3776 | 6.3878e-10 | 7.4701e-06 | 0.00025
0.6 | 7.0082e-10 | 0.6 | 1966.08 | 0.013605 | 7.0082e-10 | 11084 | 1965 | 0.0136 | 7.0082e-10 | 11084 | 0.010011 | 0.0136 | 00028974 | 00000028 | 40 | 0.01 | 0.10859 | 7.0082e-10 | 6.9196e-06 | 0.00025
0.7 | 7.6887e-10 | 0.7 | 2293.76 | 0.015122 | 7.6887e-10 | 12161 | 2293 | 0.015119 | 7.6887e-10 | 12161 | 0.010984 | 0.015119 | 0002C88F | 0000002C | 44 | 0.011 | -0.14928 | 7.6887e-10 | 6.5929e-06 | 0.00025
0.8 | 8.4353e-10 | 0.8 | 2621.44 | 0.01681 | 8.4353e-10 | 13342 | 2621 | 0.016807 | 8.4353e-10 | 13342 | 0.01205 | 0.016807 | 00030DC2 | 00000030 | 48 | 0.012 | 0.41787 | 8.4353e-10 | 6.4124e-06 | 0.00025
0.9 | 9.2545e-10 | 0.9 | 2949.12 | 0.018685 | 9.2545e-10 | 14637 | 2948 | 0.018679 | 9.2545e-10 | 14637 | 0.01322 | 0.018679 | 000359A3 | 00000035 | 53 | 0.01325 | -0.22766 | 9.2545e-10 | 6.3359e-06 | 0.00025
1 | 1.0153e-09 | 1 | 3276.8 | 0.02077 | 1.0153e-09 | 16059 | 3276 | 0.020765 | 1.0153e-09 | 16059 | 0.014504 | 0.020765 | 0003ACF5 | 0000003A | 58 | 0.0145 | 0.028919 | 1.0153e-09 | 6.3385e-06 | 0.00025
1.1 | 1.1139e-09 | 1.1 | 3604.48 | 0.023087 | 1.1139e-09 | 17618 | 3604 | 0.023084 | 1.1139e-09 | 17618 | 0.015912 | 0.023084 | 0004084E | 00000040 | 64 | 0.016 | -0.54992 | 1.1139e-09 | 6.4052e-06 | 0.00025
1.2 | 1.2221e-09 | 1.2 | 3932.16 | 0.025663 | 1.2221e-09 | 19329 | 3932 | 0.025662 | 1.2221e-09 | 19329 | 0.017458 | 0.025662 | 00046C8F | 00000046 | 70 | 0.0175 | -0.24259 | 1.2221e-09 | 6.5265e-06 | 0.00025
1.3 | 1.3408e-09 | 1.3 | 4259.84 | 0.028527 | 1.3408e-09 | 21206 | 4259 | 0.028519 | 1.3408e-09 | 21206 | 0.019153 | 0.028519 | 0004DA8A | 0000004D | 77 | 0.01925 | -0.50585 | 1.3408e-09 | 6.6967e-06 | 0.00025
1.4 | 1.471e-09 | 1.4 | 4587.52 | 0.03171 | 1.471e-09 | 23266 | 4587 | 0.031704 | 1.471e-09 | 23266 | 0.021013 | 0.031704 | 0005533E | 00000055 | 85 | 0.02125 | -1.1195 | 1.471e-09 | 6.9122e-06 | 0.00025
1.5 | 1.6138e-09 | 1.5 | 4915.2 | 0.035248 | 1.6138e-09 | 25525 | 4915 | 0.035245 | 1.6138e-09 | 25525 | 0.023054 | 0.035245 | 0005D79B | 0000005D | 93 | 0.02325 | -0.84783 | 1.6138e-09 | 7.1712e-06 | 0.00025
1.6 | 1.7705e-09 | 1.6 | 5242.88 | 0.03918 | 1.7705e-09 | 28004 | 5242 | 0.039169 | 1.7705e-09 | 28004 | 0.025293 | 0.039169 | 000668DC | 00000066 | 102 | 0.0255 | -0.81626 | 1.7705e-09 | 7.4731e-06 | 0.00025
1.7 | 1.9425e-09 | 1.7 | 5570.56 | 0.043552 | 1.9425e-09 | 30723 | 5570 | 0.043544 | 1.9425e-09 | 30723 | 0.027748 | 0.043544 | 0007082D | 00000070 | 112 | 0.028 | -0.90245 | 1.9425e-09 | 7.8182e-06 | 0.00025
1.8 | 2.1311e-09 | 1.8 | 5898.24 | 0.048411 | 2.1311e-09 | 33707 | 5898 | 0.048407 | 2.1311e-09 | 33707 | 0.030444 | 0.048407 | 0007B705 | 0000007B | 123 | 0.03075 | -1.0016 | 2.1311e-09 | 8.2077e-06 | 0.00025
1.9 | 2.338e-09 | 1.9 | 6225.92 | 0.053813 | 2.338e-09 | 36980 | 6225 | 0.053797 | 2.338e-09 | 36980 | 0.0334 | 0.053797 | 000876CC | 00000087 | 135 | 0.03375 | -1.0435 | 2.338e-09 | 8.6433e-06 | 0.00025
2 | 2.5651e-09 | 2 | 6553.6 | 0.059817 | 2.5651e-09 | 40572 | 6553 | 0.059805 | 2.5651e-09 | 40572 | 0.036644 | 0.059805 | 00094944 | 00000094 | 148 | 0.037 | -0.96713 | 2.5651e-09 | 9.1273e-06 | 0.00025
2.1 | 2.8142e-09 | 2.1 | 6881.28 | 0.066491 | 2.8142e-09 | 44511 | 6881 | 0.066485 | 2.8142e-09 | 44511 | 0.040202 | 0.066485 | 000A3011 | 000000A3 | 163 | 0.04075 | -1.3551 | 2.8142e-09 | 9.6625e-06 | 0.00025
2.2 | 3.0874e-09 | 2.2 | 7208.96 | 0.073909 | 3.0874e-09 | 48834 | 7208 | 0.073886 | 3.0874e-09 | 48834 | 0.044106 | 0.073886 | 000B2D5E | 000000B2 | 178 | 0.0445 | -0.88939 | 3.0874e-09 | 1.0252e-05 | 0.00025
2.3 | 3.3873e-09 | 2.3 | 7536.64 | 0.082156 | 3.3873e-09 | 53576 | 7536 | 0.082139 | 3.3873e-09 | 53576 | 0.048389 | 0.082139 | 000C4338 | 000000C4 | 196 | 0.049 | -1.2551 | 3.3873e-09 | 1.0901e-05 | 0.00025
2.4 | 3.7162e-09 | 2.4 | 7864.32 | 0.091322 | 3.7162e-09 | 58779 | 7864 | 0.091313 | 3.7162e-09 | 58779 | 0.053088 | 0.091313 | 000D7415 | 000000D7 | 215 | 0.05375 | -1.239 | 3.7162e-09 | 1.1612e-05 | 0.00025
2.5 | 4.0771e-09 | 2.5 | 8192 | 0.10151 | 4.0771e-09 | 64487 | 8192 | 0.10151 | 4.0771e-09 | 64487 | 0.058243 | 0.10151 | 000EC289 | 000000EC | 236 | 0.059 | -1.2905 | 4.0771e-09 | 1.2392e-05 | 0.00025
2.6 | 4.473e-09 | 2.6 | 8519.68 | 0.11284 | 4.473e-09 | 70749 | 8519 | 0.11281 | 4.473e-09 | 70749 | 0.063899 | 0.11281 | 00103173 | 00000103 | 259 | 0.06475 | -1.3227 | 4.473e-09 | 1.3244e-05 | 0.00025
2.7 | 4.9074e-09 | 2.7 | 8847.36 | 0.12543 | 4.9074e-09 | 77620 | 8847 | 0.12541 | 4.9074e-09 | 77620 | 0.070105 | 0.12541 | 0011C40C | 0000011C | 284 | 0.071 | -1.2686 | 4.9074e-09 | 1.4177e-05 | 0.00025
2.8 | 5.3839e-09 | 2.8 | 9175.04 | 0.13942 | 5.3839e-09 | 85157 | 9175 | 0.13942 | 5.3839e-09 | 85157 | 0.076912 | 0.13942 | 00137DAB | 00000137 | 311 | 0.07775 | -1.0833 | 5.3839e-09 | 1.5196e-05 | 0.00025
2.9 | 5.9067e-09 | 2.9 | 9502.72 | 0.15498 | 5.9067e-09 | 93427 | 9502 | 0.15494 | 5.9067e-09 | 93427 | 0.084382 | 0.15494 | 0015623D | 00000156 | 342 | 0.0855 | -1.3167 | 5.9067e-09 | 1.6309e-05 | 0.00025
3 | 6.4803e-09 | 3 | 9830.4 | 0.17227 | 6.4803e-09 | 102499 | 9830 | 0.17225 | 6.4803e-09 | 102499 | 0.092575 | 0.17225 | 001775CD | 00000177 | 375 | 0.09375 | -1.261 | 6.4803e-09 | 1.7524e-05 | 0.00025
3.1 | 7.1096e-09 | 3.1 | 10158.08 | 0.19149 | 7.1096e-09 | 112453 | 10158 | 0.19148 | 7.1096e-09 | 112453 | 0.10157 | 0.19148 | 0019BD0B | 0000019B | 411 | 0.10275 | -1.1595 | 7.1096e-09 | 1.8851e-05 | 0.00025
3.2 | 7.8e-09 | 3.2 | 10485.76 | 0.21285 | 7.8e-09 | 123373 | 10485 | 0.2128 | 7.8e-09 | 123373 | 0.11143 | 0.2128 | 001C3CE3 | 000001C3 | 451 | 0.11275 | -1.1792 | 7.8e-09 | 2.0299e-05 | 0.00025
3.3 | 8.5575e-09 | 3.3 | 10813.44 | 0.2366 | 8.5575e-09 | 135354 | 10813 | 0.23657 | 8.5575e-09 | 135354 | 0.12225 | 0.23657 | 001EFAE6 | 000001EF | 495 | 0.12375 | -1.2201 | 8.5575e-09 | 2.1881e-05 | 0.00025
3.4 | 9.3885e-09 | 3.4 | 11141.12 | 0.263 | 9.3885e-09 | 148498 | 11141 | 0.26299 | 9.3885e-09 | 148498 | 0.13412 | 0.26299 | 0021FD0E | 0000021F | 543 | 0.13575 | -1.2075 | 9.3885e-09 | 2.3607e-05 | 0.00025
3.5 | 1.03e-08 | 3.5 | 11468.8 | 0.29235 | 1.03e-08 | 162919 | 11468 | 0.29227 | 1.03e-08 | 162919 | 0.14715 | 0.29227 | 00254A09 | 00000254 | 596 | 0.149 | -1.2525 | 1.03e-08 | 2.5491e-05 | 0.00025
3.6 | 1.13e-08 | 3.6 | 11796.48 | 0.32497 | 1.13e-08 | 178740 | 11796 | 0.32492 | 1.13e-08 | 178740 | 0.16143 | 0.32492 | 0028E90C | 0000028E | 654 | 0.1635 | -1.2712 | 1.13e-08 | 2.7548e-05 | 0.00025
3.7 | 1.2398e-08 | 3.7 | 12124.16 | 0.36122 | 1.2398e-08 | 196097 | 12124 | 0.3612 | 1.2398e-08 | 196097 | 0.17711 | 0.3612 | 002CE20F | 000002CE | 718 | 0.1795 | -1.3397 | 1.2398e-08 | 2.9794e-05 | 0.00025
3.8 | 1.3602e-08 | 3.8 | 12451.84 | 0.40153 | 1.3602e-08 | 215140 | 12451 | 0.40142 | 1.3602e-08 | 215140 | 0.19431 | 0.40142 | 00313DDC | 00000313 | 787 | 0.19675 | -1.2476 | 1.3602e-08 | 3.2246e-05 | 0.00025
3.9 | 1.4923e-08 | 3.9 | 12779.52 | 0.44633 | 1.4923e-08 | 236032 | 12779 | 0.44625 | 1.4923e-08 | 236032 | 0.21318 | 0.44625 | 00360600 | 00000360 | 864 | 0.216 | -1.3142 | 1.4923e-08 | 3.4925e-05 | 0.00025
4 | 1.6372e-08 | 4 | 13107.2 | 0.49613 | 1.6372e-08 | 258953 | 13107 | 0.49609 | 1.6372e-08 | 258953 | 0.23388 | 0.49609 | 003B4507 | 000003B4 | 948 | 0.237 | -1.3245 | 1.6372e-08 | 3.7851e-05 | 0.00025
4.1 | 1.7962e-08 | 4.1 | 13434.88 | 0.55148 | 1.7962e-08 | 284099 | 13434 | 0.55132 | 1.7962e-08 | 284099 | 0.25659 | 0.55132 | 0041066D | 00000410 | 1040 | 0.26 | -1.319 | 1.7962e-08 | 4.1048e-05 | 0.00025
4.2 | 1.9706e-08 | 4.2 | 13762.56 | 0.61301 | 1.9706e-08 | 311688 | 13762 | 0.6129 | 1.9706e-08 | 311688 | 0.28151 | 0.6129 | 004756F8 | 00000475 | 1141 | 0.28525 | -1.3195 | 1.9706e-08 | 4.4542e-05 | 0.00025
4.3 | 2.1619e-08 | 4.3 | 14090.24 | 0.68141 | 2.1619e-08 | 341955 | 14090 | 0.68135 | 2.1619e-08 | 341955 | 0.30885 | 0.68135 | 004E446D | 000004E4 | 1252 | 0.313 | -1.3355 | 2.1619e-08 | 4.836e-05 | 0.00025
4.4 | 2.3719e-08 | 4.4 | 14417.92 | 0.75743 | 2.3719e-08 | 375162 | 14417 | 0.75721 | 2.3719e-08 | 375162 | 0.33884 | 0.75721 | 0055DE26 | 0000055D | 1373 | 0.34325 | -1.2933 | 2.3719e-08 | 5.2534e-05 | 0.00025
4.5 | 2.6022e-08 | 4.5 | 14745.6 | 0.84195 | 2.6022e-08 | 411594 | 14745 | 0.84178 | 2.6022e-08 | 411594 | 0.37174 | 0.84178 | 005E34D6 | 000005E3 | 1507 | 0.37675 | -1.3376 | 2.6022e-08 | 5.7098e-05 | 0.00025
4.6 | 2.8549e-08 | 4.6 | 15073.28 | 0.93588 | 2.8549e-08 | 451563 | 15073 | 0.9358 | 2.8549e-08 | 451563 | 0.40784 | 0.9358 | 00675AC5 | 00000675 | 1653 | 0.41325 | -1.3169 | 2.8549e-08 | 6.2089e-05 | 0.00025
4.7 | 3.1321e-08 | 4.7 | 15400.96 | 1.0403 | 3.1321e-08 | 495414 | 15400 | 1.04 | 3.1321e-08 | 495414 | 0.44745 | 1.04 | 0071642A | 00000716 | 1814 | 0.4535 | -1.3433 | 3.1321e-08 | 6.7548e-05 | 0.00025
4.8 | 3.4363e-08 | 4.8 | 15728.64 | 1.1564 | 3.4363e-08 | 543524 | 15728 | 1.1561 | 3.4363e-08 | 543524 | 0.4909 | 1.1561 | 007C671C | 000007C6 | 1990 | 0.4975 | -1.3353 | 3.4363e-08 | 7.352e-05 | 0.00025
4.9 | 3.77e-08 | 4.9 | 16056.32 | 1.2854 | 3.77e-08 | 596304 | 16056 | 1.2853 | 3.77e-08 | 596304 | 0.53857 | 1.2853 | 00887BB0 | 00000887 | 2183 | 0.54575 | -1.3242 | 3.77e-08 | 8.0056e-05 | 0.00025
5 | 4.1361e-08 | 5 | 16384 | 1.4288 | 4.1361e-08 | 654211 | 16384 | 1.4288 | 4.1361e-08 | 654211 | 0.59087 | 1.4288 | 0095BCAD | 0000095B | 2395 | 0.59875 | -1.3246 | 4.1361e-08 | 8.7208e-05 | 0.00025
5.1 | 4.5378e-08 | 5.1 | 16711.68 | 1.5882 | 4.5378e-08 | 717741 | 16711 | 1.5879 | 4.5378e-08 | 717741 | 0.64825 | 1.5879 | 00A44723 | 00000A44 | 2628 | 0.657 | -1.3407 | 4.5378e-08 | 9.5037e-05 | 0.00025
5.2 | 4.9784e-08 | 5.2 | 17039.36 | 1.7654 | 4.9784e-08 | 787440 | 17039 | 1.7652 | 4.9784e-08 | 787440 | 0.7112 | 1.7652 | 00B43B10 | 00000B43 | 2883 | 0.72075 | -1.3337 | 4.9784e-08 | 0.00010361 | 0.00025
5.3 | 5.4619e-08 | 5.3 | 17367.04 | 1.9624 | 5.4619e-08 | 863907 | 17367 | 1.9624 | 5.4619e-08 | 863907 | 0.78026 | 1.9624 | 00C5BB8D | 00000C5B | 3163 | 0.79075 | -1.3348 | 5.4619e-08 | 0.000113 | 0.00025
5.4 | 5.9923e-08 | 5.4 | 17694.72 | 2.1814 | 5.9923e-08 | 947800 | 17694 | 2.1809 | 5.9923e-08 | 947800 | 0.85604 | 2.1809 | 00D8EF28 | 00000D8E | 3470 | 0.8675 | -1.3303 | 5.9923e-08 | 0.00012328 | 0.00025
5.5 | 6.5742e-08 | 5.5 | 18022.4 | 2.4248 | 6.5742e-08 | 1039840 | 18022 | 2.4244 | 6.5742e-08 | 1039840 | 0.93916 | 2.4244 | 00EE0020 | 00000EE0 | 3808 | 0.952 | -1.3574 | 6.5742e-08 | 0.00013454 | 0.00025
5.6 | 7.2126e-08 | 5.6 | 18350.08 | 2.6953 | 7.2126e-08 | 1140818 | 18350 | 2.6952 | 7.2126e-08 | 1140818 | 1.0304 | 2.6952 | 01051CCE | 00001051 | 4177 | 1.0442 | -1.3385 | 7.2126e-08 | 0.00014688 | 0.00025
5.7 | 7.913e-08 | 5.7 | 18677.76 | 2.996 | 7.913e-08 | 1251601 | 18677 | 2.9953 | 7.913e-08 | 1251601 | 1.1304 | 2.9953 | 011E77FF | 000011E7 | 4583 | 1.1458 | -1.3467 | 7.913e-08 | 0.00016041 | 0.00025
5.8 | 8.6814e-08 | 5.8 | 19005.44 | 3.3303 | 8.6814e-08 | 1373143 | 19005 | 3.3298 | 8.6814e-08 | 1373143 | 1.2402 | 3.3298 | 013A4999 | 000013A4 | 5028 | 1.257 | -1.3457 | 8.6814e-08 | 0.00017523 | 0.00025
5.9 | 9.5244e-08 | 5.9 | 19333.12 | 3.7019 | 9.5244e-08 | 1506487 | 19333 | 3.7017 | 9.5244e-08 | 1506487 | 1.3606 | 3.7017 | 0158CEB9 | 0000158C | 5516 | 1.379 | -1.341 | 9.5244e-08 | 0.00019148 | 0.00025
6 | 1.0449e-07 | 6 | 19660.8 | 4.1149 | 1.0449e-07 | 1652781 | 19660 | 4.1138 | 1.0449e-07 | 1652781 | 1.4928 | 4.1138 | 017A4AA3 | 000017A4 | 6052 | 1.513 | -1.3467 | 1.0449e-07 | 0.0002093 | 0.00025
6.1 | 1.1464e-07 | 6.1 | 19988.48 | 4.574 | 1.1464e-07 | 1813280 | 19988 | 4.5733 | 1.1464e-07 | 1813280 | 1.6377 | 4.5733 | 019F06E0 | 000019F0 | 6640 | 1.66 | -1.3512 | 1.1464e-07 | 0.00022883 | 0.00025
6.2 | 1.2577e-07 | 6.2 | 20316.16 | 5.0844 | 1.2577e-07 | 1989366 | 20316 | 5.0841 | 1.2577e-07 | 1989366 | 1.7968 | 5.0841 | 01C7546A | 00001C75 | 7285 | 1.8213 | -1.3538 | 1.2577e-07 | 0.00025026 | 0.00025
6.3 | 1.3799e-07 | 6.3 | 20643.84 | 5.6517 | 1.3799e-07 | 2182551 | 20643 | 5.6501 | 1.3799e-07 | 2182551 | 1.9712 | 5.6501 | 01F38BD9 | 00001F38 | 7992 | 1.998 | -1.3484 | 1.3799e-07 | 0.00027377 | 0.00025
6.4 | 1.5139e-07 | 6.4 | 20971.52 | 6.2822 | 1.5139e-07 | 2394496 | 20971 | 6.2812 | 1.5139e-07 | 2394496 | 2.1627 | 6.2812 | 02240E80 | 00002240 | 8768 | 2.192 | -1.3473 | 1.5139e-07 | 0.00029956 | 0.00025
6.5 | 1.6609e-07 | 6.5 | 21299.2 | 6.9832 | 1.6609e-07 | 2627023 | 21299 | 6.9827 | 1.6609e-07 | 2627023 | 2.3727 | 6.9827 | 02594721 | 00002594 | 9620 | 2.405 | -1.353 | 1.6609e-07 | 0.00032786 | 0.00025
6.6 | 1.8222e-07 | 6.6 | 21626.88 | 7.7623 | 1.8222e-07 | 2882130 | 21626 | 7.7601 | 1.8222e-07 | 2882130 | 2.6031 | 7.7601 | 0293AACE | 0000293A | 10554 | 2.6385 | -1.3512 | 1.8222e-07 | 0.00035892 | 0.00025
6.7 | 1.9991e-07 | 6.7 | 21954.56 | 8.6284 | 1.9991e-07 | 3162011 | 21954 | 8.6268 | 1.9991e-07 | 3162011 | 2.8559 | 8.6268 | 02D3BA15 | 00002D3B | 11579 | 2.8948 | -1.3522 | 1.9991e-07 | 0.00039301 | 0.00025
6.8 | 2.1932e-07 | 6.8 | 22282.24 | 9.5911 | 2.1932e-07 | 3469070 | 22282 | 9.5904 | 2.1932e-07 | 3469070 | 3.1332 | 9.5904 | 031A01D2 | 000031A0 | 12704 | 3.176 | -1.3567 | 2.1932e-07 | 0.00043044 | 0.00025
6.9 | 2.4062e-07 | 6.9 | 22609.92 | 10.6612 | 2.4062e-07 | 3805948 | 22609 | 10.658 | 2.4062e-07 | 3805948 | 3.4375 | 10.658 | 03671CC4 | 00003671 | 13937 | 3.4842 | -1.3519 | 2.4062e-07 | 0.00047153 | 0.00025
7 | 2.6399e-07 | 7 | 22937.6 | 11.8507 | 2.6399e-07 | 4175539 | 22937 | 11.8484 | 2.6399e-07 | 4175539 | 3.7713 | 11.8484 | 03BBB47D | 00003BBB | 15291 | 3.8228 | -1.3558 | 2.6399e-07 | 0.00051665 | 0.00025
7.1 | 2.8962e-07 | 7.1 | 23265.28 | 13.173 | 2.8962e-07 | 4581021 | 23265 | 13.1718 | 2.8962e-07 | 4581021 | 4.1375 | 13.1718 | 04188333 | 00004188 | 16776 | 4.194 | -1.3564 | 2.8962e-07 | 0.00056621 | 0.00025
7.2 | 3.1775e-07 | 7.2 | 23592.96 | 14.6427 | 3.1775e-07 | 5025878 | 23593 | 14.6429 | 3.1775e-07 | 5025878 | 4.5393 | 14.6429 | 047E550A | 000047E5 | 18405 | 4.6013 | -1.3559 | 3.1775e-07 | 0.00062064 | 0.00025
7.35 | 3.6514e-07 | 7.35 | 24084.48 | 17.1605 | 3.6514e-07 | 5775460 | 24084 | 17.1578 | 3.6514e-07 | 5775460 | 5.2163 | 17.1578 | 0529E5DC | 0000529E | 21150 | 5.2875 | -1.3559 | 3.6514e-07 | 0.00071251 | 0.00025
7.4 | 3.8246e-07 | 7.4 | 24248.32 | 18.0925 | 3.8246e-07 | 6049388 | 24248 | 18.0906 | 3.8246e-07 | 6049388 | 5.4637 | 18.0906 | 05689854 | 00005689 | 22153 | 5.5382 | -1.3553 | 3.8246e-07 | 0.00074613 | 0.00025
7.5 | 4.196e-07 | 7.5 | 24576 | 20.1112 | 4.196e-07 | 6636837 | 24576 | 20.1112 | 4.196e-07 | 6636837 | 5.9943 | 20.1112 | 05EF0D2B | 00005EF0 | 24304 | 6.076 | -1.3542 | 4.196e-07 | 0.00081833 | 0.00025
7.6 | 4.6035e-07 | 7.6 | 24903.68 | 22.3551 | 4.6035e-07 | 7281333 | 24903 | 22.3502 | 4.6035e-07 | 7281333 | 6.5764 | 22.3502 | 0682909B | 00006829 | 26665 | 6.6662 | -1.3575 | 4.6035e-07 | 0.00089766 | 0.00025
7.7 | 5.0505e-07 | 7.7 | 25231.36 | 24.8493 | 5.0505e-07 | 7988415 | 25231 | 24.8464 | 5.0505e-07 | 7988415 | 7.215 | 24.8464 | 07246731 | 00007246 | 29254 | 7.3135 | -1.3561 | 5.0505e-07 | 0.00098486 | 0.00025
7.8 | 5.5409e-07 | 7.8 | 25559.04 | 27.6218 | 5.5409e-07 | 8764160 | 25559 | 27.6215 | 5.5409e-07 | 8764160 | 7.9156 | 27.6215 | 07D5F500 | 00007D5F | 32095 | 8.0237 | -1.3567 | 5.5409e-07 | 0.0010807 | 0.00025
7.9 | 6.079e-07 | 7.9 | 25886.72 | 30.7037 | 6.079e-07 | 9615238 | 25886 | 30.6966 | 6.079e-07 | 9615238 | 8.6843 | 30.6966 | 0898C0DA | 0000898C | 35212 | 8.803 | -1.3575 | 6.079e-07 | 0.0011861 | 0.00025
8 | 6.6693e-07 | 8 | 26214.4 | 34.1295 | 6.6693e-07 | 10548963 | 26214 | 34.1251 | 6.6693e-07 | 10548963 | 9.5276 | 34.1251 | 096E774D | 000096E7 | 38631 | 9.6578 | -1.3565 | 6.6693e-07 | 0.0013019 | 0.00025
8.1 | 7.317e-07 | 8.1 | 26542.08 | 37.9374 | 7.317e-07 | 11573360 | 26542 | 37.9365 | 7.317e-07 | 11573360 | 10.4528 | 37.9365 | 0A58EE90 | 0000A58E | 42382 | 10.5955 | -1.3555 | 7.317e-07 | 0.0014293 | 0.00025
8.2 | 8.0275e-07 | 8.2 | 26869.76 | 42.1703 | 8.0275e-07 | 12697236 | 26869 | 42.1599 | 8.0275e-07 | 12697236 | 11.4679 | 42.1599 | 0B5A2AAC | 0000B5A2 | 46498 | 11.6245 | -1.3562 | 8.0275e-07 | 0.0015694 | 0.00025
8.3 | 8.8071e-07 | 8.3 | 27197.44 | 46.8754 | 8.8071e-07 | 13930250 | 27197 | 46.8687 | 8.8071e-07 | 13930250 | 12.5815 | 46.8687 | 0C746196 | 0000C746 | 51014 | 12.7535 | -1.3574 | 8.8071e-07 | 0.0017235 | 0.00025
8.4 | 9.6623e-07 | 8.4 | 27525.12 | 52.1055 | 9.6623e-07 | 15283001 | 27525 | 52.1035 | 9.6623e-07 | 15283001 | 13.8033 | 52.1035 | 0DAA0057 | 0000DAA0 | 55968 | 13.992 | -1.3576 | 9.6623e-07 | 0.001893 | 0.00025
8.5 | 1.0601e-06 | 8.5 | 27852.8 | 57.9191 | 1.0601e-06 | 16767116 | 27852 | 57.9042 | 1.0601e-06 | 16767116 | 15.1438 | 57.9042 | 0EFDB034 | 0000EFDB | 61403 | 15.3507 | -1.3576 | 1.0601e-06 | 0.0020795 | 0.00025
The 16 bit range limit
8.6 | 1.163e-06 | 8.6 | 28180.48 | 64.3814 | 1.163e-06 | 18395351 | 28180 | 64.3714 | 1.163e-06 | 18395351 | 16.6143 | 64.3714 | 10725C99 | 00010725 | 67365 | 16.8412 | -1.3565
8.7 | 1.2759e-06 | 8.7 | 28508.16 | 71.5647 | 1.2759e-06 | 20181703 | 28508 | 71.561 | 1.2759e-06 | 20181703 | 18.2277 | 71.561 | 120B39A9 | 000120B3 | 73907 | 18.4767 | -1.3568
8.8 | 1.3998e-06 | 8.8 | 28835.84 | 79.5495 | 1.3998e-06 | 22141525 | 28835 | 79.5279 | 1.3998e-06 | 22141525 | 19.9978 | 79.5279 | 13CBCAFB | 00013CBC | 81084 | 20.271 | -1.3568
8.9 | 1.5358e-06 | 8.9 | 29163.52 | 88.4251 | 1.5358e-06 | 24291663 | 29163 | 88.4103 | 1.5358e-06 | 24291663 | 21.9398 | 88.4103 | 15B7EBA1 | 00015B7E | 88958 | 22.2395 | -1.3568
9 | 1.6849e-06 | 9 | 29491.2 | 98.2911 | 1.6849e-06 | 26650599 | 29491 | 98.2848 | 1.6849e-06 | 26650599 | 24.0703 | 98.2848 | 17D3D689 | 00017D3D | 97597 | 24.3992 | -1.3572
9.1 | 1.8485e-06 | 9.1 | 29818.88 | 109.2578 | 1.8485e-06 | 29238608 | 29818 | 109.2268 | 1.8485e-06 | 29238608 | 26.4078 | 109.2268 | 1A242FB0 | 0001A242 | 107074 | 26.7685 | -1.3567
9.2 | 2.0281e-06 | 9.2 | 30146.56 | 121.4482 | 2.0281e-06 | 32077936 | 30146 | 121.4263 | 2.0281e-06 | 32077936 | 28.9722 | 121.4263 | 1CAE0E90 | 0001CAE0 | 117472 | 29.368 | -1.3569
9.3 | 2.225e-06 | 9.3 | 30474.24 | 134.9987 | 2.225e-06 | 35192988 | 30474 | 134.9882 | 2.225e-06 | 35192988 | 31.7857 | 134.9882 | 1F770924 | 0001F770 | 128880 | 32.22 | -1.3572
9.4 | 2.4411e-06 | 9.4 | 30801.92 | 150.0611 | 2.4411e-06 | 38610538 | 30801 | 150.0165 | 2.4411e-06 | 38610538 | 34.8723 | 150.0165 | 22854036 | 00022854 | 141396 | 35.349 | -1.3576
9.5 | 2.6781e-06 | 9.5 | 31129.6 | 166.804 | 2.6781e-06 | 42359964 | 31129 | 166.7717 | 2.6781e-06 | 42359964 | 38.2587 | 166.7717 | 25DF6D24 | 00025DF6 | 155126 | 38.7815 | -1.3571
9.6 | 2.9382e-06 | 9.6 | 31457.28 | 185.415 | 2.9382e-06 | 46473491 | 31457 | 185.3983 | 2.9382e-06 | 46473491 | 41.974 | 185.3983 | 298CF01D | 000298CF | 170191 | 42.5478 | -1.3576
9.7 | 3.2235e-06 | 9.7 | 31784.96 | 206.1026 | 3.2235e-06 | 50986479 | 31785 | 206.1052 | 3.2235e-06 | 50986479 | 46.0501 | 206.1052 | 2D95E101 | 0002D95E | 186718 | 46.6795 | -1.3576
9.8 | 3.5365e-06 | 9.8 | 32112.64 | 229.0983 | 3.5365e-06 | 55937717 | 32112 | 229.051 | 3.5365e-06 | 55937717 | 50.5219 | 229.051 | 3203209B | 00032032 | 204850 | 51.2125 | -1.3576
9.9 | 3.88e-06 | 9.9 | 32440.32 | 254.6597 | 3.88e-06 | 61369765 | 32440 | 254.6334 | 3.88e-06 | 61369765 | 55.428 | 254.6334 | 36DE6CAB | 00036DE6 | 224742 | 56.1855 | -1.3573
10 | 4.2567e-06 | 10 | 32768 | 283.0732 | 4.2567e-06 | 67329313 | 32768 | 283.0732 | 4.2567e-06 | 67329313 | 60.8106 | 283.0732 | 3C3274EF | 0003C327 | 246567 | 61.6418 | -1.3575
Given that we know that the Voltage output range of the Log Integrators has a range of 0.0002 Volts to 10.0 Volts we can use the calibration coefficients we have measured to determine the corresponding amount of charge from the detector that produces the Voltage output.
Applying the scaling of the MADC digitizers and the scaling performed by the ACNET database, we can compute the range of Rad/Sec values that can result.
Using the range of charge determined for the Log Integrators, we can compute the range of Rad/Sec values that could be managed using the new Integrator digitizers. This is done here when using the new readouts with an equivalent Log conversion of the data being delivered to ACNET.
The draw back here is the complex math needed to be done quickly in the front-ends.
Using the new readout system we can convert the linear data directly to a 16 bit representation of the number of Rads. For our established range of charge in for the BLM detector, this results in a range in Rads from 0.0057 to 60.8 Rads. We have determined that by multiplying these values up by 4000 we can deliver to ACNET a 16 bit value with a bit more resolution and a range of 15.1438 Rads (57.9042 equivalent Rads/Sec). This top end is higher than we would ever allow a Booster BLM to experience, and is sufficiently high.
The determination of the scale factor between charge Qsum and Rads (x4000) is detailed in the specification document "Booster Beam Loss Monitor Data Acquisition and Presentation Specification". Besides converting Rads to a 16 bit integer for passing to ACNET, an approximation of the final scale factor is made to simplify the math needed in the front-end. The final columns below evaluate the error due to this approximation.

## COMPARISON
ACNET Log to Linear
conversion coefficients
d1 = | 0.007212 | [7 empty] | data column> | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25
d2 = | 1.0578 | [7 empty] | Rads | 0.0003 | 0.0007 | 0.001 | 0.002 | 0.004 | 0.007 | 0.01 | 0.02 | 0.04 | 0.07 | 0.1 | 0.2 | 0.4 | 0.7 | 1 | 2 | 4 | 7 | 10 | 16
[6 empty] | Choose Period | m selected |  | Rads/Sec | 0.00032301 | 0.00085022 | 0.0012778 | 0.0028204 | 0.0062253 | 0.011797 | 0.01773 | 0.039133 | 0.086375 | 0.16368 | 0.246 | 0.54297 | 1.1985 | 2.2711 | 3.4132 | 7.5337 | 16.6285 | 31.5108 | 47.358 | 81.0114
[6 empty] | L02 | 4.8688 |  | func(Rads) | 0.0014606 | 0.0034082 | 0.0048688 | 0.0097376 | 0.019475 | 0.034082 | 0.048688 | 0.097376 | 0.19475 | 0.34082 | 0.48688 | 0.97376 | 1.9475 | 3.4082 | 4.8688 | 9.7376 | 19.4752 | 34.0816 | 48.688 | 77.9008
C A L I B R A T I O N D A T A
[7 empty] | Linear Fit Rads/Sec = f(Rads) |  | Qsum=70nC/Rad | 2.1e-11 | 4.9e-11 | 7e-11 | 1.4e-10 | 2.8e-10 | 4.9e-10 | 7e-10 | 1.4e-09 | 2.8e-09 | 4.9e-09 | 7e-09 | 1.4e-08 | 2.8e-08 | 4.9e-08 | 7e-08 | 1.4e-07 | 2.8e-07 | 4.9e-07 | 7e-07 | 1.12e-06
channel | C1 | C2 | Channel | C1 | C2 | Period | m |  | Rads | 0.0003 | 0.0007 | 0.001 | 0.002 | 0.004 | 0.007 | 0.01 | 0.02 | 0.04 | 0.07 | 0.1 | 0.2 | 0.4 | 0.7 | 1 | 2 | 4 | 7 | 10 | 16
L10 | 1.0721 | 22.896 | 21 | 1.0517 | 22.157 | 21 | 1.5803 | Rads/Sec | 21 | 0.00014378 | 0.00036903 | 0.00054877 | 0.0011865 | 0.0025655 | 0.0047814 | 0.0071102 | 0.015374 | 0.033241 | 0.061951 | 0.092124 | 0.19919 | 0.43069 | 0.80268 | 1.1936 | 2.5808 | 5.5803 | 10.4 | 15.4654 | 26.0881
S10 | 1.0894 | 22.9734 | 23 | 1.0742 | 22.3916 | 23 | 1.4529 | Rads/Sec | 23 | 0.00010268 | 0.00026892 | 0.0004033 | 0.00088651 | 0.0019487 | 0.0036804 | 0.0055196 | 0.012133 | 0.02667 | 0.05037 | 0.075542 | 0.16605 | 0.365 | 0.68937 | 1.0339 | 2.2726 | 4.9954 | 9.4348 | 14.1496 | 24.137
L11 | 1.0679 | 22.6948 | 24 | 1.0529 | 22.5771 | 24 | 2.4209 | Rads/Sec | 24 | 0.0002173 | 0.00055834 | 0.00083066 | 0.0017976 | 0.0038903 | 0.0072555 | 0.010794 | 0.02336 | 0.050553 | 0.094283 | 0.14027 | 0.30355 | 0.65692 | 1.2252 | 1.8227 | 3.9446 | 8.5365 | 15.9209 | 23.686 | 39.9792
S11 | 1.0684 | 22.5817 | 25 | 1.0564 | 22.806 | 25 | 2.9302 | Rads/Sec | 25 | 0.00025297 | 0.00065203 | 0.00097131 | 0.0021074 | 0.0045721 | 0.0085447 | 0.012729 | 0.027617 | 0.059917 | 0.11198 | 0.16681 | 0.36191 | 0.7852 | 1.4674 | 2.186 | 4.7428 | 10.2899 | 19.2305 | 28.6473 | 48.4366
L08 | 1.0563 | 22.2452 | 26 | 1.044 | 22.7278 | 26 | 3.2415 | Rads/Sec | 26 | 0.00032176 | 0.00082013 | 0.001216 | 0.0026143 | 0.0056206 | 0.010427 | 0.015461 | 0.033239 | 0.071461 | 0.13257 | 0.19657 | 0.42261 | 0.90857 | 1.6856 | 2.4992 | 5.3731 | 11.5517 | 21.4303 | 31.7749 | 53.3939
S08 | 1.0567 | 22.3405 | 61 | 1.0903 | 24.1701 | 61 | 7.5163 | Rads/Sec | 61 | 0.00044335 | 0.001178 | 0.0017774 | 0.0039533 | 0.0087929 | 0.016766 | 0.025298 | 0.056267 | 0.12515 | 0.23863 | 0.36006 | 0.80085 | 1.7813 | 3.3965 | 5.1248 | 11.3986 | 25.353 | 48.342 | 72.9412 | 125.4251
L09 | 1.0739 | 22.8623 | 62 | 1.0854 | 23.8656 | 62 | 5.8561 | Rads/Sec | 62 | 0.000365 | 0.00096555 | 0.0014542 | 0.0032228 | 0.0071425 | 0.013579 | 0.020452 | 0.045325 | 0.10045 | 0.19098 | 0.28763 | 0.63745 | 1.4127 | 2.6859 | 4.0452 | 8.9651 | 19.8688 | 37.7751 | 56.8918 | 97.5892
S09 | 1.0947 | 23.3568 | 71 | 1.137 | 24.8525 | 71 | 7.762 | Rads/Sec | 71 | 0.0002708 | 0.00075026 | 0.0011522 | 0.002652 | 0.0061041 | 0.011965 | 0.018375 | 0.042294 | 0.097349 | 0.19082 | 0.29305 | 0.67451 | 1.5525 | 3.0433 | 4.6736 | 10.7572 | 24.7601 | 48.5352 | 74.5348 | 131.1766
L06 | 1.0665 | 21.7027 | 72 | 1.1181 | 24.4306 | 72 | 6.5722 | Rads/Sec | 72 | 0.00028379 | 0.00077302 | 0.0011787 | 0.0026755 | 0.0060731 | 0.011772 | 0.017949 | 0.040743 | 0.092483 | 0.17926 | 0.27333 | 0.62044 | 1.4083 | 2.7299 | 4.1623 | 9.4482 | 21.4467 | 41.5708 | 63.3846 | 110.506
S06 | 1.0551 | 22.3607 | L01 | 1.0587 | 23.1389 | L01 | 4.0266 | Rads/Sec | L01 | 0.00033866 | 0.00087471 | 0.0013042 | 0.0028344 | 0.00616 | 0.011528 | 0.017188 | 0.037355 | 0.081184 | 0.15193 | 0.22653 | 0.49231 | 1.0699 | 2.0023 | 2.9854 | 6.4883 | 14.101 | 26.3892 | 39.346 | 66.6028
L07 | 1.0592 | 22.1628 | L02 | 1.0798 | 23.6135 | L02 | 4.8688 | Rads/Sec | L02 | 0.00032301 | 0.00085022 | 0.0012778 | 0.0028204 | 0.0062253 | 0.011797 | 0.01773 | 0.039133 | 0.086375 | 0.16368 | 0.246 | 0.54297 | 1.1985 | 2.2711 | 3.4132 | 7.5337 | 16.6285 | 31.5108 | 47.358 | 81.0114
S07 | 1.0665 | 22.1066 | L03 | 1.0661 | 23.2804 | L03 | 4.1923 | Rads/Sec | L03 | 0.00032448 | 0.00084365 | 0.0012614 | 0.0027563 | 0.0060229 | 0.011321 | 0.016927 | 0.036987 | 0.08082 | 0.15191 | 0.22714 | 0.49632 | 1.0845 | 2.0385 | 3.0479 | 6.6601 | 14.5532 | 27.355 | 40.8999 | 69.4883
L16 | 1.0246 | 22.346 | L04 | 1.0854 | 23.8134 | L04 | 5.544 | Rads/Sec | L04 | 0.00034569 | 0.00091442 | 0.0013772 | 0.003052 | 0.0067638 | 0.012859 | 0.019367 | 0.04292 | 0.095118 | 0.18084 | 0.27235 | 0.60358 | 1.3376 | 2.5431 | 3.83 | 8.488 | 18.8108 | 35.7629 | 53.8606 | 92.388
S16 | 1.0848 | 23.6356 | L05 | 1.0852 | 23.6941 | L05 | 4.8983 | Rads/Sec | L05 | 0.00030596 | 0.00080923 | 0.0012187 | 0.0027005 | 0.005984 | 0.011376 | 0.017131 | 0.037961 | 0.084119 | 0.15991 | 0.24082 | 0.53364 | 1.1825 | 2.2479 | 3.3853 | 7.5015 | 16.6228 | 31.6001 | 47.5884 | 81.6227
L17 | 1.0554 | 22.9564 | L06 | 1.0665 | 21.7027 | L06 | 0.78535 | Rads/Sec | L06 | 6.0507e-05 | 0.00015738 | 0.00023534 | 0.0005144 | 0.0011244 | 0.0021139 | 0.0031611 | 0.0069095 | 0.015103 | 0.028395 | 0.042461 | 0.092811 | 0.20286 | 0.38141 | 0.57035 | 1.2467 | 2.7249 | 5.1232 | 7.6612 | 13.0188
S17 | 1.0718 | 23.25 | L07 | 1.0592 | 22.1628 | L07 | 1.4233 | Rads/Sec | L07 | 0.00011904 | 0.0003076 | 0.00045872 | 0.0009973 | 0.0021682 | 0.0040589 | 0.006053 | 0.01316 | 0.028611 | 0.053559 | 0.079871 | 0.17365 | 0.37753 | 0.70673 | 1.0539 | 2.2913 | 4.9816 | 9.3255 | 13.9068 | 23.5465
L14 | 1.097 | 23.695 | L08 | 1.0563 | 22.2452 | L08 | 1.622 | Rads/Sec | L08 | 0.00014021 | 0.00036135 | 0.00053827 | 0.0011677 | 0.0025333 | 0.0047341 | 0.007052 | 0.015299 | 0.033189 | 0.062022 | 0.092389 | 0.20043 | 0.43482 | 0.81256 | 1.2104 | 2.6259 | 5.6966 | 10.6455 | 15.8577 | 26.8105
S14 | 1.1033 | 23.7398 | L09 | 1.0739 | 22.8623 | L09 | 2.4018 | Rads/Sec | L09 | 0.00017036 | 0.00044604 | 0.00066886 | 0.0014699 | 0.0032302 | 0.0060997 | 0.0091468 | 0.020101 | 0.044174 | 0.083415 | 0.12508 | 0.27489 | 0.60409 | 1.1407 | 1.7105 | 3.7591 | 8.261 | 15.5994 | 23.392 | 39.8964
L15 | 1.1107 | 23.8796 | L10 | 1.0721 | 22.896 | L10 | 2.5553 | Rads/Sec | L10 | 0.00018492 | 0.00048338 | 0.00072436 | 0.0015898 | 0.0034891 | 0.0065817 | 0.0098629 | 0.021646 | 0.047508 | 0.089616 | 0.13429 | 0.29474 | 0.64688 | 1.2202 | 1.8285 | 4.0132 | 8.8079 | 16.6146 | 24.8975 | 42.4266
S15 | 1.1051 | 23.7679 | L11 | 1.0679 | 22.6948 | L11 | 2.1991 | Rads/Sec | L11 | 0.00016691 | 0.00043465 | 0.00065029 | 0.0014228 | 0.003113 | 0.0058573 | 0.0087634 | 0.019174 | 0.04195 | 0.078934 | 0.1181 | 0.25838 | 0.56533 | 1.0637 | 1.5915 | 3.482 | 7.6184 | 14.3347 | 21.4467 | 36.4692
L12 | 1.0493 | 22.7909 | L12 | 1.0493 | 22.7909 | L12 | 3.2008 | Rads/Sec | L12 | 0.00029912 | 0.0007661 | 0.0011382 | 0.0024567 | 0.0053026 | 0.0098686 | 0.014662 | 0.031646 | 0.068306 | 0.12712 | 0.18887 | 0.40765 | 0.87989 | 1.6375 | 2.4329 | 5.2512 | 11.3343 | 21.094 | 31.3397 | 52.8033
S12 | 1.1017 | 23.825 | L13 | 1.0718 | 23.1069 | L13 | 3.2085 | Rads/Sec | L13 | 0.00023298 | 0.00060884 | 0.00091226 | 0.0020017 | 0.0043923 | 0.0082838 | 0.012412 | 0.027235 | 0.059761 | 0.11271 | 0.16888 | 0.37056 | 0.81311 | 1.5335 | 2.2978 | 5.0419 | 11.0631 | 20.8649 | 31.2632 | 53.2661
L13 | 1.0718 | 23.1069 | L14 | 1.097 | 23.695 | L14 | 4.118 | Rads/Sec | L14 | 0.00022524 | 0.00060207 | 0.00091074 | 0.0020357 | 0.0045502 | 0.0087106 | 0.013176 | 0.029452 | 0.065831 | 0.12602 | 0.19063 | 0.42611 | 0.95243 | 1.8233 | 2.7581 | 6.1648 | 13.7797 | 26.3791 | 39.9033 | 68.8447
S13 | 1.0894 | 23.3996 | L15 | 1.1107 | 23.8796 | L15 | 4.0927 | Rads/Sec | L15 | 0.00019203 | 0.00051962 | 0.00079008 | 0.0017837 | 0.0040271 | 0.0077718 | 0.011817 | 0.026679 | 0.060232 | 0.11624 | 0.17674 | 0.39902 | 0.90086 | 1.7385 | 2.6434 | 5.968 | 13.4738 | 26.0027 | 39.5369 | 68.6766
L22 | 1.0724 | 23.8485 | L16 | 1.0246 | 22.346 | L16 | 2.8822 | Rads/Sec | L16 | 0.00035563 | 0.00089085 | 0.0013113 | 0.0027793 | 0.005891 | 0.010804 | 0.015903 | 0.033708 | 0.071446 | 0.13103 | 0.19287 | 0.4088 | 0.8665 | 1.5892 | 2.3391 | 4.958 | 10.5089 | 19.2734 | 28.3687 | 47.213
S22 | 1.1021 | 24.2856 | L17 | 1.0554 | 22.9564 | L17 | 3.4881 | Rads/Sec | L17 | 0.00030464 | 0.00078447 | 0.0011681 | 0.0025325 | 0.0054904 | 0.010255 | 0.01527 | 0.033105 | 0.071771 | 0.13405 | 0.19961 | 0.43275 | 0.93819 | 1.7523 | 2.6093 | 5.657 | 12.2641 | 22.906 | 34.1094 | 57.6424
L23 | 1.062 | 23.4057 | L18 | 1.0405 | 23.0732 | L18 | 4.9189 | Rads/Sec | L18 | 0.00050784 | 0.0012904 | 0.0019107 | 0.0040974 | 0.0087865 | 0.016267 | 0.024087 | 0.051652 | 0.11076 | 0.20506 | 0.30364 | 0.65114 | 1.3963 | 2.585 | 3.8278 | 8.2083 | 17.6021 | 32.5872 | 48.2536 | 80.9432
S23 | 1.1105 | 24.2963 | L19 | 1.0617 | 23.379 | L19 | 4.9692 | Rads/Sec | L19 | 0.00040432 | 0.001047 | 0.0015629 | 0.003404 | 0.0074139 | 0.013899 | 0.020746 | 0.045185 | 0.098414 | 0.1845 | 0.27539 | 0.59981 | 1.3064 | 2.4491 | 3.6557 | 7.9621 | 17.3415 | 32.5104 | 48.5266 | 82.2639
L20 | 1.0488 | 23.1454 | L20 | 1.0488 | 23.1454 | L20 | 4.693 | Rads/Sec | L20 | 0.00044113 | 0.0011293 | 0.0016775 | 0.0036193 | 0.0078089 | 0.014529 | 0.021581 | 0.046563 | 0.10046 | 0.18692 | 0.27765 | 0.59905 | 1.2925 | 2.4047 | 3.572 | 7.707 | 16.6286 | 30.9376 | 45.9556 | 77.4093
S20 | 1.1113 | 24.4039 | L21 | 1.0427 | 23.0798 | L21 | 4.791 | Rads/Sec | L21 | 0.00048227 | 0.0012279 | 0.0018197 | 0.0039087 | 0.0083958 | 0.015564 | 0.023066 | 0.049546 | 0.10642 | 0.19729 | 0.29238 | 0.62802 | 1.349 | 2.5007 | 3.7061 | 7.9605 | 17.0988 | 31.6978 | 46.9765 | 78.8894
L21 | 1.0427 | 23.0798 | L22 | 1.0724 | 23.8485 | L22 | 6.9645 | Rads/Sec | L22 | 0.00050211 | 0.0013129 | 0.0019677 | 0.0043196 | 0.0094828 | 0.017891 | 0.026814 | 0.058864 | 0.12922 | 0.2438 | 0.36539 | 0.80214 | 1.7609 | 3.3223 | 4.9793 | 10.9308 | 23.9962 | 45.2736 | 67.8526 | 115.6433
S21 | 1.0801 | 23.756 | L23 | 1.062 | 23.4057 | L23 | 5.0849 | Rads/Sec | L23 | 0.00041208 | 0.0010675 | 0.0015936 | 0.0034718 | 0.0075636 | 0.014183 | 0.021172 | 0.046126 | 0.10049 | 0.18843 | 0.28129 | 0.61282 | 1.3351 | 2.5034 | 3.7372 | 8.1418 | 17.7375 | 33.2598 | 49.6518 | 84.1862
L18 | 1.0405 | 23.0732 | L24 | 1.0503 | 23.0933 | L24 | 4.3456 | Rads/Sec | L24 | 0.00040179 | 0.0010299 | 0.0015307 | 0.0033063 | 0.0071413 | 0.013298 | 0.019764 | 0.042689 | 0.092204 | 0.1717 | 0.25518 | 0.55117 | 1.1905 | 2.2168 | 3.2948 | 7.1164 | 15.3709 | 28.6226 | 42.5404 | 71.7089
S18 | 1.0514 | 23.1911 | S01 | 1.0566 | 22.9983 | S01 | 3.5795 | Rads/Sec | S01 | 0.00030826 | 0.00079468 | 0.0011839 | 0.0025691 | 0.0055747 | 0.01042 | 0.015523 | 0.033685 | 0.073096 | 0.13662 | 0.20354 | 0.44168 | 0.95843 | 1.7914 | 2.6688 | 5.7913 | 12.5669 | 23.4889 | 34.9938 | 59.1738
L19 | 1.0617 | 23.379 | S02 | 1.0751 | 23.4548 | S02 | 4.417 | Rads/Sec | S02 | 0.00030916 | 0.00081029 | 0.0012156 | 0.0026737 | 0.0058809 | 0.011113 | 0.016672 | 0.03667 | 0.080656 | 0.15241 | 0.22865 | 0.50292 | 1.1062 | 2.0903 | 3.1359 | 6.8974 | 15.1711 | 28.6679 | 43.008 | 73.396
S19 | 1.1037 | 24.39 | S03 | 1.1018 | 24.1617 | S03 | 6.2873 | Rads/Sec | S03 | 0.00032594 | 0.00087497 | 0.0013259 | 0.0029741 | 0.006671 | 0.012807 | 0.019408 | 0.043532 | 0.097644 | 0.18745 | 0.28407 | 0.63718 | 1.4292 | 2.7438 | 4.1579 | 9.3263 | 20.9193 | 40.1602 | 60.8592 | 105.2491
L04 | 1.0854 | 23.8134 | S04 | 1.0886 | 23.8275 | S04 | 5.3634 | Rads/Sec | S04 | 0.00032242 | 0.00085538 | 0.0012898 | 0.0028653 | 0.0063651 | 0.012125 | 0.018283 | 0.040614 | 0.090224 | 0.17186 | 0.25915 | 0.57569 | 1.2789 | 2.4361 | 3.6734 | 8.1603 | 18.1277 | 34.5307 | 52.0687 | 89.4588
S04 | 1.0886 | 23.8275 | S05 | 1.096 | 23.7838 | S05 | 4.5928 | Rads/Sec | S05 | 0.00025412 | 0.00067865 | 0.0010262 | 0.002292 | 0.0051192 | 0.009794 | 0.01481 | 0.033077 | 0.073879 | 0.14134 | 0.21373 | 0.47736 | 1.0662 | 2.0398 | 3.0845 | 6.8892 | 15.3871 | 29.4385 | 44.5139 | 76.7602
L05 | 1.0852 | 23.6941 | S06 | 1.0551 | 22.3607 | S06 | 1.8654 | Rads/Sec | S06 | 0.00016344 | 0.00042075 | 0.00062647 | 0.0013579 | 0.0029433 | 0.0054963 | 0.0081837 | 0.017738 | 0.038448 | 0.071798 | 0.1069 | 0.23171 | 0.50225 | 0.9379 | 1.3965 | 3.0269 | 6.5608 | 12.2518 | 18.2422 | 30.8237
S05 | 1.096 | 23.7838 | S07 | 1.0665 | 22.1066 | S07 | 1.2051 | Rads/Sec | S07 | 9.2911e-05 | 0.00024164 | 0.00036134 | 0.00078977 | 0.0017262 | 0.0032453 | 0.0048529 | 0.010607 | 0.023183 | 0.043585 | 0.065175 | 0.14245 | 0.31135 | 0.58536 | 0.87531 | 1.9132 | 4.1815 | 7.8615 | 11.7556 | 19.976
L02 | 1.0798 | 23.6135 | S08 | 1.0567 | 22.3405 | S08 | 1.7837 | Rads/Sec | S08 | 0.00015353 | 0.0003958 | 0.00058968 | 0.0012796 | 0.0027768 | 0.0051903 | 0.0077327 | 0.01678 | 0.036414 | 0.068063 | 0.1014 | 0.22005 | 0.47751 | 0.89255 | 1.3297 | 2.8856 | 6.2618 | 11.7044 | 17.4376 | 29.4873
S02 | 1.0751 | 23.4548 | S09 | 1.0947 | 23.3568 | S09 | 2.9785 | Rads/Sec | S09 | 0.00016716 | 0.0004459 | 0.00067393 | 0.0015038 | 0.0033557 | 0.0064154 | 0.0096961 | 0.021636 | 0.04828 | 0.0923 | 0.1395 | 0.31129 | 0.69463 | 1.328 | 2.0071 | 4.4786 | 9.9939 | 19.1059 | 28.8763 | 49.7633
L03 | 1.0661 | 23.2804 | S10 | 1.0894 | 22.9734 | S10 | 2.1484 | Rads/Sec | S10 | 0.00012802 | 0.00033987 | 0.00051264 | 0.0011395 | 0.0025328 | 0.0048267 | 0.0072804 | 0.016182 | 0.035969 | 0.068548 | 0.10339 | 0.22982 | 0.51082 | 0.97349 | 1.4684 | 3.2638 | 7.2545 | 13.8251 | 20.853 | 35.8412
S03 | 1.1018 | 24.1617 | S11 | 1.0684 | 22.5817 | S11 | 1.9345 | Rads/Sec | S11 | 0.00014587 | 0.00038007 | 0.00056875 | 0.0012449 | 0.002725 | 0.005129 | 0.0076754 | 0.0168 | 0.036773 | 0.069217 | 0.10358 | 0.22672 | 0.49626 | 0.93408 | 1.3978 | 3.0596 | 6.6971 | 12.6055 | 18.8637 | 32.0862
L24 | 1.0503 | 23.0933 | S12 | 1.1017 | 23.825 | S12 | 4.4093 | Rads/Sec | S12 | 0.0002288 | 0.00061417 | 0.00093069 | 0.0020874 | 0.0046819 | 0.0089877 | 0.01362 | 0.030547 | 0.068514 | 0.13152 | 0.19931 | 0.44702 | 1.0026 | 1.9247 | 2.9166 | 6.5416 | 14.6722 | 28.1657 | 42.6812 | 73.809
S24 | 0.95399 | 20.7421 | S13 | 1.0894 | 23.3996 | S13 | 3.3736 | Rads/Sec | S13 | 0.00020111 | 0.00053389 | 0.00080527 | 0.0017899 | 0.0039783 | 0.0075814 | 0.011435 | 0.025417 | 0.056494 | 0.10766 | 0.16239 | 0.36093 | 0.80224 | 1.5288 | 2.306 | 5.1254 | 11.3922 | 21.7101 | 32.7458 | 56.2812
L01 | 1.0587 | 23.1389 | S14 | 1.1033 | 23.7398 | S14 | 3.9365 | Rads/Sec | S14 | 0.00020068 | 0.00053943 | 0.00081791 | 0.0018366 | 0.0041241 | 0.0079243 | 0.012015 | 0.02698 | 0.060584 | 0.11641 | 0.17651 | 0.39635 | 0.88999 | 1.7101 | 2.5929 | 5.8224 | 13.0741 | 25.1215 | 38.0908 | 65.9226
S01 | 1.0566 | 22.9983 | S15 | 1.1051 | 23.7679 | S15 | 3.9484 | Rads/Sec | S15 | 0.00019723 | 0.00053103 | 0.00080574 | 0.0018117 | 0.0040735 | 0.0078354 | 0.011889 | 0.026731 | 0.060105 | 0.11561 | 0.17542 | 0.39442 | 0.88685 | 1.7059 | 2.5883 | 5.8198 | 13.0855 | 25.1703 | 38.1909 | 66.1553
24 | 1.0529 | 22.5771 | S16 | 1.0848 | 23.6356 | S16 | 4.6304 | Rads/Sec | S16 | 0.00029048 | 0.00076801 | 0.0011564 | 0.0025618 | 0.0056752 | 0.010786 | 0.016241 | 0.035979 | 0.079704 | 0.15148 | 0.2281 | 0.5053 | 1.1194 | 2.1275 | 3.2034 | 7.0966 | 15.721 | 29.8791 | 44.9902 | 77.1517
23 | 1.0742 | 22.3916 | S17 | 1.0718 | 23.25 | S17 | 3.7317 | Rads/Sec | S17 | 0.00027092 | 0.000708 | 0.0010608 | 0.0023278 | 0.0051078 | 0.0096334 | 0.014434 | 0.031673 | 0.0695 | 0.13108 | 0.1964 | 0.43096 | 0.94566 | 1.7835 | 2.6724 | 5.8639 | 12.8671 | 24.2675 | 36.3617 | 61.9533
11 | 1.0604 | 22.6922 | S18 | 1.0514 | 23.1911 | S18 | 4.7428 | Rads/Sec | S18 | 0.00043321 | 0.0011115 | 0.0016527 | 0.0035726 | 0.0077226 | 0.01439 | 0.021395 | 0.046249 | 0.099973 | 0.18628 | 0.27698 | 0.59872 | 1.2942 | 2.4115 | 3.5856 | 7.7508 | 16.7543 | 31.2187 | 46.4177 | 78.2871
21 | 1.0517 | 22.157 | S19 | 1.1037 | 24.39 | S19 | 7.7871 | Rads/Sec | S19 | 0.00039532 | 0.001063 | 0.001612 | 0.0036207 | 0.0081325 | 0.01563 | 0.023702 | 0.053238 | 0.11958 | 0.22981 | 0.34851 | 0.78278 | 1.7582 | 3.3791 | 5.1243 | 11.5096 | 25.8518 | 49.6842 | 75.345 | 130.4213
52 | 1.1129 | 23.846 | S20 | 1.1113 | 24.4039 | S20 | 7.0607 | Rads/Sec | S20 | 0.00032898 | 0.00089067 | 0.0013546 | 0.0030596 | 0.0069107 | 0.013342 | 0.020291 | 0.045831 | 0.10352 | 0.19985 | 0.30394 | 0.68652 | 1.5506 | 2.9937 | 4.5529 | 10.2837 | 23.2277 | 44.8432 | 68.1998 | 118.5019
51 | 1.126 | 23.9854 | S21 | 1.0801 | 23.756 | S21 | 5.6366 | Rads/Sec | S21 | 0.00037272 | 0.00098132 | 0.001475 | 0.0032563 | 0.007189 | 0.013625 | 0.02048 | 0.045214 | 0.099819 | 0.18919 | 0.28437 | 0.62779 | 1.386 | 2.6269 | 3.9484 | 8.7168 | 19.2442 | 36.4738 | 54.823 | 93.7947
26 | 1.044 | 22.7278 | S22 | 1.1021 | 24.2856 | S22 | 7.1348 | Rads/Sec | S22 | 0.00036856 | 0.00098968 | 0.0014999 | 0.0033652 | 0.00755 | 0.014497 | 0.021972 | 0.049294 | 0.11059 | 0.21235 | 0.32184 | 0.72207 | 1.62 | 3.1106 | 4.7144 | 10.577 | 23.73 | 45.5646 | 69.0573 | 119.4453
25 | 1.0564 | 22.806 | S23 | 1.1105 | 24.2963 | S23 | 6.3742 | Rads/Sec | S23 | 0.00029961 | 0.00081059 | 0.0012324 | 0.0027821 | 0.0062803 | 0.012119 | 0.018426 | 0.041595 | 0.093897 | 0.18119 | 0.27549 | 0.62188 | 1.4038 | 2.709 | 4.1188 | 9.2977 | 20.9888 | 40.502 | 61.5794 | 106.9569
71 | 1.137 | 24.8525 | S24 | 0.95399 | 20.7421 | S24 | 1.5012 | Rads/Sec | S24 | 0.00040893 | 0.00096155 | 0.0013781 | 0.0027737 | 0.0055825 | 0.0098193 | 0.014073 | 0.028325 | 0.057008 | 0.10027 | 0.14371 | 0.28925 | 0.58216 | 1.024 | 1.4676 | 2.9538 | 5.945 | 10.4569 | 14.987 | 24.082
72 | 1.1181 | 24.4306
62 | 1.0854 | 23.8656
61 | 1.0903 | 24.1701
rms | 1.0723 | 23.2652
choose | 1.079 | 23.344
Choose a Period to plot.

## BLM ALARMS
Expected
|  | Rad/Sec | scale fact. | Rad
Period |  | Max Limit | m | Max Limit
L | 01 | 25 | 4.0266 | 6.2088
L | 02 | 28 | 4.8688 | 5.7509
L | 03 | 150 | 4.1923 | 35.7801
L | 04 | 50 | 5.544 | 9.0187
L | 05 | 50 | 4.8983 | 10.2077
L | 06 | 600 | 0.78535 | 763.9949
L | 07 | 280 | 1.4233 | 196.7214
L | 08 | 100 | 1.622 | 61.653
L | 09 | 40 | 2.4018 | 16.6545
L | 10 | 100 | 2.5553 | 39.1336
L | 11 | 28 | 2.1991 | 12.7323
L | 12 | 26 | 3.2008 | 8.123
L | 13 | 500 | 3.2085 | 155.8373
L | 14 | 55 | 4.118 | 13.356
L | 15 | 40 | 4.0927 | 9.7736
L | 16 | 50 | 2.8822 | 17.3476
L | 17 | 25 | 3.4881 | 7.1671
L | 18 | 42 | 4.9189 | 8.5385
L | 19 | 95 | 4.9692 | 19.1176
L | 20 | 75 | 4.693 | 15.9813
L | 21 | 50 | 4.791 | 10.4362
L | 22 | 43 | 6.9645 | 6.1742
L | 23 | 50 | 5.0849 | 9.8331
L | 24 | 38 | 4.3456 | 8.7445
S | 01 | 340 | 3.5795 | 94.984
S | 02 | 1300 | 4.417 | 294.3197
S | 03 | 163 | 6.2873 | 25.9251
S | 04 | 100 | 5.3634 | 18.6447
S | 05 | 300 | 4.5928 | 65.32
S | 06 | 1800 | 1.8654 | 964.9404
S | 07 | 200 | 1.2051 | 165.9679
S | 08 | 120 | 1.7837 | 67.2749
S | 09 | 100 | 2.9785 | 33.5739
S | 10 | 100 | 2.1484 | 46.5468
S | 11 | 85 | 1.9345 | 43.9385
S | 12 | 100 | 4.4093 | 22.6794
S | 13 | 500 | 3.3736 | 148.2093
S | 14 | 118 | 3.9365 | 29.9762
S | 15 | 190 | 3.9484 | 48.1207
S | 16 | 120 | 4.6304 | 25.9155
S | 17 | 120 | 3.7317 | 32.1567
S | 18 | 288 | 4.7428 | 60.7239
S | 19 | 300 | 7.7871 | 38.5252
S | 20 | 100 | 7.0607 | 14.1628
S | 21 | 225 | 5.6366 | 39.9175
S | 22 | 225 | 7.1348 | 31.5357
S | 23 | 135 | 6.3742 | 21.1791
S | 24 | 225 | 1.5012 | 149.8814
| 021 | 55 | 1.5803 | 34.8033
| 023 | 100 | 1.4529 | 68.8274
| 024 | 50 | 2.4209 | 20.6531
| 025 | 85 | 2.9302 | 29.0081
| 026 | 525 | 3.2415 | 161.9615
| 061 | 500 | 7.5163 | 66.5222
| 062 | 1600 | 5.8561 | 273.2191
| 071 | 280 | 7.762 | 36.0734
| 072 | 280 | 6.5722 | 42.6034
| 121 | 200
| 122 | 30
| 123 | 1500
| 124 | 30
| 125 | 180
| 126 | 325

## G11-RR6-1
BLM Locations and Control |  |  | MADC | MADC |  | RESET TRIG |  |  | L06 |  |  | S06 |  |  | L07 |  |  | S07
CHASSIS LOCATION | UNIT # | BLM NAME | Crate-Slot-Ch | # | RESET TRIG | CRATE-SLOT-CH | IRM NAME | Qsum=70nC/Rad | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads)
G11-RR6-1 | TWO | B:BLMS11 | 60_16_61 | 05 | B:BLMRT2 | 60-1-0 | B:IRMS11 | 2.1e-11 | 0.0003 | 6.0507e-05 | 0.0002356 | 0.0003 | 0.00016344 | 0.00055962 | 0.0003 | 0.00011904 | 0.000427 | 0.0003 | 9.2911e-05 | 0.00036152
G11-RR6-1 | TWO | B:BLML11 | 60_16_60 | 05 | B:BLMRT2 | 60-1-0 | B:IRML11 | 4.9e-11 | 0.0007 | 0.00015738 | 0.00054974 | 0.0007 | 0.00042075 | 0.0013058 | 0.0007 | 0.0003076 | 0.00099633 | 0.0007 | 0.00024164 | 0.00084354
G11-RR6-1 | TWO | B:BLMS10 | 60_16_59 | 05 | B:BLMRT2 | 60-1-0 | B:IRMS10 | 7e-11 | 0.001 | 0.00023534 | 0.00078535 | 0.001 | 0.00062647 | 0.0018654 | 0.001 | 0.00045872 | 0.0014233 | 0.001 | 0.00036134 | 0.0012051
G11-RR6-1 | TWO | B:BLML10 | 60_16_58 | 05 | B:BLMRT2 | 60-1-0 | B:IRML10 | 1.4e-10 | 0.002 | 0.0005144 | 0.0015707 | 0.002 | 0.0013579 | 0.0037308 | 0.002 | 0.0009973 | 0.0028467 | 0.002 | 0.00078977 | 0.0024101
G11-RR6-1 | TWO | B:BLMS09 | 60_16_57 | 05 | B:BLMRT2 | 60-1-0 | B:IRMS09 | 2.8e-10 | 0.004 | 0.0011244 | 0.0031414 | 0.004 | 0.0029433 | 0.0074616 | 0.004 | 0.0021682 | 0.0056933 | 0.004 | 0.0017262 | 0.0048202
G11-RR6-1 | TWO | B:BLML09 | 60_16_56 | 05 | B:BLMRT2 | 60-1-0 | B:IRML09 | 4.9e-10 | 0.007 | 0.0021139 | 0.0054974 | 0.007 | 0.0054963 | 0.013058 | 0.007 | 0.0040589 | 0.0099633 | 0.007 | 0.0032453 | 0.0084354
G11-RR6-1 | TWO | B:BLMS08 | 60_16_55 | 05 | B:BLMRT2 | 60-1-0 | B:IRMS08 | 7e-10 | 0.01 | 0.0031611 | 0.0078535 | 0.01 | 0.0081837 | 0.018654 | 0.01 | 0.006053 | 0.014233 | 0.01 | 0.0048529 | 0.012051
G11-RR6-1 | TWO | B:BLML08 | 60_16_54 | 05 | B:BLMRT2 | 60-1-0 | B:IRML08 | 1.4e-09 | 0.02 | 0.0069095 | 0.015707 | 0.02 | 0.017738 | 0.037308 | 0.02 | 0.01316 | 0.028467 | 0.02 | 0.010607 | 0.024101
G11-RR6-1 | TWO | B:BLMS07 | 60_16_53 | 05 | B:BLMRT2 | 60-1-0 | B:IRMS07 | 2.8e-09 | 0.04 | 0.015103 | 0.031414 | 0.04 | 0.038448 | 0.074616 | 0.04 | 0.028611 | 0.056933 | 0.04 | 0.023183 | 0.048202
G11-RR6-1 | TWO | B:BLML07 | 60_16_52 | 05 | B:BLMRT2 | 60-1-0 | B:IRML07 | 4.9e-09 | 0.07 | 0.028395 | 0.054974 | 0.07 | 0.071798 | 0.13058 | 0.07 | 0.053559 | 0.099633 | 0.07 | 0.043585 | 0.084354
G11-RR6-1 | TWO | B:BLMS06 | 60_16_51 | 05 | B:BLMRT2 | 60-1-0 | B:IRMS06 | 7e-09 | 0.1 | 0.042461 | 0.078535 | 0.1 | 0.1069 | 0.18654 | 0.1 | 0.079871 | 0.14233 | 0.1 | 0.065175 | 0.12051
G11-RR6-1 | TWO | B:BLML06 | 60_16_50 | 05 | B:BLMRT2 | 60-1-0 | B:IRML06 | 1.4e-08 | 0.2 | 0.092811 | 0.15707 | 0.2 | 0.23171 | 0.37308 | 0.2 | 0.17365 | 0.28467 | 0.2 | 0.14245 | 0.24101
[8 empty] | 2.8e-08 | 0.4 | 0.20286 | 0.31414 | 0.4 | 0.50225 | 0.74616 | 0.4 | 0.37753 | 0.56933 | 0.4 | 0.31135 | 0.48202
[8 empty] | 4.9e-08 | 0.7 | 0.38141 | 0.54974 | 0.7 | 0.9379 | 1.3058 | 0.7 | 0.70673 | 0.99633 | 0.7 | 0.58536 | 0.84354
[8 empty] | 7e-08 | 1 | 0.57035 | 0.78535 | 1 | 1.3965 | 1.8654 | 1 | 1.0539 | 1.4233 | 1 | 0.87531 | 1.2051
[8 empty] | 1.4e-07 | 2 | 1.2467 | 1.5707 | 2 | 3.0269 | 3.7308 | 2 | 2.2913 | 2.8467 | 2 | 1.9132 | 2.4101
[8 empty] | 2.8e-07 | 4 | 2.7249 | 3.1414 | 4 | 6.5608 | 7.4616 | 4 | 4.9816 | 5.6933 | 4 | 4.1815 | 4.8202
[8 empty] | 4.9e-07 | 7 | 5.1232 | 5.4974 | 7 | 12.2518 | 13.0578 | 7 | 9.3255 | 9.9633 | 7 | 7.8615 | 8.4354
[8 empty] | 7e-07 | 10 | 7.6612 | 7.8535 | 10 | 18.2422 | 18.654 | 10 | 13.9068 | 14.2333 | 10 | 11.7556 | 12.0505
[8 empty] | 1.12e-06 | 16 | 13.0188 | 12.5655 | 16 | 30.8237 | 29.8464 | 16 | 23.5465 | 22.7733 | 16 | 19.976 | 19.2808
[9 empty] | L08 |  |  | S08 |  |  |  | L09 |  |  | S09
[8 empty] | Qsum=70nC/Rad | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads)
[8 empty] | 2.1e-11 | 0.0003 | 0.00014021 | 0.00048659 | 0.0003 | 0.00015353 | 0.00053512 | 0.0003 | 0.00017036 | 0.00072053 | 0.0003 | 0.00016716 | 0.00089355
[8 empty] | 4.9e-11 | 0.0007 | 0.00036135 | 0.0011354 | 0.0007 | 0.0003958 | 0.0012486 | 0.0007 | 0.00044604 | 0.0016812 | 0.0007 | 0.0004459 | 0.002085
[8 empty] | 7e-11 | 0.001 | 0.00053827 | 0.001622 | 0.001 | 0.00058968 | 0.0017837 | 0.001 | 0.00066886 | 0.0024018 | 0.001 | 0.00067393 | 0.0029785
[8 empty] | 1.4e-10 | 0.002 | 0.0011677 | 0.003244 | 0.002 | 0.0012796 | 0.0035675 | 0.002 | 0.0014699 | 0.0048035 | 0.002 | 0.0015038 | 0.005957
[8 empty] | 2.8e-10 | 0.004 | 0.0025333 | 0.0064879 | 0.004 | 0.0027768 | 0.0071349 | 0.004 | 0.0032302 | 0.009607 | 0.004 | 0.0033557 | 0.011914
[8 empty] | 4.9e-10 | 0.007 | 0.0047341 | 0.011354 | 0.007 | 0.0051903 | 0.012486 | 0.007 | 0.0060997 | 0.016812 | 0.007 | 0.0064154 | 0.02085
[8 empty] | 7e-10 | 0.01 | 0.007052 | 0.01622 | 0.01 | 0.0077327 | 0.017837 | 0.01 | 0.0091468 | 0.024018 | 0.01 | 0.0096961 | 0.029785
[8 empty] | 1.4e-09 | 0.02 | 0.015299 | 0.03244 | 0.02 | 0.01678 | 0.035675 | 0.02 | 0.020101 | 0.048035 | 0.02 | 0.021636 | 0.05957
[8 empty] | 2.8e-09 | 0.04 | 0.033189 | 0.064879 | 0.04 | 0.036414 | 0.071349 | 0.04 | 0.044174 | 0.09607 | 0.04 | 0.04828 | 0.11914
[8 empty] | 4.9e-09 | 0.07 | 0.062022 | 0.11354 | 0.07 | 0.068063 | 0.12486 | 0.07 | 0.083415 | 0.16812 | 0.07 | 0.0923 | 0.2085
[8 empty] | 7e-09 | 0.1 | 0.092389 | 0.1622 | 0.1 | 0.1014 | 0.17837 | 0.1 | 0.12508 | 0.24018 | 0.1 | 0.1395 | 0.29785
[8 empty] | 1.4e-08 | 0.2 | 0.20043 | 0.3244 | 0.2 | 0.22005 | 0.35675 | 0.2 | 0.27489 | 0.48035 | 0.2 | 0.31129 | 0.5957
[8 empty] | 2.8e-08 | 0.4 | 0.43482 | 0.64879 | 0.4 | 0.47751 | 0.71349 | 0.4 | 0.60409 | 0.9607 | 0.4 | 0.69463 | 1.1914
[8 empty] | 4.9e-08 | 0.7 | 0.81256 | 1.1354 | 0.7 | 0.89255 | 1.2486 | 0.7 | 1.1407 | 1.6812 | 0.7 | 1.328 | 2.085
[8 empty] | 7e-08 | 1 | 1.2104 | 1.622 | 1 | 1.3297 | 1.7837 | 1 | 1.7105 | 2.4018 | 1 | 2.0071 | 2.9785
[8 empty] | 1.4e-07 | 2 | 2.6259 | 3.244 | 2 | 2.8856 | 3.5675 | 2 | 3.7591 | 4.8035 | 2 | 4.4786 | 5.957
[8 empty] | 2.8e-07 | 4 | 5.6966 | 6.4879 | 4 | 6.2618 | 7.1349 | 4 | 8.261 | 9.607 | 4 | 9.9939 | 11.914
[8 empty] | 4.9e-07 | 7 | 10.6455 | 11.3539 | 7 | 11.7044 | 12.4861 | 7 | 15.5994 | 16.8123 | 7 | 19.1059 | 20.8495
[8 empty] | 7e-07 | 10 | 15.8577 | 16.2198 | 10 | 17.4376 | 17.8373 | 10 | 23.392 | 24.0176 | 10 | 28.8763 | 29.785
[8 empty] | 1.12e-06 | 16 | 26.8105 | 25.9517 | 16 | 29.4873 | 28.5396 | 16 | 39.8964 | 38.4281 | 16 | 49.7633 | 47.6561
[9 empty] | L10 |  |  | S10 |  |  | L11 |  |  | S11
[8 empty] | Qsum=70nC/Rad | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads) | Rads | Rads/Sec | func(Rads)
[8 empty] | 2.1e-11 | 0.0003 | 0.00018492 | 0.0007666 | 0.0003 | 0.00012802 | 0.00064451 | 0.0003 | 0.00018492 | 0.0007666 | 0.0003 | 0.00014587 | 0.00058036
[8 empty] | 4.9e-11 | 0.0007 | 0.00048338 | 0.0017887 | 0.0007 | 0.00033987 | 0.0015039 | 0.0007 | 0.00048338 | 0.0017887 | 0.0007 | 0.00038007 | 0.0013542
[8 empty] | 7e-11 | 0.001 | 0.00072436 | 0.0025553 | 0.001 | 0.00051264 | 0.0021484 | 0.001 | 0.00072436 | 0.0025553 | 0.001 | 0.00056875 | 0.0019345
[8 empty] | 1.4e-10 | 0.002 | 0.0015898 | 0.0051107 | 0.002 | 0.0011395 | 0.0042968 | 0.002 | 0.0015898 | 0.0051107 | 0.002 | 0.0012449 | 0.003869
[8 empty] | 2.8e-10 | 0.004 | 0.0034891 | 0.010221 | 0.004 | 0.0025328 | 0.0085935 | 0.004 | 0.0034891 | 0.010221 | 0.004 | 0.002725 | 0.0077381
[8 empty] | 4.9e-10 | 0.007 | 0.0065817 | 0.017887 | 0.007 | 0.0048267 | 0.015039 | 0.007 | 0.0065817 | 0.017887 | 0.007 | 0.005129 | 0.013542
[8 empty] | 7e-10 | 0.01 | 0.0098629 | 0.025553 | 0.01 | 0.0072804 | 0.021484 | 0.01 | 0.0098629 | 0.025553 | 0.01 | 0.0076754 | 0.019345
[8 empty] | 1.4e-09 | 0.02 | 0.021646 | 0.051107 | 0.02 | 0.016182 | 0.042968 | 0.02 | 0.021646 | 0.051107 | 0.02 | 0.0168 | 0.03869
[8 empty] | 2.8e-09 | 0.04 | 0.047508 | 0.10221 | 0.04 | 0.035969 | 0.085935 | 0.04 | 0.047508 | 0.10221 | 0.04 | 0.036773 | 0.077381
[8 empty] | 4.9e-09 | 0.07 | 0.089616 | 0.17887 | 0.07 | 0.068548 | 0.15039 | 0.07 | 0.089616 | 0.17887 | 0.07 | 0.069217 | 0.13542
[8 empty] | 7e-09 | 0.1 | 0.13429 | 0.25553 | 0.1 | 0.10339 | 0.21484 | 0.1 | 0.13429 | 0.25553 | 0.1 | 0.10358 | 0.19345
[8 empty] | 1.4e-08 | 0.2 | 0.29474 | 0.51107 | 0.2 | 0.22982 | 0.42968 | 0.2 | 0.29474 | 0.51107 | 0.2 | 0.22672 | 0.3869
[8 empty] | 2.8e-08 | 0.4 | 0.64688 | 1.0221 | 0.4 | 0.51082 | 0.85935 | 0.4 | 0.64688 | 1.0221 | 0.4 | 0.49626 | 0.77381
[8 empty] | 4.9e-08 | 0.7 | 1.2202 | 1.7887 | 0.7 | 0.97349 | 1.5039 | 0.7 | 1.2202 | 1.7887 | 0.7 | 0.93408 | 1.3542
[8 empty] | 7e-08 | 1 | 1.8285 | 2.5553 | 1 | 1.4684 | 2.1484 | 1 | 1.8285 | 2.5553 | 1 | 1.3978 | 1.9345
[8 empty] | 1.4e-07 | 2 | 4.0132 | 5.1107 | 2 | 3.2638 | 4.2968 | 2 | 4.0132 | 5.1107 | 2 | 3.0596 | 3.869
[8 empty] | 2.8e-07 | 4 | 8.8079 | 10.2214 | 4 | 7.2545 | 8.5935 | 4 | 8.8079 | 10.2214 | 4 | 6.6971 | 7.7381
[8 empty] | 4.9e-07 | 7 | 16.6146 | 17.8874 | 7 | 13.8251 | 15.0386 | 7 | 16.6146 | 17.8874 | 7 | 12.6055 | 13.5417
[8 empty] | 7e-07 | 10 | 24.8975 | 25.5535 | 10 | 20.853 | 21.4838 | 10 | 24.8975 | 25.5535 | 10 | 18.8637 | 19.3452
[8 empty] | 1.12e-06 | 16 | 42.4266 | 40.8855 | 16 | 35.8412 | 34.374 | 16 | 42.4266 | 40.8855 | 16 | 32.0862 | 30.9524