## Slide 1: Status of Profile Monitors @ Fermilab
11 June 2013
Jim Zagel & Randy Thurman-Keup

## Slide 2: IPM’s in the Tev ERA
Booster Long 5
Antiproton Source
Debuncher
Main Injector MI-10
Tevatron E0

## Slide 3: IPM Basic Types
Booster (400MeV – 8 GeV)
  Electrostatic 10KV Clearing Field (Good at injection.)
Main Ring Original (8 GeV -150GeV)
  Electrostatic 30KV Clearing Field (Good at injection.)
Recycler Ultra High Vacuum (8 GeV)
  Electrostatic 30KV Clearing Field, e-11 Torr vacuum.
Main Injector Mark-II (8 GeV -150GeV)
  Permanent Magnetic Field 1KG and 10KV Clearing Field.
Tevatron (150 GeV – 1 TeV)
  Electro Magnet 1 KG and 10KV Clearing Field.

## Slide 4: IPM Concept
Magnet with vertical B field
Cathode
Field shaping electrodes
Electron Suppression Grid
Wire mesh gate
Microchannel Plate (MCP)
Anode strips
250 mm to 1.5mm spacing
Beam
(into page)
Ions
Electrons
Ionization Happens

## Slide 5: Why the Magnetic Field
Ions - E field only
Electrons - E and B field
Ion / electron paths with E and B field

## Slide 6: Booster
Horizontal and Vertical co-located in Long 5

## Slide 7: Main Ring/MI Original
Vertical at Q103
Also a Horizontal at Q102

## Slide 8: Main Injector Mark-II
Horizontal Measurement Permanent Magnet at Q104
Independent up and downstream +/- 25mm in horizontal plane for alignment and MCP Exposure

## Slide 9: Tevatron
Horizontal Detector
Horizontal Correction Magnet
Measured 36 Proton and 36 Anti-Proton bunches per turn using QIE Chips in tunnel.

## Slide 10: Mark-II, and Tevatron Internals
Tray design for quick extraction for MCP replacement, assures accurate realignment.

## Slide 11: Mark-III Internals Under Construction

## Slide 12: Mark-III IPM
Magnet with vertical B field
Cathode
Field shaping electrodes
Electron Suppression Grid
Wire mesh gate
Microchannel Plate (MCP)
Anode strips 500 mm
Beam
(into page)
Ions
Electrons
Ionization Happens

## Slide 13: Mark-III Control Grid

## Slide 14: Anode Strip Board
Ceramic with mass terminated connectors.
  80% copper, 20% space.
Booster
  Beam sigma 4.5 mm
  60 Strips at 1.5 mm
  1.2E12 to 4.5E12 protons
Main Injector/Recycler
  Beam sigma 4.5 - 1.5 mm
  120 Strips, pitch 0.5 mm
  MI up to 6 booster batches
  RR slip stack up to 12 batches
  RR max intensity 5E13
Tevatron
  Beam sigma 1.5 mm
  128 strips, pitch 0.25 mm

## Slide 15: Mark-III Model
Vacuum Vessel with Instrument
Now being assembled for installation.
Typical vertical installation.
Same magnet and internals for both horizontal and vertical.
The vacuum vessel moves in the plane of the measurement, while the magnet is fixed.

## Slide 16: Main Injector Mark-III Magnet
Magnet MIIPM001 on measurement stand for integral field map.
Similar to Mark-II but smaller.
1KG center field and half correction up and downstream,
Local 3 bump shunted, so beam sees close to zero integrated field.

## Slide 17: MARK-III Magnet
31.5”
14”
4.25”
Mounted on measurement stand
for field quality map.
Hall Probe shown.

## Slide 18: MI Orbit Perturbation
Measured magnet integrated field is -0.001T-m
Maximum displacement around the ring for the measured field integral is
For the Main Injector ρm ≈ 27 T-m
and the maximum β is 50, Tune, ν is 0.43,
D ≈ 0.001 m

## Slide 19: Magnet Measurements
B Field Line
Maximum Deviation

## Slide 20: Magnet Measurements
0.002 T
0.004 T
Old Shunt
New Shunt
IPM Active Region

## Slide 21: Magnet Measurements
B Field Line
Deviation from top to bottom
Average value of 200 mm could be hall probe rotation; corresponds to ~0.1 degrees

## Slide 22: Mark-III Vacuum Vessel

## Slide 23: IPM New Installation’s
New Main Injector
  Magnetic Mark-III vertical at Q103
  Mark-II internal parts will eventually be retrofitted.
New Recycler Magnetic Mark-III
  Horizontal at Q104
  Vertical at Q103
Booster
  Will have 2 – 30KV Electrostatic cans available

## Slide 24: IPM Measurement Capability
All Systems
  Turn by turn measurements.
    Turns could be averaged for any accuracy desired.
  Used for injection tuning/matching.
    Routinely used for first 500 turns to see injection oscillations.
    Sigma measurements anywhere in the cycle.
  Collected 65K samples @ 1 per revolution
    Booster 19900 turns for a full cycle.

## Slide 25: IPM Measurement New Features
New Main Injector
  Higher speed 16 channel digitizers 80MHz
    multiple sample each batch for better accuracy/sample
    Allows for digital filtering of signals on A/D
  96 Channels to be sampled using new Brian Fellenz 20 channel preamp module.
  Control Grid to gate off electrons for unmeasured batches
    Should significantly increase MCP life time.
  2000 samples at either
    1 Batch per revolution
    Spread across all batches for about 300 turns

## Slide 26: Typical Data Display
Main Injector:
Injection tuning study.
Showing injection oscillations for the first 300 turns.

## Slide 27: Typical Data Display cont’d
Main Injector
P-Bar injection tuning.

## Slide 28: Typical Data Display Last
Booster LabView Front End
Main Injector Console Application
Horizontal Vertical
Top left trace indicates intensity.
Bottom left 2 plots –
can plot sigma and position,
or individual turn profiles.

## Slide 29: MCP Test Chamber
Facility to scan MCPs for suitability and look at areas of reduced gain

## Slide 30: Test Chamber Measurements

## Slide 31: The Players
Instrumentation
  Dave Slimmer, Carl Lundberg, Jim Galloway, Brian Fellenz, Dan Schoo, John VanBogaert, Alexei Semenov
Main Injector, etc…
  Bruce Brown, Denton Morris, Jim Volk
Mechanical Support
  Matt Alvarez, Tom McLaughton, Dave Tinsley, Kevin Duel, Linda Valerio, Eric Pirtle, Jim Wilson, James Williams, Sali Sylejmani, Scott McCormick, Debbie Bonifas, Tom Lassiter
Technical Division
  David Harding, Oliver Kiemschies, Bill Robotham, Vladimir Kashikhin, Bill Robotham, Michael A. Tartaglia, Mark D Thompson, Gueorgui Velev, Dana Walbridge

## Slide 32: Intermission

## Slide 33: Gated IPM Concept
Problem with MCP is short lifetime
  Plate is using up lifetime whenever beam is in the machine and the IPM voltage is on
  Voltage takes a while to raise and lower
Would like to be able to gate the charge to preserve the MCP
  Stop the electrons and ions from reaching the MCP
  Allow the electrons and ions an escape path from the IPM active region
    i.e. no Penning traps

## Slide 34: Gated IPM Concept
  The force on a charged particle is
  Assume that and
  The solution to this is circular motion in the plane, constant motion along and a drift along which in this case is , i.e. along the beam
  Putting in the values for the electric and magnetic fields gives us a drift velocity of ~10 cm/ms along the proton beam direction
    The electrons will have drifted beyond the MCP in ~1-2 ms

## Slide 35: MATLAB Simulation
Simulation tracks particles through arbitrary E and B fields
Uses interpolation to obtain the fields at any point from previously calculated field distributions
Propagates using a relativistic formula
Invert

## Slide 36: Matlab Simulation
Once the acceleration is determined, a discrete evaluation of the differential equation of motion is used to step the particles
The magnetic and electric fields are handled separately
  Magnetic contribution to the motion is only applied to the components perpendicular to the B field
  Magnitude of the velocity perpendicular to the B field is forced to be preserved, since the B field does no work
    This in particular helps with the tight spirals along the field lines

## Slide 37: Matlab Simulation
The electric and magnetic fields of the bunch are calculated before hand for various bunch parameters
  Shifted as a function of time to represent the moving beam
Electric field of IPM from a 2-D Poisson calculation
Magnetic field from 3-D magnet model
Ionized particle distributions are random in emission angle with 1/E2 energy distribution

## Slide 38: Magnetic Field in Simulation
Measured
Model
0.0004 T
0.0005 T

## Slide 39: Gated-on IPM
Magnet with vertical B field
Cathode
Field shaping electrodes
Electron Suppression Grid
Wire mesh gate
Microchannel Plate (MCP)
Anode strips
B Field ~ 1 kg
E Field ~ 100 kV/m
ON
Electrons spiral down helically

## Slide 40: Gated-on IPM
Particles originating from single point
(resolution contribution)
Elapsed time ~ 1.7 ns
Anode Strip

## Slide 41: Gated-on IPM
Bunch offset refers to x
Particles originating from single point
(resolution contribution)

## Slide 42: Gated-on Expected Signal
From figure 7 of Sauli #, the number of primary ion pairs produced in one centimeter of a gas species i at one atmosphere of pressure by one minimum ionizing particle can be roughly parameterized as
Expressing this in terms of the proton bunch parameters and partial pressures in the beampipe one arrives at
At the peak of a Main Injector bunch, the number of ionization electrons is ~10 per anode strip (no MCP gain)
#F. Sauli, “Principles of Operation of Multiwire Proportional and Drift Chambers”, CERN 77-09, 3/5/77.

## Slide 43: Gated-off IPM
Magnet with vertical B field
Cathode
Field shaping electrodes
Electron Suppression Grid
Wire mesh gate
Microchannel Plate (MCP)
Anode strips
B Field ~ 1 kg
E Field ~ 10 kV/m
OFF
Electrons propagate into or out of the page

## Slide 44: Gated-off Fields
X Component of E field
Y Component of E field

## Slide 45: Gated-off Motion
Electrons drift along beam direction
Single particle
Particle origination point

## Slide 46: Gated-off Behavior
Bunch Centers
Drift Velocity
1.2 cm / 150 ns = 8 cm/ms
Compared to
10 cm/ms analytically estimated
Y motion vs time
What happens when electrons reach the end of the field region?

## Slide 47: Gated-off Ion Paths
Elapsed time is ~1.5 ms
Initially appears ok, since ions do not go much beyond the gating grid
-- However --
Secondaries from ion impacts on gating grid could be a problem

## Slide 48: Electron Beam Profiler
Deflection vs. Angle provides information about the proton beam transverse profile
Proton beam out of page
Electron beam
Angle
Deflection
Electron beam is deflected by electric and magnetic fields of the protons
Increasing beam power in MI/RR implies the need for non-invasive instrumentation
  Electron beam deflection technique is one choice (working implementation at SNS)

## Slide 49: Techniques for Main Injector
Various techniques for measuring deflection
  Fast scan through peak of bunch (similar to SNS)
    Requires fast deflector (< 1 ns sweep time)
  Slow scan, akin to flying wires (most likely solution for Nova)
    Position the beam and record the maximum deflection as the beam passes by
      Leave the electron beam stationary
      Sweep the beam along the proton direction
        Obtain longitudinal distribution
Collaborating with Wim Blokland at SNS who has done simulations of the various techniques

## Slide 50: Electron Deflection
Plots courtesy of Wim Blokland
Slow electron sweep
Position the electron beam
Record the deflection of a bunch
Move the electron beam and repeat

## Slide 51: Electron Deflection Simulation
Plots courtesy of Wim Blokland
Step the electron beam through the proton beam and record maximum deflections
Derivative of deflection vs. position is nominally beam profile
Derivative

## Slide 52: Alternative Deflection Scheme
Electron Sweep
Sweep the electron beam along the proton bunch
Sweep duration coincides with the duration of the proton bunch
Beam Sim. Longitudinal s = 2 ns
Meas. Sim. Longitudinal s = 2.3 ns
Beam Simulated Transverse s = 3 mm
Meas. Simulated Transverse s = 3.5 mm

## Slide 53: Simulated Camera Image
Camera frames are ~30 ms
Main Injector cycle is ~1 s
Need to step many times per frame to accumulate data fast enough for measurement
  Complicated to extract each step

## Slide 54: Electron Beam Device
Ion Pump
60 KeV Electron Gun
Kimball Physics
Beam Valve
Electrostatic Deflector
Ion Gauge
Ion Gauge
Pneumatic
Insertion Device
with OTR Stainless
Steel Mirror
Phosphor Screen
P47, 400 nm, 60 ns decay(?)
Optical Breadboard
~ 60 cm x 150 cm
Elliptical Main Injector
beampipe
Optic components box

## Slide 55: Optics
Two optical paths: OTR screen and Phosphor screen with some shared elements OTR screen is inserted at location of proton beam (sans proton beam) and used to focus the electron beam and measure the electron beam spot size
SS Mirror
45 Degree Mirror
Lens on moving
stage
Filters and polarizers
Image Intensifier
(Hamamatsu)
Camera Objective
CID Camera
Megarad version
Phosphor screen
45 Degree Mirror
Lens on moving stage
45 Degree Mirror on moving stage
Stationary Lens
45 Degree Mirror

## Slide 56: Devices
Solenoid and steering magnets
Cathode
Thermionic Electron Gun
Electrostatic Deflector
Kimball Physics up to 60 KeV
(we will use up to 15 KeV)
6 mA, pulsed, 1 ms to DC
LaB6 cathode, 100 mm spot size
15 cm long ‘circular’ plates
~2.5 cm diameter
Plates

## Slide 57: Test Stand at NWA
Optical Transition Radiation Screens
Electron Gun
Lens / Digital Camera Imaging Systems
Faraday Cup

## Slide 58: Gun Tests
Gun has internal solenoid
  Scanned beam through waist at first screen
Scanned beam sizes from Ce:YAG screens (1 A beam)
Scanned beam sizes from OTR screens (1 mA beam)

## Slide 59: Test of Electrostatic Deflector
Deflector Pulse
Deflector Pulse
15 cm long plates
~120 V
~190 V
Deflecting Voltage vs. Deflector Length
500 V
80 ns

## Slide 60: Electrostatic Deflector Test
Short sweep
Effect is similar proton bunch passing by
Longer sweep
Bright part off screen
Beam size not uniform
  Possibly due to poor pulse quality

## Slide 61: Magnetic Fields in Tunnel
Electron beam
Bx
By
5 gauss
0 g
Need to shield beamline
First attempt will be mu metal
Quad busses
3500 A
Dipole busses
9000 A

## Slide 62: Electron Beam Profiler
Would like to install as soon as possible
But…
  Priority is “Very Low” (to put it politely)
  Relies on the “kindness of strangers”

## Slide 63: Strangers (and not so strangers)
Instrumentation
  Amber Johnson, Carl Lundberg, Jim Galloway, Jim Fitzgerald, Peter Prieto, Pierpaolo Stabile, Andrea Saewart, Dave Slimmer, Dehong Zhang, Brian Fellenz, Alex Lumpkin
Mechanical Support
  Wade Muranyi, Brad Tennis, Elias Lopez, Debbie Bonifas, Scott McCormick, Ryan Montiel, Sali Sylejmani, Tom Lassiter, James Williams, John Sobolewski, Matt Alvarez, Kevin Duel
Summer Students
  Paul Butkovich, Khalida Hendricks, Danila Nikiforov
Others
  Charles Thangaraj, Dave Burk, Dennis Schmitt

## Slide 64: Extras

## Slide 65: Optics
Phosphor screen
Image Intensifier

## Slide 66: Optical Transition Radiation
Light yield over the 2 ms electron pulse
Initial beam images determined to be blackbody
  No polarization
  Intensity increased nonlinearly with duration
  Damage to stainless steel mirror observed
Electron energy low
  Broad angular distribution
  Mirror should be 15 instead of 45
(E. Bravin, private communication)

## Slide 67: Wire Tests
Wire to simulate proton beam
e Beam pulsed on for 40 ms
Wire pulsed for 20 ms
Half the time the beam is deflected
0V
150V
50V
250V
200V
300V
100V