## Slide 1
John Kuharik
Booster 2nd Harmonic Cavity Meeting
March 16, 2017
Garnet Witness Sample Testing

## Slide 2: Calculations for B, H, Mu, with and without Air Gap Correction
Calculation parameters:
nprim = 183 turns on the excitation coil (183turns/0.0179m)*(0.016m)=163 turns
nsec = 80 turns on the measurement coil
A = 0.000249 m2 Area of the garnet sample
Acoil = 0.000269 m2 Area of the measurement coil
L = 0.01603 m Length of sample. In a solenoid field calculation, L = length of solenoid. Because of the ferrite box, we use L=sample length.
  R = 3.92 ohm primary resistor.
Oefact = 0.01256637 factor for conversion from A/m to Oersted.
Gaussf = 10000 factor for conversion from Tesla to Gauss.
Vrp = Voltage across the primary side resistor
Vs = Measurement signal voltage

## Slide 3: Data Processing
Read raw data from file and select region of interest t= 0.000256 ->0.00172s

## Slide 4: Calculate scope DC offset in data files
Take the mean of each signal in the pre-pulse region and subtract from the whole signals.
In this case, pre-pulse interval is t= -0.0005s to t=0.000255s, and the offsets are:
Vs offset correction: -0.00664906
Vrp offset correction: 0.00660004

## Slide 5: Calculate H in the region of interest

## Slide 6: Calculate dB from Faraday’s Law
If we take into account the air gap between measurement coil and sample:

## Slide 7: Calculate B
With no air gap correction:
With the air gap correction:
Multiply B by Gaussf = 10,000 to get B in Gauss.

## Slide 8: Plot B(H) and Bcorrected(H)
Calculate Mu and Mucorrected:

## Slide 9: Compare results to theory
N=183 turns
N=163 turns

## Slide 10: Interpolate and Consolidate Data for Multiple Runs
Five runs at 19 amps processed as described above. H and B for each run interpolated and Mu calculated (without the air gap correction). The average measurement with standard error is shown below along side the theory prediction.
Chart: Mu Measure vs Mu Theory vs B(G)
Mu Theory | 50 | 49.7 | 46.5 | 25 | 9 | 5 | 3 | 2.3 | 2 | 1.66 | 1.47
This is with N=183 turns