## Slide 1: Miscellaneous.
2nd Harmonic cavity meeting
June 8, 2017
Gennady Romanov

## Slide 2: Angular recess
The shim with tiny rectangular edges is slightly recessed in the main alumina substrate. Then the shim edges are covered with additional alumina disks
Requires two “recess” alumina disks with angular edges.
Obviously a good tight connection between parts without air pockets is hard to make
The best recess design, but the most complicated
The most practical recess (if we decide to make a recess at all). Depth is 0.02” .
The model sent to Kevin.

## Slide 3
Unloaded Q
Chart: Unloaded Q
| 0 | 25 | 50 | 75 | 100 | 150 | 200 | 300 | 400 | 500 | 600 | 700 | 800
Q_0 | 3161.7804 | 3922.7792 | 4559.628 | 5038.2417 | 5381.2201 | 5524.029 | 5581.6632 | 5285.2427 | 4915.9065
Chart: Unloaded Q
| 0 | 25 | 50 | 75 | 100 | 150 | 200 | 300 | 400 | 500 | 600 | 700 | 800
 | 3252.71 | 3693.6962 | 4075.2987 | 4378.1168 | 4637.3462 | 5002.1535 | 5226.909 | 5497.0726 | 5636.1151 | 5733.0775 | 5752.9633 | 5777.4392 | 5779.1032
Discrete port in the tetrode,
reflection method. S11 is interpreted as a scalar?
Two probes, discrete port removed, transmission method
Two probes
Transmission method.
The final table of the cavity RF parameters will be posted in our share point folder after revision.

## Slide 4: MDC probe
0.25”
0.55”
1300 V/m
10 V
MDC Type-N, up to 1 GHz,
500 V dc, 3 A
0.68”

## Slide 5: Slots
Chart: Tuning curve
| 0 | 25 | 50 | 75 | 100 | 150 | 200 | 300 | 400 | 500 | 600 | 700 | 800
With slots | 75.687 | 80.643 | 84.542 | 87.714 | 90.286 | 94.181 | 96.963 | 100.7 | 103.15 | 104.86 | 106.1 | 107.044 | 107.78
Without slots | 75.767 | 80.663 | 84.599 | 87.758 | 90.323 | 94.205 | 96.979 | 100.71 | 103.16 | 104.864 | 106.1 | 107.09 | 107.78
Chart: Q0 (I_dyn)
| 0 | 25 | 50 | 75 | 100 | 150 | 200 | 300 | 400 | 500 | 600 | 700 | 800
With slots | 3255 | 3694.8 | 4073.151 | 4359.8 | 4624.1 | 4996.8 | 5232 | 5509 | 5641 | 5687.2 | 5742.5 | 5762.9 | 5772.9
Without slots | 3252.71 | 3693.6962 | 4075.2987 | 4378.1168 | 4637.3462 | 5002.1535 | 5226.909 | 5497.0726 | 5636.1151 | 5733.0775 | 5752.9633 | 5777.4392 | 5779.1032
Chart: Rsh (I_dyn)
| 0 | 25 | 50 | 75 | 100 | 150 | 200 | 300 | 400 | 500 | 600 | 700 | 800
With slots | 185.5 | 210.4 | 231.97 | 252.35 | 268.5 | 295.4 | 313.1 | 336.15 | 350.4 | 356.8 | 360 | 362.9 | 365.7
Without slots | 185.5 | 210.55 | 232.1 | 252.52 | 269 | 296 | 313.9 | 336.4 | 349.1 | 359.7 | 363.9 | 370.2 | 368.5
Barely noticeable degradation with perfectly aligned slots