NuMI Horn Longitudinal Field-Mapping System
Adam Watts
TSD Topical Meeting
19 September 2019

---
Comparison Slide [28pt Bold]
Motivation
For quality control, new NuMI horns go through a field-mapping process to ensure required field quality while the horn pulses in the MI-8 test cage.
This talk focuses on the mapping system for the field along the center axis of the horn, where a 3-D hall probe is scanned longitudinally through the pulsing horn to measure what should be a flat and very low magnetic field.
The original implementation of this system relied on proprietary electronics and software that have been obsoleted by the vendor, and field measurements were written down by hand during the measurement. An open-source alternative was sought to allow continued development and use of the measurement system without relying on the vendor; unification of the motor drive computer and field probe DAQ will greatly speed up the mapping process.

---
Original System, Mechanical
Three-axis control
Parker Zeta83-135-MO stepper motors
382 oz-in (Nm) static torque
NEMA 34 frame
10.47 oz-in rotor inertia
3/8" shafts
Not well documented at all. Unclear what the motor phase current rating is, though NEMA 34 standard is a clue.
Linear encoders, one per axis. Connectors appear damaged.
No apparent motion stops or limit switches. Unclear if motors are slip-clutch coupled.

---
Original System, Electronics
Stepper motor electronics:
ZETA4 stepper motor system from Parker Motion
Proprietary, uses “Motion Architect” software to talk to controllers
Motion architect only runs on Windows XP or older
Requires: legacy Windows OS, PC with serial ports or USB-to-serial adapters
One controller module (pictured) per motorMagnetic field probes:
Single 3D probe
Probe reading manually recorded (i.e. paper and pencil!)

---
New System, Overview

---
New System, CPU
“Beaglebone Black”
Credit-card-sized single-board Linux computer, open-source hardware with open-source software and OS
Debian or Ubuntu Linux
1 GHz ARM Cortex-A8 CPU
512MB DDR3 RAM
4GB 8-bit eMMC on-board flash storage
2x PRU 32-bit microcontrollers
Ethernet, micro-HDMI, USB
8 ADC pins, 0-1.8V, 12-bit resolution
62 GPIO pins
GPIO/ADC open-source Python library freely available and supported
External storage: microSD or USB flash

---
New System, Motor Controllers
Geckodrive GM210X Step Motor Drive
80VDC, 7A maximum output to motor, one controller per motor
Motor power can be set with onboard DIP switches
Built-in pulse multiplier, handles acceleration/deceleration curves, opto-isolated inputs to protect CPU board
Driven by CPU board GPIO pin; set GPIO output pin high for ~10 milliseconds, motor controller drives motor set number of steps based on internal DIP switches
Power supply, one per motor 36V 7.0A

---
Motor Motion Sample Code
Python GPIO library for the CPU board is very straightforward
On the right is example code that waits for user input; “s” key moves motor one way, “d” key moves motor the other way.
Move amount per button press determined by DIP switch settings for the controller’s pulse-multiplier board. Can also modify the code to send multiple pulses per button press and speed up movement if necessary.

---
New System, Linear Position Sensors
- Linear encoders (old system) difficult to read, as magnetic transitions must be digitized at high rate and counted.
- Simpler solution: industrial linear slide potentiometer.
- Megatron MSL38, enviromentally-sealed for industrial applications, models from 100mm to 2m effective electrical travel length.
- By applying a reference voltage and reading the divider ration from the potentiometer, the position can be measured asynchronously with the slow built-in ADCs in the CPU board.
- Linear to ±0.05% over electrical length. Resolution limited by ADC of CPU board (12-bit).

---
New System, Linear Position Sensors
A buffer circuit is necessary to prevent too much current drawn through the wiper, as per datasheet. Reference voltage is taken from the CPU board’s 3.3V rail, which is also the reference for the ADC pins.

---
New System, Linear Position Sensors
Sensor mounted to axial measurement sled. Thanks Clay!
The plan is to verify this sensor is what we want to use, then purchase the other two and install them.

---
Data Acquisition
Jim would like the 3D hall probe signals and horn stripline currents digitized @ >100 points per pulse
Most reliable DAQ for this is an ethernet-capable oscilliscope
Network switch creates LAN between CPU board and two scopes with static IP addresses
Simple Python code to communicate to scopes, configure their settings, and capture data

---
Progress report
Completed
Motor movement verified with new controller and CPU board
DAQ verified at requested rate using scopes over local Ethernet LAN, acquired and saved by CPU board
Housekeeping: old electronics uninstalled, new monitor, keyboard/mouse, power supplies, etc. installed.
X-axis linear sensor installed (thanks Clay!)
Linear sensor buffer circuit built and connected to CPU board
Sensor cables terminated, terminal block and patch panel made for PC cart so there’s adequate strain relief.
To do
Terminate new cables for motors, create patch panel for their connectors into motor controllers (extra cables for limit switch?)
Verify linearity of X-axis sensor and calibrate to position on sled
Order Y and Z sensors, connect to sled
Cobble existing code together for a user-friendly scan script
Dedicated scopes for measurement setup (discussed with Bob)
Determine whether we need to use hard stops, slip clutches, limit switches (software or “dead-man”) or some combination thereof
GUI

---
Thank you!