## Meson checklist
Last edited: 10/30/19 | Meson Startup Checklist
Shutdown Work
All shutdown work completed
Vacuum
All roughing pumps turned on with roughing valves open and pressures at operational levels
I55, Switchyard, <Meson T>
I55, Switchyard, <Meson C>
LCW Systems
MS2LCW system up an running (S18 MESON <14>)
Verify that all operational cooling tower fans run successfully (F:MS2CTB off to enable manual control)
Make sure temperature regulation is enabled (F:MS2CTB)
F:MS2RSYS > 5 MOhm to be able to run power supplies
M01 Raw system up and running (local control @ MS1); needs MS2 up first
Instrumentation
BLM high voltage supplies on; remote readback only, control is local
(S18 M_TUNE <21>) ~500 V
Profile monitors
MWMS1 front-end up and running (D31)
Sequences enabled (S45 Meson Line and Meson Center)
Verify timing event configuration, background seen upon trigger (S45 Meson Line and Meson Center)
Monitors moved in and out of beam path to verify motion (S45 Meson Line and Meson Center)
High voltages on (S18 M_Quads <10>)
Scintillator counter high voltages (local @ MS4)
Ion chamber and SEM high voltages
MW1SEM (local @ MS1)
MC1SEM (local @ MS1)
MT4IC (local @ MS3)
MC6IC (local @ MS4)
ArCO2 valved-in and operational (local @ MS2, pressure readback F:MS2GAS)
Motion Control
Verify position readbacks and motion
F-seps (S18 SY120 <15>)
MT4 momentum collimators: F:MT4CH1, CH2, CV1, CV2 (S18 SY120 <9>)
Meson primary target F:MW1TGT (S17 MTEST1 <4>)
Meson secondary target F:MT4TGT (S17 MTEST1 <4>)
M03 high-rate tracking area motion table: F:MT3HT, F:MT3VT (S17 MTEST1 <4>)
MTest pinhole collimator: F:MT3PUH, F:MT3PUV, F:MT3PDH, F:MT3PDV (S17 MTEST1 <4>)
MTest lead electron absorber F:MT4PB (S17 MTEST1 <4>)
Magnet power supplies
Power supplies configuration control unlocked by MCR
All magnets verified valved-in to LCW
All magnet power supplies verified valved-in to LCW
Ask MCR Crew Chief to put a SY event in the timeline to test power supply ramps
Ask MCR to run turn-on sequencers for Meson, placing MTest in proton, HE pion, and LE pion modes, and MCenter at 64 GeV.
Alarms
Validate all Meson alarm lists and active/bypassed statuses
Signature/Date:

## Copy of Power Supply
Power Supply Checkout Form
Beamline: | Meson
Enclosure: | M01, M02, M03, M05, MC6
Initials/Date:
Power Supply Location | Power Supply Name | Power Supply Labeled | LCW valved on @ supply | P.S. Leads Connected | Date | Comments
MS1 | F:M00U
F:M01D-A
F:MW1W
F:MC1D
F:M00H
F:M00V
F:Q230
MS2 | F:MT2WD2
F:MC2V
F:MC2H
F:MT2V
F:MT2WD1
F:MT2WU
F:MC2Q2
F:MC2Q1
F:MT2Q2
F:MT2Q1
MS3 | F:MT3U
F:MT3Q1
F:MT3Q2
F:MT3Q3
F:MT3Q4
F:MT4Q1
F:MT4Q2
F:MT4Q3
F:MT4Q4
F:MT4Q5
F:MT4HT
F:MT3V
F:MT4VT
F:MT3W
F:MT3SW
F:MT4W
MS4 | F:MT4Q6
F:MT5Q1
F:MC6Q1
F:MC6Q4
F:MC6Q3
F:MC6Q6
F:MC5Q1
F:MC5Q2
F:MC5H1
F:MC5V1
F:MT5Q2
F:MC5U
F:MC6H2
F:MC6V1
F:MT5VT1
F:MC6V2
F:MC6H1
F:MT5E
F:MT5VT2
F:MT5HT2
F:MC6D
F:MC6Q2
F:MC6Q5